## cv_conc_table_2sd_filtered_labs
SampleType | ceramideName | Calibration | Number_Of_Labs | Lab_Median_IntraCV_Conc_% | Lab_Average_Conc | Lab_Standard_Deviation_Conc | Lab_CV_Conc_% | Concentration_Unit
NIST hTAG | Cer d18:1/16:0 | Calibration Curve | 26 | 6 | 0.13954 | 0.027658 | 20 | pmol/µL [µM]
NIST hTAG | Cer d18:1/16:0 | Single Point | 25 | 6 | 0.20187 | 0.026322 | 13 | pmol/µL [µM]
NIST hTAG | Cer d18:1/18:0 | Calibration Curve | 26 | 5 | 0.047903 | 0.0096679 | 20 | pmol/µL [µM]
NIST hTAG | Cer d18:1/18:0 | Single Point | 26 | 6 | 0.032927 | 0.0045304 | 14 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:0 | Calibration Curve | 27 | 5 | 1.8454 | 0.36848 | 20 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:0 | Single Point | 25 | 4 | 1.967 | 0.25691 | 13 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:1 | Calibration Curve | 26 | 5 | 0.62722 | 0.10388 | 17 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:1 | Single Point | 25 | 5 | 0.61374 | 0.11623 | 19 | pmol/µL [µM]
NIST SRM | Cer d18:1/16:0 | Calibration Curve | 26 | 5 | 0.12259 | 0.022595 | 18 | pmol/µL [µM]
NIST SRM | Cer d18:1/16:0 | Single Point | 26 | 4 | 0.17558 | 0.024562 | 14 | pmol/µL [µM]
NIST SRM | Cer d18:1/18:0 | Calibration Curve | 26 | 5 | 0.04024 | 0.0079856 | 20 | pmol/µL [µM]
NIST SRM | Cer d18:1/18:0 | Single Point | 26 | 5 | 0.027931 | 0.0030502 | 11 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:0 | Calibration Curve | 27 | 4 | 1.1865 | 0.24093 | 20 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:0 | Single Point | 25 | 4 | 1.2823 | 0.14009 | 11 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:1 | Calibration Curve | 27 | 4 | 0.43245 | 0.081488 | 19 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:1 | Single Point | 24 | 4 | 0.41278 | 0.063041 | 15 | pmol/µL [µM]
NIST T1D | Cer d18:1/16:0 | Calibration Curve | 25 | 5 | 0.10867 | 0.022715 | 21 | pmol/µL [µM]
NIST T1D | Cer d18:1/16:0 | Single Point | 25 | 5 | 0.15646 | 0.024774 | 16 | pmol/µL [µM]
NIST T1D | Cer d18:1/18:0 | Calibration Curve | 26 | 5 | 0.046323 | 0.0083236 | 18 | pmol/µL [µM]
NIST T1D | Cer d18:1/18:0 | Single Point | 26 | 5 | 0.031901 | 0.0036032 | 11 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:0 | Calibration Curve | 27 | 4 | 1.1679 | 0.24206 | 21 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:0 | Single Point | 25 | 4 | 1.2583 | 0.14225 | 11 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:1 | Calibration Curve | 27 | 5 | 0.42661 | 0.085965 | 20 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:1 | Single Point | 25 | 4 | 0.41932 | 0.087135 | 21 | pmol/µL [µM]
NIST YAA | Cer d18:1/16:0 | Calibration Curve | 26 | 8 | 0.071704 | 0.020582 | 29 | pmol/µL [µM]
NIST YAA | Cer d18:1/16:0 | Single Point | 25 | 7 | 0.10961 | 0.019206 | 18 | pmol/µL [µM]
NIST YAA | Cer d18:1/18:0 | Calibration Curve | 27 | 7 | 0.022819 | 0.0076475 | 34 | pmol/µL [µM]
NIST YAA | Cer d18:1/18:0 | Single Point | 25 | 6 | 0.016688 | 0.0020853 | 12 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:0 | Calibration Curve | 27 | 4 | 0.8428 | 0.29369 | 35 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:0 | Single Point | 26 | 4 | 0.85674 | 0.17364 | 20 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:1 | Calibration Curve | 27 | 6 | 0.259 | 0.078324 | 30 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:1 | Single Point | 27 | 5 | 0.262 | 0.086174 | 33 | pmol/µL [µM]

## cv_conc_table_all_labs
SampleType | ceramideName | Calibration | Number_Of_Labs | Lab_Median_IntraCV_Conc_% | Lab_Average_Conc | Lab_Standard_Deviation_Conc | Lab_CV_Conc_% | Lab_Filtered_CV_Conc_% | Concentration_Unit
NIST hTAG | Cer d18:1/16:0 | Calibration Curve | 28 | 6 | 0.13912 | 0.034265 | 25 | 20 | pmol/µL [µM]
NIST hTAG | Cer d18:1/16:0 | Single Point | 28 | 6 | 0.19436 | 0.051707 | 27 | 13 | pmol/µL [µM]
NIST hTAG | Cer d18:1/18:0 | Calibration Curve | 28 | 6 | 0.047868 | 0.012538 | 26 | 20 | pmol/µL [µM]
NIST hTAG | Cer d18:1/18:0 | Single Point | 28 | 6 | 0.033291 | 0.0093765 | 28 | 14 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:0 | Calibration Curve | 28 | 5 | 2.0918 | 1.3531 | 65 | 20 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:0 | Single Point | 28 | 5 | 1.907 | 0.50408 | 26 | 13 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:1 | Calibration Curve | 28 | 5 | 0.62275 | 0.14659 | 24 | 17 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:1 | Single Point | 28 | 5 | 0.59856 | 0.18726 | 31 | 19 | pmol/µL [µM]
NIST SRM | Cer d18:1/16:0 | Calibration Curve | 28 | 5 | 0.11688 | 0.030341 | 26 | 18 | pmol/µL [µM]
NIST SRM | Cer d18:1/16:0 | Single Point | 28 | 5 | 0.16545 | 0.044098 | 27 | 14 | pmol/µL [µM]
NIST SRM | Cer d18:1/18:0 | Calibration Curve | 28 | 6 | 0.04026 | 0.011383 | 28 | 20 | pmol/µL [µM]
NIST SRM | Cer d18:1/18:0 | Single Point | 28 | 5 | 0.028325 | 0.0084698 | 30 | 11 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:0 | Calibration Curve | 28 | 4 | 1.3516 | 0.90531 | 67 | 20 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:0 | Single Point | 28 | 4 | 1.2376 | 0.32344 | 26 | 11 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:1 | Calibration Curve | 28 | 4 | 0.42009 | 0.10332 | 25 | 19 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:1 | Single Point | 28 | 4 | 0.41192 | 0.13158 | 32 | 15 | pmol/µL [µM]
NIST T1D | Cer d18:1/16:0 | Calibration Curve | 28 | 6 | 0.10547 | 0.031909 | 30 | 21 | pmol/µL [µM]
NIST T1D | Cer d18:1/16:0 | Single Point | 28 | 6 | 0.15052 | 0.044014 | 29 | 16 | pmol/µL [µM]
NIST T1D | Cer d18:1/18:0 | Calibration Curve | 28 | 6 | 0.046565 | 0.013131 | 28 | 18 | pmol/µL [µM]
NIST T1D | Cer d18:1/18:0 | Single Point | 28 | 6 | 0.03255 | 0.0105 | 32 | 11 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:0 | Calibration Curve | 28 | 4 | 1.3321 | 0.9012 | 68 | 21 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:0 | Single Point | 28 | 4 | 1.2175 | 0.32163 | 26 | 11 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:1 | Calibration Curve | 28 | 5 | 0.41461 | 0.10559 | 25 | 20 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:1 | Single Point | 28 | 5 | 0.40676 | 0.13039 | 32 | 21 | pmol/µL [µM]
NIST YAA | Cer d18:1/16:0 | Calibration Curve | 28 | 7 | 0.071458 | 0.02721 | 38 | 29 | pmol/µL [µM]
NIST YAA | Cer d18:1/16:0 | Single Point | 28 | 7 | 0.10574 | 0.032206 | 30 | 18 | pmol/µL [µM]
NIST YAA | Cer d18:1/18:0 | Calibration Curve | 28 | 6 | 0.023837 | 0.0092395 | 39 | 34 | pmol/µL [µM]
NIST YAA | Cer d18:1/18:0 | Single Point | 28 | 6 | 0.017431 | 0.0056514 | 32 | 12 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:0 | Calibration Curve | 28 | 5 | 0.95382 | 0.65437 | 69 | 35 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:0 | Single Point | 28 | 5 | 0.88399 | 0.34915 | 39 | 20 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:1 | Calibration Curve | 28 | 6 | 0.27447 | 0.11228 | 41 | 30 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:1 | Single Point | 28 | 5 | 0.27713 | 0.11645 | 42 | 33 | pmol/µL [µM]

## avg_conc_tbl_not_filtered_labs
SampleType | ceramideName | Calibration | LabId | MassAnalyzerType | Number_Of_Samples | Sample_Average_Conc | Sample_Standard_Deviation_Conc | Sample_CV_Conc_% | IsInDataRange | Lab_Average_Conc | Lab_Left_Bound | Lab_Right_Bound | Concentration_Unit
NIST hTAG | Cer d18:1/16:0 | Calibration Curve | 02a | QQQ | 6 | 0.15944 | 0.0020107 | 1.2611 | TRUE | 0.13912 | 0.070594 | 0.20765 | pmol/µL [µM]
NIST hTAG | Cer d18:1/16:0 | Calibration Curve | 02b | QQQ | 18 | 0.15981 | 0.0074313 | 4.65 | TRUE | 0.13912 | 0.070594 | 0.20765 | pmol/µL [µM]
NIST hTAG | Cer d18:1/16:0 | Calibration Curve | 10a | QQQ | 6 | 0.075616 | 0.0050338 | 6.6571 | TRUE | 0.13912 | 0.070594 | 0.20765 | pmol/µL [µM]
NIST hTAG | Cer d18:1/16:0 | Calibration Curve | 10b | QQQ | 6 | 0.093287 | 0.010958 | 11.7469 | TRUE | 0.13912 | 0.070594 | 0.20765 | pmol/µL [µM]
NIST hTAG | Cer d18:1/16:0 | Calibration Curve | 12 | QQQ | 18 | 0.1565 | 0.013982 | 8.934 | TRUE | 0.13912 | 0.070594 | 0.20765 | pmol/µL [µM]
NIST hTAG | Cer d18:1/16:0 | Calibration Curve | 14 | QQQ | 18 | 0.13762 | 0.0053647 | 3.8983 | TRUE | 0.13912 | 0.070594 | 0.20765 | pmol/µL [µM]
NIST hTAG | Cer d18:1/16:0 | Calibration Curve | 15 | QQQ | 18 | 0.14446 | 0.0047995 | 3.3224 | TRUE | 0.13912 | 0.070594 | 0.20765 | pmol/µL [µM]
NIST hTAG | Cer d18:1/16:0 | Calibration Curve | 16 | Orbitrap | 18 | 0.1366 | 0.014191 | 10.3892 | TRUE | 0.13912 | 0.070594 | 0.20765 | pmol/µL [µM]
NIST hTAG | Cer d18:1/16:0 | Calibration Curve | 17a | TOF | 6 | 0.17526 | 0.010323 | 5.8902 | TRUE | 0.13912 | 0.070594 | 0.20765 | pmol/µL [µM]
NIST hTAG | Cer d18:1/16:0 | Calibration Curve | 18a | QQQ | 18 | 0.14276 | 0.0074252 | 5.2013 | TRUE | 0.13912 | 0.070594 | 0.20765 | pmol/µL [µM]
NIST hTAG | Cer d18:1/16:0 | Calibration Curve | 18b | QQQ | 18 | 0.14555 | 0.0087163 | 5.9886 | TRUE | 0.13912 | 0.070594 | 0.20765 | pmol/µL [µM]
NIST hTAG | Cer d18:1/16:0 | Calibration Curve | 19 | QQQ | 6 | 0.1472 | 0.022334 | 15.1726 | TRUE | 0.13912 | 0.070594 | 0.20765 | pmol/µL [µM]
NIST hTAG | Cer d18:1/16:0 | Calibration Curve | 22 | QQQ | 18 | 0.13811 | 0.0042224 | 3.0573 | TRUE | 0.13912 | 0.070594 | 0.20765 | pmol/µL [µM]
NIST hTAG | Cer d18:1/16:0 | Calibration Curve | 23 | QQQ | 6 | 0.15237 | 0.00436 | 2.8615 | TRUE | 0.13912 | 0.070594 | 0.20765 | pmol/µL [µM]
NIST hTAG | Cer d18:1/16:0 | Calibration Curve | 24 | QQQ | 18 | 0.14614 | 0.0082731 | 5.6609 | TRUE | 0.13912 | 0.070594 | 0.20765 | pmol/µL [µM]
NIST hTAG | Cer d18:1/16:0 | Calibration Curve | 27 | QQQ | 18 | 0.096361 | 0.006184 | 6.4175 | TRUE | 0.13912 | 0.070594 | 0.20765 | pmol/µL [µM]
NIST hTAG | Cer d18:1/16:0 | Calibration Curve | 28 | QQQ | 18 | 0.084247 | 0.036116 | 42.8698 | TRUE | 0.13912 | 0.070594 | 0.20765 | pmol/µL [µM]
NIST hTAG | Cer d18:1/16:0 | Calibration Curve | 29a | TOF | 18 | 0.13752 | 0.0025973 | 1.8887 | TRUE | 0.13912 | 0.070594 | 0.20765 | pmol/µL [µM]
NIST hTAG | Cer d18:1/16:0 | Calibration Curve | 3 | QQQ | 6 | 0.13065 | 0.010837 | 8.2945 | TRUE | 0.13912 | 0.070594 | 0.20765 | pmol/µL [µM]
NIST hTAG | Cer d18:1/16:0 | Calibration Curve | 30 | QQQ | 18 | 0.1936 | 0.027455 | 14.1814 | TRUE | 0.13912 | 0.070594 | 0.20765 | pmol/µL [µM]
NIST hTAG | Cer d18:1/16:0 | Calibration Curve | 31 | QQQ | 18 | 0.14316 | 0.0088642 | 6.1918 | TRUE | 0.13912 | 0.070594 | 0.20765 | pmol/µL [µM]
NIST hTAG | Cer d18:1/16:0 | Calibration Curve | 32 | TOF | 18 | 0.16812 | 0.040772 | 24.252 | TRUE | 0.13912 | 0.070594 | 0.20765 | pmol/µL [µM]
NIST hTAG | Cer d18:1/16:0 | Calibration Curve | 35 | QQQ | 18 | 0.11005 | 0.0089219 | 8.1068 | TRUE | 0.13912 | 0.070594 | 0.20765 | pmol/µL [µM]
NIST hTAG | Cer d18:1/16:0 | Calibration Curve | 36 | QQQ | 18 | 0.14213 | 0.0060151 | 4.2322 | TRUE | 0.13912 | 0.070594 | 0.20765 | pmol/µL [µM]
NIST hTAG | Cer d18:1/16:0 | Calibration Curve | 37 | QQQ | 18 | 0.15844 | 0.01199 | 7.5678 | TRUE | 0.13912 | 0.070594 | 0.20765 | pmol/µL [µM]
NIST hTAG | Cer d18:1/16:0 | Calibration Curve | 38 | QQQ | 18 | 0.054555 | 0.0084132 | 15.4214 | FALSE | 0.13912 | 0.070594 | 0.20765 | pmol/µL [µM]
NIST hTAG | Cer d18:1/16:0 | Calibration Curve | 4 | QQQ | 18 | 0.15316 | 0.003798 | 2.4798 | TRUE | 0.13912 | 0.070594 | 0.20765 | pmol/µL [µM]
NIST hTAG | Cer d18:1/16:0 | Calibration Curve | 7 | Orbitrap | 18 | 0.21274 | 0.01446 | 6.797 | FALSE | 0.13912 | 0.070594 | 0.20765 | pmol/µL [µM]
NIST hTAG | Cer d18:1/16:0 | Single Point | 02a | QQQ | 6 | 0.21343 | 0.0026437 | 1.2387 | TRUE | 0.19436 | 0.090949 | 0.29778 | pmol/µL [µM]
NIST hTAG | Cer d18:1/16:0 | Single Point | 02b | QQQ | 18 | 0.22298 | 0.010286 | 4.613 | TRUE | 0.19436 | 0.090949 | 0.29778 | pmol/µL [µM]
NIST hTAG | Cer d18:1/16:0 | Single Point | 10a | QQQ | 6 | 0.18557 | 0.0086984 | 4.6873 | TRUE | 0.19436 | 0.090949 | 0.29778 | pmol/µL [µM]
NIST hTAG | Cer d18:1/16:0 | Single Point | 10b | QQQ | 6 | 0.18594 | 0.016964 | 9.1233 | TRUE | 0.19436 | 0.090949 | 0.29778 | pmol/µL [µM]
NIST hTAG | Cer d18:1/16:0 | Single Point | 12 | QQQ | 18 | 0.15147 | 0.013541 | 8.9396 | TRUE | 0.19436 | 0.090949 | 0.29778 | pmol/µL [µM]
NIST hTAG | Cer d18:1/16:0 | Single Point | 14 | QQQ | 18 | 0.22313 | 0.0085459 | 3.83 | TRUE | 0.19436 | 0.090949 | 0.29778 | pmol/µL [µM]
NIST hTAG | Cer d18:1/16:0 | Single Point | 15 | QQQ | 18 | 0.19748 | 0.0065023 | 3.2926 | TRUE | 0.19436 | 0.090949 | 0.29778 | pmol/µL [µM]
NIST hTAG | Cer d18:1/16:0 | Single Point | 16 | Orbitrap | 18 | 0.2121 | 0.022446 | 10.5827 | TRUE | 0.19436 | 0.090949 | 0.29778 | pmol/µL [µM]
NIST hTAG | Cer d18:1/16:0 | Single Point | 17a | TOF | 6 | 0.19499 | 0.011325 | 5.8083 | TRUE | 0.19436 | 0.090949 | 0.29778 | pmol/µL [µM]
NIST hTAG | Cer d18:1/16:0 | Single Point | 18a | QQQ | 18 | 0.20108 | 0.010269 | 5.1067 | TRUE | 0.19436 | 0.090949 | 0.29778 | pmol/µL [µM]
NIST hTAG | Cer d18:1/16:0 | Single Point | 18b | QQQ | 18 | 0.19798 | 0.011473 | 5.7953 | TRUE | 0.19436 | 0.090949 | 0.29778 | pmol/µL [µM]
NIST hTAG | Cer d18:1/16:0 | Single Point | 19 | QQQ | 6 | 0.13771 | 0.019273 | 13.9951 | TRUE | 0.19436 | 0.090949 | 0.29778 | pmol/µL [µM]
NIST hTAG | Cer d18:1/16:0 | Single Point | 22 | QQQ | 18 | 0.18696 | 0.0056846 | 3.0405 | TRUE | 0.19436 | 0.090949 | 0.29778 | pmol/µL [µM]
NIST hTAG | Cer d18:1/16:0 | Single Point | 23 | QQQ | 6 | 0.20588 | 0.0060345 | 2.9311 | TRUE | 0.19436 | 0.090949 | 0.29778 | pmol/µL [µM]
NIST hTAG | Cer d18:1/16:0 | Single Point | 24 | QQQ | 18 | 0.23493 | 0.013064 | 5.5608 | TRUE | 0.19436 | 0.090949 | 0.29778 | pmol/µL [µM]
NIST hTAG | Cer d18:1/16:0 | Single Point | 27 | QQQ | 18 | 0.17487 | 0.0089034 | 5.0913 | TRUE | 0.19436 | 0.090949 | 0.29778 | pmol/µL [µM]
NIST hTAG | Cer d18:1/16:0 | Single Point | 28 | QQQ | 18 | 0.23083 | 0.047326 | 20.5024 | TRUE | 0.19436 | 0.090949 | 0.29778 | pmol/µL [µM]
NIST hTAG | Cer d18:1/16:0 | Single Point | 29a | TOF | 18 | 0.17853 | 0.0033188 | 1.8589 | TRUE | 0.19436 | 0.090949 | 0.29778 | pmol/µL [µM]
NIST hTAG | Cer d18:1/16:0 | Single Point | 3 | QQQ | 6 | 0.21632 | 0.017531 | 8.1043 | TRUE | 0.19436 | 0.090949 | 0.29778 | pmol/µL [µM]
NIST hTAG | Cer d18:1/16:0 | Single Point | 30 | QQQ | 18 | 0.2484 | 0.035122 | 14.1395 | TRUE | 0.19436 | 0.090949 | 0.29778 | pmol/µL [µM]
NIST hTAG | Cer d18:1/16:0 | Single Point | 31 | QQQ | 18 | 0.2312 | 0.01392 | 6.0209 | TRUE | 0.19436 | 0.090949 | 0.29778 | pmol/µL [µM]
NIST hTAG | Cer d18:1/16:0 | Single Point | 32 | TOF | 18 | 0.23695 | 0.057238 | 24.1563 | TRUE | 0.19436 | 0.090949 | 0.29778 | pmol/µL [µM]
NIST hTAG | Cer d18:1/16:0 | Single Point | 35 | QQQ | 18 | 0.18934 | 0.012223 | 6.4557 | TRUE | 0.19436 | 0.090949 | 0.29778 | pmol/µL [µM]
NIST hTAG | Cer d18:1/16:0 | Single Point | 36 | QQQ | 18 | 0.19968 | 0.0083987 | 4.2061 | TRUE | 0.19436 | 0.090949 | 0.29778 | pmol/µL [µM]
NIST hTAG | Cer d18:1/16:0 | Single Point | 37 | QQQ | 18 | 0.043473 | 0.0032061 | 7.3749 | FALSE | 0.19436 | 0.090949 | 0.29778 | pmol/µL [µM]
NIST hTAG | Cer d18:1/16:0 | Single Point | 38 | QQQ | 18 | 0.052052 | 0.0082398 | 15.8299 | FALSE | 0.19436 | 0.090949 | 0.29778 | pmol/µL [µM]
NIST hTAG | Cer d18:1/16:0 | Single Point | 4 | QQQ | 18 | 0.18903 | 0.004676 | 2.4736 | TRUE | 0.19436 | 0.090949 | 0.29778 | pmol/µL [µM]
NIST hTAG | Cer d18:1/16:0 | Single Point | 7 | Orbitrap | 18 | 0.29986 | 0.019493 | 6.5008 | FALSE | 0.19436 | 0.090949 | 0.29778 | pmol/µL [µM]
NIST hTAG | Cer d18:1/18:0 | Calibration Curve | 02a | QQQ | 6 | 0.051565 | 0.0010241 | 1.986 | TRUE | 0.047868 | 0.022792 | 0.072944 | pmol/µL [µM]
NIST hTAG | Cer d18:1/18:0 | Calibration Curve | 02b | QQQ | 18 | 0.052654 | 0.0017748 | 3.3706 | TRUE | 0.047868 | 0.022792 | 0.072944 | pmol/µL [µM]
NIST hTAG | Cer d18:1/18:0 | Calibration Curve | 10a | QQQ | 6 | 0.030397 | 0.006541 | 21.5186 | TRUE | 0.047868 | 0.022792 | 0.072944 | pmol/µL [µM]
NIST hTAG | Cer d18:1/18:0 | Calibration Curve | 10b | QQQ | 6 | 0.027336 | 0.0041714 | 15.2599 | TRUE | 0.047868 | 0.022792 | 0.072944 | pmol/µL [µM]
NIST hTAG | Cer d18:1/18:0 | Calibration Curve | 12 | QQQ | 18 | 0.061742 | 0.0045777 | 7.4142 | TRUE | 0.047868 | 0.022792 | 0.072944 | pmol/µL [µM]
NIST hTAG | Cer d18:1/18:0 | Calibration Curve | 14 | QQQ | 18 | 0.050697 | 0.0019476 | 3.8416 | TRUE | 0.047868 | 0.022792 | 0.072944 | pmol/µL [µM]
NIST hTAG | Cer d18:1/18:0 | Calibration Curve | 15 | QQQ | 18 | 0.049388 | 0.0052777 | 10.6861 | TRUE | 0.047868 | 0.022792 | 0.072944 | pmol/µL [µM]
NIST hTAG | Cer d18:1/18:0 | Calibration Curve | 16 | Orbitrap | 18 | 0.046201 | 0.0066929 | 14.4866 | TRUE | 0.047868 | 0.022792 | 0.072944 | pmol/µL [µM]
NIST hTAG | Cer d18:1/18:0 | Calibration Curve | 17a | TOF | 6 | 0.069748 | 0.0091487 | 13.1167 | TRUE | 0.047868 | 0.022792 | 0.072944 | pmol/µL [µM]
NIST hTAG | Cer d18:1/18:0 | Calibration Curve | 18a | QQQ | 18 | 0.050552 | 0.0024982 | 4.9418 | TRUE | 0.047868 | 0.022792 | 0.072944 | pmol/µL [µM]
NIST hTAG | Cer d18:1/18:0 | Calibration Curve | 18b | QQQ | 18 | 0.047603 | 0.0019485 | 4.0933 | TRUE | 0.047868 | 0.022792 | 0.072944 | pmol/µL [µM]
NIST hTAG | Cer d18:1/18:0 | Calibration Curve | 19 | QQQ | 6 | 0.044599 | 0.010557 | 23.671 | TRUE | 0.047868 | 0.022792 | 0.072944 | pmol/µL [µM]
NIST hTAG | Cer d18:1/18:0 | Calibration Curve | 22 | QQQ | 18 | 0.049463 | 0.0018116 | 3.6627 | TRUE | 0.047868 | 0.022792 | 0.072944 | pmol/µL [µM]
NIST hTAG | Cer d18:1/18:0 | Calibration Curve | 23 | QQQ | 6 | 0.058456 | 0.0030815 | 5.2715 | TRUE | 0.047868 | 0.022792 | 0.072944 | pmol/µL [µM]
NIST hTAG | Cer d18:1/18:0 | Calibration Curve | 24 | QQQ | 18 | 0.054421 | 0.0029424 | 5.4067 | TRUE | 0.047868 | 0.022792 | 0.072944 | pmol/µL [µM]
NIST hTAG | Cer d18:1/18:0 | Calibration Curve | 27 | QQQ | 18 | 0.028659 | 0.0027726 | 9.6744 | TRUE | 0.047868 | 0.022792 | 0.072944 | pmol/µL [µM]
NIST hTAG | Cer d18:1/18:0 | Calibration Curve | 28 | QQQ | 18 | 0.04222 | 0.0014013 | 3.319 | TRUE | 0.047868 | 0.022792 | 0.072944 | pmol/µL [µM]
NIST hTAG | Cer d18:1/18:0 | Calibration Curve | 29a | TOF | 18 | 0.046077 | 0.0017479 | 3.7935 | TRUE | 0.047868 | 0.022792 | 0.072944 | pmol/µL [µM]
NIST hTAG | Cer d18:1/18:0 | Calibration Curve | 3 | QQQ | 6 | 0.051339 | 0.0028126 | 5.4784 | TRUE | 0.047868 | 0.022792 | 0.072944 | pmol/µL [µM]
NIST hTAG | Cer d18:1/18:0 | Calibration Curve | 30 | QQQ | 18 | 0.045167 | 0.0044071 | 9.7571 | TRUE | 0.047868 | 0.022792 | 0.072944 | pmol/µL [µM]
NIST hTAG | Cer d18:1/18:0 | Calibration Curve | 31 | QQQ | 18 | 0.045807 | 0.0028914 | 6.312 | TRUE | 0.047868 | 0.022792 | 0.072944 | pmol/µL [µM]
NIST hTAG | Cer d18:1/18:0 | Calibration Curve | 32 | TOF | 18 | 0.078296 | 0.013862 | 17.7046 | FALSE | 0.047868 | 0.022792 | 0.072944 | pmol/µL [µM]
NIST hTAG | Cer d18:1/18:0 | Calibration Curve | 35 | QQQ | 18 | 0.033295 | 0.0029065 | 8.7295 | TRUE | 0.047868 | 0.022792 | 0.072944 | pmol/µL [µM]
NIST hTAG | Cer d18:1/18:0 | Calibration Curve | 36 | QQQ | 18 | 0.050737 | 0.0019883 | 3.9188 | TRUE | 0.047868 | 0.022792 | 0.072944 | pmol/µL [µM]
NIST hTAG | Cer d18:1/18:0 | Calibration Curve | 37 | QQQ | 18 | 0.052181 | 0.0044214 | 8.4731 | TRUE | 0.047868 | 0.022792 | 0.072944 | pmol/µL [µM]
NIST hTAG | Cer d18:1/18:0 | Calibration Curve | 38 | QQQ | 18 | 0.016535 | 0.0032478 | 19.6419 | FALSE | 0.047868 | 0.022792 | 0.072944 | pmol/µL [µM]
NIST hTAG | Cer d18:1/18:0 | Calibration Curve | 4 | QQQ | 18 | 0.053603 | 0.001065 | 1.9868 | TRUE | 0.047868 | 0.022792 | 0.072944 | pmol/µL [µM]
NIST hTAG | Cer d18:1/18:0 | Calibration Curve | 7 | Orbitrap | 18 | 0.051568 | 0.0022078 | 4.2814 | TRUE | 0.047868 | 0.022792 | 0.072944 | pmol/µL [µM]
NIST hTAG | Cer d18:1/18:0 | Single Point | 02a | QQQ | 6 | 0.033351 | 0.00064875 | 1.9452 | TRUE | 0.033291 | 0.014538 | 0.052044 | pmol/µL [µM]
NIST hTAG | Cer d18:1/18:0 | Single Point | 02b | QQQ | 18 | 0.03444 | 0.0011465 | 3.3289 | TRUE | 0.033291 | 0.014538 | 0.052044 | pmol/µL [µM]
NIST hTAG | Cer d18:1/18:0 | Single Point | 10a | QQQ | 6 | 0.031622 | 0.0049768 | 15.7385 | TRUE | 0.033291 | 0.014538 | 0.052044 | pmol/µL [µM]
NIST hTAG | Cer d18:1/18:0 | Single Point | 10b | QQQ | 6 | 0.026803 | 0.0031854 | 11.8843 | TRUE | 0.033291 | 0.014538 | 0.052044 | pmol/µL [µM]
NIST hTAG | Cer d18:1/18:0 | Single Point | 12 | QQQ | 18 | 0.037538 | 0.0027771 | 7.3982 | TRUE | 0.033291 | 0.014538 | 0.052044 | pmol/µL [µM]
NIST hTAG | Cer d18:1/18:0 | Single Point | 14 | QQQ | 18 | 0.033558 | 0.0012525 | 3.7324 | TRUE | 0.033291 | 0.014538 | 0.052044 | pmol/µL [µM]
NIST hTAG | Cer d18:1/18:0 | Single Point | 15 | QQQ | 18 | 0.032456 | 0.0034673 | 10.6829 | TRUE | 0.033291 | 0.014538 | 0.052044 | pmol/µL [µM]
NIST hTAG | Cer d18:1/18:0 | Single Point | 16 | Orbitrap | 18 | 0.031086 | 0.0048865 | 15.7189 | TRUE | 0.033291 | 0.014538 | 0.052044 | pmol/µL [µM]
NIST hTAG | Cer d18:1/18:0 | Single Point | 17a | TOF | 6 | 0.043813 | 0.0056432 | 12.8801 | TRUE | 0.033291 | 0.014538 | 0.052044 | pmol/µL [µM]
NIST hTAG | Cer d18:1/18:0 | Single Point | 18a | QQQ | 18 | 0.036729 | 0.0017585 | 4.7876 | TRUE | 0.033291 | 0.014538 | 0.052044 | pmol/µL [µM]
NIST hTAG | Cer d18:1/18:0 | Single Point | 18b | QQQ | 18 | 0.027257 | 0.0010574 | 3.8792 | TRUE | 0.033291 | 0.014538 | 0.052044 | pmol/µL [µM]
NIST hTAG | Cer d18:1/18:0 | Single Point | 19 | QQQ | 6 | 0.027055 | 0.0054639 | 20.1952 | TRUE | 0.033291 | 0.014538 | 0.052044 | pmol/µL [µM]
NIST hTAG | Cer d18:1/18:0 | Single Point | 22 | QQQ | 18 | 0.031441 | 0.0011417 | 3.6312 | TRUE | 0.033291 | 0.014538 | 0.052044 | pmol/µL [µM]
NIST hTAG | Cer d18:1/18:0 | Single Point | 23 | QQQ | 6 | 0.038634 | 0.0021851 | 5.6558 | TRUE | 0.033291 | 0.014538 | 0.052044 | pmol/µL [µM]
NIST hTAG | Cer d18:1/18:0 | Single Point | 24 | QQQ | 18 | 0.043016 | 0.0022518 | 5.2347 | TRUE | 0.033291 | 0.014538 | 0.052044 | pmol/µL [µM]
NIST hTAG | Cer d18:1/18:0 | Single Point | 27 | QQQ | 18 | 0.028852 | 0.0020153 | 6.985 | TRUE | 0.033291 | 0.014538 | 0.052044 | pmol/µL [µM]
NIST hTAG | Cer d18:1/18:0 | Single Point | 28 | QQQ | 18 | 0.033597 | 0.00089474 | 2.6632 | TRUE | 0.033291 | 0.014538 | 0.052044 | pmol/µL [µM]
NIST hTAG | Cer d18:1/18:0 | Single Point | 29a | TOF | 18 | 0.029442 | 0.0010996 | 3.7348 | TRUE | 0.033291 | 0.014538 | 0.052044 | pmol/µL [µM]
NIST hTAG | Cer d18:1/18:0 | Single Point | 3 | QQQ | 6 | 0.034144 | 0.001854 | 5.4299 | TRUE | 0.033291 | 0.014538 | 0.052044 | pmol/µL [µM]
NIST hTAG | Cer d18:1/18:0 | Single Point | 30 | QQQ | 18 | 0.029706 | 0.0029803 | 10.0325 | TRUE | 0.033291 | 0.014538 | 0.052044 | pmol/µL [µM]
NIST hTAG | Cer d18:1/18:0 | Single Point | 31 | QQQ | 18 | 0.036779 | 0.0021756 | 5.9154 | TRUE | 0.033291 | 0.014538 | 0.052044 | pmol/µL [µM]
NIST hTAG | Cer d18:1/18:0 | Single Point | 32 | TOF | 18 | 0.068122 | 0.011968 | 17.5685 | FALSE | 0.033291 | 0.014538 | 0.052044 | pmol/µL [µM]
NIST hTAG | Cer d18:1/18:0 | Single Point | 35 | QQQ | 18 | 0.027624 | 0.0018011 | 6.5198 | TRUE | 0.033291 | 0.014538 | 0.052044 | pmol/µL [µM]
NIST hTAG | Cer d18:1/18:0 | Single Point | 36 | QQQ | 18 | 0.03253 | 0.0012432 | 3.8217 | TRUE | 0.033291 | 0.014538 | 0.052044 | pmol/µL [µM]
NIST hTAG | Cer d18:1/18:0 | Single Point | 37 | QQQ | 18 | 0.029047 | 0.0023756 | 8.1786 | TRUE | 0.033291 | 0.014538 | 0.052044 | pmol/µL [µM]
NIST hTAG | Cer d18:1/18:0 | Single Point | 38 | QQQ | 18 | 0.0079125 | 0.0014999 | 18.9561 | FALSE | 0.033291 | 0.014538 | 0.052044 | pmol/µL [µM]
NIST hTAG | Cer d18:1/18:0 | Single Point | 4 | QQQ | 18 | 0.030475 | 0.00060436 | 1.9832 | TRUE | 0.033291 | 0.014538 | 0.052044 | pmol/µL [µM]
NIST hTAG | Cer d18:1/18:0 | Single Point | 7 | Orbitrap | 18 | 0.035112 | 0.0014019 | 3.9925 | TRUE | 0.033291 | 0.014538 | 0.052044 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:0 | Calibration Curve | 02a | QQQ | 6 | 2.1199 | 0.043922 | 2.0719 | TRUE | 2.0918 | -0.61446 | 4.798 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:0 | Calibration Curve | 02b | QQQ | 18 | 2.0611 | 0.030389 | 1.4744 | TRUE | 2.0918 | -0.61446 | 4.798 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:0 | Calibration Curve | 10a | QQQ | 6 | 1.5978 | 0.093498 | 5.8518 | TRUE | 2.0918 | -0.61446 | 4.798 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:0 | Calibration Curve | 10b | QQQ | 6 | 1.3441 | 0.061902 | 4.6056 | TRUE | 2.0918 | -0.61446 | 4.798 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:0 | Calibration Curve | 12 | QQQ | 18 | 2.0848 | 0.14357 | 6.8864 | TRUE | 2.0918 | -0.61446 | 4.798 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:0 | Calibration Curve | 14 | QQQ | 18 | 2.1501 | 0.086937 | 4.0433 | TRUE | 2.0918 | -0.61446 | 4.798 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:0 | Calibration Curve | 15 | QQQ | 18 | 2.3593 | 0.23029 | 9.7606 | TRUE | 2.0918 | -0.61446 | 4.798 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:0 | Calibration Curve | 16 | Orbitrap | 18 | 1.8175 | 0.20557 | 11.3104 | TRUE | 2.0918 | -0.61446 | 4.798 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:0 | Calibration Curve | 17a | TOF | 6 | 1.9604 | 0.032446 | 1.6551 | TRUE | 2.0918 | -0.61446 | 4.798 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:0 | Calibration Curve | 18a | QQQ | 18 | 1.8437 | 0.046701 | 2.533 | TRUE | 2.0918 | -0.61446 | 4.798 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:0 | Calibration Curve | 18b | QQQ | 18 | 2.1705 | 0.14764 | 6.8023 | TRUE | 2.0918 | -0.61446 | 4.798 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:0 | Calibration Curve | 19 | QQQ | 6 | 1.8335 | 0.18876 | 10.2952 | TRUE | 2.0918 | -0.61446 | 4.798 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:0 | Calibration Curve | 22 | QQQ | 18 | 1.8104 | 0.053789 | 2.9712 | TRUE | 2.0918 | -0.61446 | 4.798 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:0 | Calibration Curve | 23 | QQQ | 6 | 2.0486 | 0.045235 | 2.2081 | TRUE | 2.0918 | -0.61446 | 4.798 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:0 | Calibration Curve | 24 | QQQ | 18 | 1.9531 | 0.094489 | 4.8379 | TRUE | 2.0918 | -0.61446 | 4.798 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:0 | Calibration Curve | 27 | QQQ | 18 | 1.541 | 0.088138 | 5.7196 | TRUE | 2.0918 | -0.61446 | 4.798 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:0 | Calibration Curve | 28 | QQQ | 18 | 2.3795 | 0.071487 | 3.0042 | TRUE | 2.0918 | -0.61446 | 4.798 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:0 | Calibration Curve | 29a | TOF | 18 | 1.7936 | 0.042085 | 2.3464 | TRUE | 2.0918 | -0.61446 | 4.798 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:0 | Calibration Curve | 3 | QQQ | 6 | 2.0248 | 0.099961 | 4.9368 | TRUE | 2.0918 | -0.61446 | 4.798 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:0 | Calibration Curve | 30 | QQQ | 18 | 1.5264 | 0.050626 | 3.3167 | TRUE | 2.0918 | -0.61446 | 4.798 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:0 | Calibration Curve | 31 | QQQ | 18 | 1.9739 | 0.081589 | 4.1335 | TRUE | 2.0918 | -0.61446 | 4.798 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:0 | Calibration Curve | 32 | TOF | 18 | 1.5059 | 0.18045 | 11.9827 | TRUE | 2.0918 | -0.61446 | 4.798 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:0 | Calibration Curve | 35 | QQQ | 18 | 1.5094 | 0.080656 | 5.3435 | TRUE | 2.0918 | -0.61446 | 4.798 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:0 | Calibration Curve | 36 | QQQ | 18 | 1.9837 | 0.060708 | 3.0604 | TRUE | 2.0918 | -0.61446 | 4.798 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:0 | Calibration Curve | 37 | QQQ | 18 | 8.745 | 0.69019 | 7.8924 | FALSE | 2.0918 | -0.61446 | 4.798 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:0 | Calibration Curve | 38 | QQQ | 18 | 0.53919 | 0.11362 | 21.0731 | TRUE | 2.0918 | -0.61446 | 4.798 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:0 | Calibration Curve | 4 | QQQ | 18 | 1.9101 | 0.032456 | 1.6991 | TRUE | 2.0918 | -0.61446 | 4.798 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:0 | Calibration Curve | 7 | Orbitrap | 18 | 1.9823 | 0.10522 | 5.3082 | TRUE | 2.0918 | -0.61446 | 4.798 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:0 | Single Point | 02a | QQQ | 6 | 2.0482 | 0.042097 | 2.0553 | TRUE | 1.907 | 0.89883 | 2.9152 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:0 | Single Point | 02b | QQQ | 18 | 1.9946 | 0.029122 | 1.4601 | TRUE | 1.907 | 0.89883 | 2.9152 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:0 | Single Point | 10a | QQQ | 6 | 2.0243 | 0.11113 | 5.4898 | TRUE | 1.907 | 0.89883 | 2.9152 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:0 | Single Point | 10b | QQQ | 6 | 1.7548 | 0.076653 | 4.3683 | TRUE | 1.907 | 0.89883 | 2.9152 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:0 | Single Point | 12 | QQQ | 18 | 1.6232 | 0.11188 | 6.8926 | TRUE | 1.907 | 0.89883 | 2.9152 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:0 | Single Point | 14 | QQQ | 18 | 2.2282 | 0.089099 | 3.9987 | TRUE | 1.907 | 0.89883 | 2.9152 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:0 | Single Point | 15 | QQQ | 18 | 1.7812 | 0.17246 | 9.6827 | TRUE | 1.907 | 0.89883 | 2.9152 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:0 | Single Point | 16 | Orbitrap | 18 | 2.0139 | 0.22964 | 11.4027 | TRUE | 1.907 | 0.89883 | 2.9152 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:0 | Single Point | 17a | TOF | 6 | 2.4739 | 0.041537 | 1.679 | TRUE | 1.907 | 0.89883 | 2.9152 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:0 | Single Point | 18a | QQQ | 18 | 2.0492 | 0.051859 | 2.5307 | TRUE | 1.907 | 0.89883 | 2.9152 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:0 | Single Point | 18b | QQQ | 18 | 2.0999 | 0.13985 | 6.66 | TRUE | 1.907 | 0.89883 | 2.9152 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:0 | Single Point | 19 | QQQ | 6 | 2.0415 | 0.20904 | 10.2394 | TRUE | 1.907 | 0.89883 | 2.9152 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:0 | Single Point | 22 | QQQ | 18 | 1.9379 | 0.057675 | 2.9762 | TRUE | 1.907 | 0.89883 | 2.9152 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:0 | Single Point | 23 | QQQ | 6 | 2.3688 | 0.053415 | 2.2549 | TRUE | 1.907 | 0.89883 | 2.9152 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:0 | Single Point | 24 | QQQ | 18 | 0.63652 | 0.030621 | 4.8107 | FALSE | 1.907 | 0.89883 | 2.9152 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:0 | Single Point | 27 | QQQ | 18 | 1.7593 | 0.09637 | 5.4776 | TRUE | 1.907 | 0.89883 | 2.9152 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:0 | Single Point | 28 | QQQ | 18 | 2.0623 | 0.060129 | 2.9157 | TRUE | 1.907 | 0.89883 | 2.9152 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:0 | Single Point | 29a | TOF | 18 | 1.9254 | 0.044878 | 2.3309 | TRUE | 1.907 | 0.89883 | 2.9152 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:0 | Single Point | 3 | QQQ | 6 | 2.303 | 0.11324 | 4.9169 | TRUE | 1.907 | 0.89883 | 2.9152 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:0 | Single Point | 30 | QQQ | 18 | 1.5133 | 0.050431 | 3.3324 | TRUE | 1.907 | 0.89883 | 2.9152 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:0 | Single Point | 31 | QQQ | 18 | 3.1234 | 0.1259 | 4.0309 | FALSE | 1.907 | 0.89883 | 2.9152 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:0 | Single Point | 32 | TOF | 18 | 1.7898 | 0.21542 | 12.036 | TRUE | 1.907 | 0.89883 | 2.9152 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:0 | Single Point | 35 | QQQ | 18 | 1.5918 | 0.080167 | 5.0363 | TRUE | 1.907 | 0.89883 | 2.9152 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:0 | Single Point | 36 | QQQ | 18 | 2.0892 | 0.063683 | 3.0482 | TRUE | 1.907 | 0.89883 | 2.9152 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:0 | Single Point | 37 | QQQ | 18 | 1.4668 | 0.11576 | 7.8917 | TRUE | 1.907 | 0.89883 | 2.9152 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:0 | Single Point | 38 | QQQ | 18 | 0.46032 | 0.094122 | 20.4469 | FALSE | 1.907 | 0.89883 | 2.9152 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:0 | Single Point | 4 | QQQ | 18 | 2.1353 | 0.03633 | 1.7014 | TRUE | 1.907 | 0.89883 | 2.9152 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:0 | Single Point | 7 | Orbitrap | 18 | 2.0999 | 0.10923 | 5.2017 | TRUE | 1.907 | 0.89883 | 2.9152 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:1 | Calibration Curve | 02a | QQQ | 6 | 0.68162 | 0.016729 | 2.4543 | TRUE | 0.62275 | 0.32957 | 0.91592 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:1 | Calibration Curve | 02b | QQQ | 18 | 0.68818 | 0.011795 | 1.714 | TRUE | 0.62275 | 0.32957 | 0.91592 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:1 | Calibration Curve | 10a | QQQ | 6 | 0.95395 | 0.040779 | 4.2747 | FALSE | 0.62275 | 0.32957 | 0.91592 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:1 | Calibration Curve | 10b | QQQ | 6 | 0.3684 | 0.029328 | 7.9608 | TRUE | 0.62275 | 0.32957 | 0.91592 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:1 | Calibration Curve | 12 | QQQ | 18 | 0.67717 | 0.049715 | 7.3416 | TRUE | 0.62275 | 0.32957 | 0.91592 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:1 | Calibration Curve | 14 | QQQ | 18 | 0.66817 | 0.033493 | 5.0127 | TRUE | 0.62275 | 0.32957 | 0.91592 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:1 | Calibration Curve | 15 | QQQ | 18 | 0.80598 | 0.077376 | 9.6003 | TRUE | 0.62275 | 0.32957 | 0.91592 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:1 | Calibration Curve | 16 | Orbitrap | 18 | 0.52018 | 0.24145 | 46.4168 | TRUE | 0.62275 | 0.32957 | 0.91592 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:1 | Calibration Curve | 17a | TOF | 6 | 0.83079 | 0.036823 | 4.4322 | TRUE | 0.62275 | 0.32957 | 0.91592 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:1 | Calibration Curve | 18a | QQQ | 18 | 0.61605 | 0.012413 | 2.0149 | TRUE | 0.62275 | 0.32957 | 0.91592 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:1 | Calibration Curve | 18b | QQQ | 18 | 0.66762 | 0.025497 | 3.8192 | TRUE | 0.62275 | 0.32957 | 0.91592 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:1 | Calibration Curve | 19 | QQQ | 6 | 0.61096 | 0.041908 | 6.8594 | TRUE | 0.62275 | 0.32957 | 0.91592 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:1 | Calibration Curve | 22 | QQQ | 18 | 0.60965 | 0.014757 | 2.4206 | TRUE | 0.62275 | 0.32957 | 0.91592 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:1 | Calibration Curve | 23 | QQQ | 6 | 0.66435 | 0.033343 | 5.0188 | TRUE | 0.62275 | 0.32957 | 0.91592 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:1 | Calibration Curve | 24 | QQQ | 18 | 0.63105 | 0.029041 | 4.6021 | TRUE | 0.62275 | 0.32957 | 0.91592 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:1 | Calibration Curve | 27 | QQQ | 18 | 0.4489 | 0.023266 | 5.1829 | TRUE | 0.62275 | 0.32957 | 0.91592 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:1 | Calibration Curve | 28 | QQQ | 18 | 0.71996 | 0.021872 | 3.0379 | TRUE | 0.62275 | 0.32957 | 0.91592 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:1 | Calibration Curve | 29a | TOF | 18 | 0.58404 | 0.012354 | 2.1153 | TRUE | 0.62275 | 0.32957 | 0.91592 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:1 | Calibration Curve | 3 | QQQ | 6 | 0.69209 | 0.037788 | 5.4599 | TRUE | 0.62275 | 0.32957 | 0.91592 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:1 | Calibration Curve | 30 | QQQ | 18 | 0.58033 | 0.037251 | 6.419 | TRUE | 0.62275 | 0.32957 | 0.91592 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:1 | Calibration Curve | 31 | QQQ | 18 | 0.61773 | 0.026407 | 4.2749 | TRUE | 0.62275 | 0.32957 | 0.91592 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:1 | Calibration Curve | 32 | TOF | 18 | 0.73085 | 0.112 | 15.325 | TRUE | 0.62275 | 0.32957 | 0.91592 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:1 | Calibration Curve | 35 | QQQ | 18 | 0.45873 | 0.026617 | 5.8022 | TRUE | 0.62275 | 0.32957 | 0.91592 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:1 | Calibration Curve | 36 | QQQ | 18 | 0.54422 | 0.020111 | 3.6953 | TRUE | 0.62275 | 0.32957 | 0.91592 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:1 | Calibration Curve | 37 | QQQ | 18 | 0.56425 | 0.038014 | 6.7371 | TRUE | 0.62275 | 0.32957 | 0.91592 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:1 | Calibration Curve | 38 | QQQ | 18 | 0.17531 | 0.03957 | 22.5709 | FALSE | 0.62275 | 0.32957 | 0.91592 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:1 | Calibration Curve | 4 | QQQ | 18 | 0.63136 | 0.014446 | 2.2881 | TRUE | 0.62275 | 0.32957 | 0.91592 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:1 | Calibration Curve | 7 | Orbitrap | 18 | 0.695 | 0.036411 | 5.239 | TRUE | 0.62275 | 0.32957 | 0.91592 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:1 | Single Point | 02a | QQQ | 6 | 0.62114 | 0.01503 | 2.4198 | TRUE | 0.59856 | 0.22404 | 0.97309 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:1 | Single Point | 02b | QQQ | 18 | 0.63307 | 0.010709 | 1.6916 | TRUE | 0.59856 | 0.22404 | 0.97309 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:1 | Single Point | 10a | QQQ | 6 | 0.75096 | 0.029896 | 3.981 | TRUE | 0.59856 | 0.22404 | 0.97309 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:1 | Single Point | 10b | QQQ | 6 | 0.5195 | 0.035143 | 6.7649 | TRUE | 0.59856 | 0.22404 | 0.97309 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:1 | Single Point | 12 | QQQ | 18 | 0.49478 | 0.036462 | 7.3694 | TRUE | 0.59856 | 0.22404 | 0.97309 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:1 | Single Point | 14 | QQQ | 18 | 0.67246 | 0.032858 | 4.8863 | TRUE | 0.59856 | 0.22404 | 0.97309 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:1 | Single Point | 15 | QQQ | 18 | 0.57145 | 0.053976 | 9.4453 | TRUE | 0.59856 | 0.22404 | 0.97309 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:1 | Single Point | 16 | Orbitrap | 18 | 0.29214 | 0.14551 | 49.8086 | TRUE | 0.59856 | 0.22404 | 0.97309 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:1 | Single Point | 17a | TOF | 6 | 1.0875 | 0.049966 | 4.5945 | FALSE | 0.59856 | 0.22404 | 0.97309 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:1 | Single Point | 18a | QQQ | 18 | 0.70896 | 0.014272 | 2.0131 | TRUE | 0.59856 | 0.22404 | 0.97309 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:1 | Single Point | 18b | QQQ | 18 | 0.69739 | 0.025333 | 3.6325 | TRUE | 0.59856 | 0.22404 | 0.97309 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:1 | Single Point | 19 | QQQ | 6 | 0.55271 | 0.038279 | 6.9258 | TRUE | 0.59856 | 0.22404 | 0.97309 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:1 | Single Point | 22 | QQQ | 18 | 0.60633 | 0.014744 | 2.4317 | TRUE | 0.59856 | 0.22404 | 0.97309 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:1 | Single Point | 23 | QQQ | 6 | 0.69678 | 0.037541 | 5.3878 | TRUE | 0.59856 | 0.22404 | 0.97309 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:1 | Single Point | 24 | QQQ | 18 | 0.18409 | 0.0083658 | 4.5443 | FALSE | 0.59856 | 0.22404 | 0.97309 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:1 | Single Point | 27 | QQQ | 18 | 0.52549 | 0.024085 | 4.5834 | TRUE | 0.59856 | 0.22404 | 0.97309 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:1 | Single Point | 28 | QQQ | 18 | 0.61686 | 0.017631 | 2.8581 | TRUE | 0.59856 | 0.22404 | 0.97309 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:1 | Single Point | 29a | TOF | 18 | 0.60111 | 0.012633 | 2.1016 | TRUE | 0.59856 | 0.22404 | 0.97309 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:1 | Single Point | 3 | QQQ | 6 | 0.68565 | 0.037844 | 5.5194 | TRUE | 0.59856 | 0.22404 | 0.97309 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:1 | Single Point | 30 | QQQ | 18 | 0.59696 | 0.038937 | 6.5226 | TRUE | 0.59856 | 0.22404 | 0.97309 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:1 | Single Point | 31 | QQQ | 18 | 0.74201 | 0.030481 | 4.1079 | TRUE | 0.59856 | 0.22404 | 0.97309 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:1 | Single Point | 32 | TOF | 18 | 0.83465 | 0.12862 | 15.4101 | TRUE | 0.59856 | 0.22404 | 0.97309 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:1 | Single Point | 35 | QQQ | 18 | 0.50325 | 0.024278 | 4.8242 | TRUE | 0.59856 | 0.22404 | 0.97309 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:1 | Single Point | 36 | QQQ | 18 | 0.624 | 0.016177 | 2.5925 | TRUE | 0.59856 | 0.22404 | 0.97309 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:1 | Single Point | 37 | QQQ | 18 | 0.41297 | 0.027599 | 6.6829 | TRUE | 0.59856 | 0.22404 | 0.97309 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:1 | Single Point | 38 | QQQ | 18 | 0.14472 | 0.02964 | 20.48 | FALSE | 0.59856 | 0.22404 | 0.97309 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:1 | Single Point | 4 | QQQ | 18 | 0.65143 | 0.015078 | 2.3146 | TRUE | 0.59856 | 0.22404 | 0.97309 | pmol/µL [µM]
NIST hTAG | Cer d18:1/24:1 | Single Point | 7 | Orbitrap | 18 | 0.73137 | 0.03645 | 4.9838 | TRUE | 0.59856 | 0.22404 | 0.97309 | pmol/µL [µM]
NIST SRM | Cer d18:1/16:0 | Calibration Curve | 02a | QQQ | 6 | 0.12798 | 0.0043396 | 3.3909 | TRUE | 0.11688 | 0.056201 | 0.17756 | pmol/µL [µM]
NIST SRM | Cer d18:1/16:0 | Calibration Curve | 02b | QQQ | 18 | 0.13269 | 0.0046796 | 3.5266 | TRUE | 0.11688 | 0.056201 | 0.17756 | pmol/µL [µM]
NIST SRM | Cer d18:1/16:0 | Calibration Curve | 10a | QQQ | 6 | 0.072839 | 0.00768 | 10.5438 | TRUE | 0.11688 | 0.056201 | 0.17756 | pmol/µL [µM]
NIST SRM | Cer d18:1/16:0 | Calibration Curve | 10b | QQQ | 6 | 0.081914 | 0.013435 | 16.4019 | TRUE | 0.11688 | 0.056201 | 0.17756 | pmol/µL [µM]
NIST SRM | Cer d18:1/16:0 | Calibration Curve | 12 | QQQ | 18 | 0.12829 | 0.0053005 | 4.1318 | TRUE | 0.11688 | 0.056201 | 0.17756 | pmol/µL [µM]
NIST SRM | Cer d18:1/16:0 | Calibration Curve | 14 | QQQ | 18 | 0.11284 | 0.0045282 | 4.0129 | TRUE | 0.11688 | 0.056201 | 0.17756 | pmol/µL [µM]
NIST SRM | Cer d18:1/16:0 | Calibration Curve | 15 | QQQ | 18 | 0.12277 | 0.0048425 | 3.9444 | TRUE | 0.11688 | 0.056201 | 0.17756 | pmol/µL [µM]
NIST SRM | Cer d18:1/16:0 | Calibration Curve | 16 | Orbitrap | 18 | 0.12144 | 0.011273 | 9.2829 | TRUE | 0.11688 | 0.056201 | 0.17756 | pmol/µL [µM]
NIST SRM | Cer d18:1/16:0 | Calibration Curve | 17a | TOF | 6 | 0.14219 | 0.013946 | 9.8085 | TRUE | 0.11688 | 0.056201 | 0.17756 | pmol/µL [µM]
NIST SRM | Cer d18:1/16:0 | Calibration Curve | 18a | QQQ | 18 | 0.11742 | 0.0038268 | 3.259 | TRUE | 0.11688 | 0.056201 | 0.17756 | pmol/µL [µM]
NIST SRM | Cer d18:1/16:0 | Calibration Curve | 18b | QQQ | 18 | 0.12245 | 0.0067716 | 5.5301 | TRUE | 0.11688 | 0.056201 | 0.17756 | pmol/µL [µM]
NIST SRM | Cer d18:1/16:0 | Calibration Curve | 19 | QQQ | 6 | 0.13398 | 0.012646 | 9.4391 | TRUE | 0.11688 | 0.056201 | 0.17756 | pmol/µL [µM]
NIST SRM | Cer d18:1/16:0 | Calibration Curve | 22 | QQQ | 18 | 0.12439 | 0.0062989 | 5.0639 | TRUE | 0.11688 | 0.056201 | 0.17756 | pmol/µL [µM]
NIST SRM | Cer d18:1/16:0 | Calibration Curve | 23 | QQQ | 6 | 0.13061 | 0.0044732 | 3.4249 | TRUE | 0.11688 | 0.056201 | 0.17756 | pmol/µL [µM]
NIST SRM | Cer d18:1/16:0 | Calibration Curve | 24 | QQQ | 18 | 0.11766 | 0.0061872 | 5.2585 | TRUE | 0.11688 | 0.056201 | 0.17756 | pmol/µL [µM]
NIST SRM | Cer d18:1/16:0 | Calibration Curve | 27 | QQQ | 18 | 0.082707 | 0.0043215 | 5.2251 | TRUE | 0.11688 | 0.056201 | 0.17756 | pmol/µL [µM]
NIST SRM | Cer d18:1/16:0 | Calibration Curve | 28 | QQQ | 18 | 0.053415 | 0.024225 | 45.3521 | FALSE | 0.11688 | 0.056201 | 0.17756 | pmol/µL [µM]
NIST SRM | Cer d18:1/16:0 | Calibration Curve | 29a | TOF | 18 | 0.12474 | 0.0030701 | 2.4611 | TRUE | 0.11688 | 0.056201 | 0.17756 | pmol/µL [µM]
NIST SRM | Cer d18:1/16:0 | Calibration Curve | 3 | QQQ | 6 | 0.10884 | 0.0083619 | 7.6829 | TRUE | 0.11688 | 0.056201 | 0.17756 | pmol/µL [µM]
NIST SRM | Cer d18:1/16:0 | Calibration Curve | 30 | QQQ | 18 | 0.15561 | 0.0044348 | 2.8499 | TRUE | 0.11688 | 0.056201 | 0.17756 | pmol/µL [µM]
NIST SRM | Cer d18:1/16:0 | Calibration Curve | 31 | QQQ | 18 | 0.11575 | 0.0034385 | 2.9706 | TRUE | 0.11688 | 0.056201 | 0.17756 | pmol/µL [µM]
NIST SRM | Cer d18:1/16:0 | Calibration Curve | 32 | TOF | 18 | 0.1689 | 0.028121 | 16.6495 | TRUE | 0.11688 | 0.056201 | 0.17756 | pmol/µL [µM]
NIST SRM | Cer d18:1/16:0 | Calibration Curve | 35 | QQQ | 18 | 0.093069 | 0.0048964 | 5.261 | TRUE | 0.11688 | 0.056201 | 0.17756 | pmol/µL [µM]
NIST SRM | Cer d18:1/16:0 | Calibration Curve | 36 | QQQ | 18 | 0.11838 | 0.0033255 | 2.8092 | TRUE | 0.11688 | 0.056201 | 0.17756 | pmol/µL [µM]
NIST SRM | Cer d18:1/16:0 | Calibration Curve | 37 | QQQ | 18 | 0.136 | 0.01277 | 9.3893 | TRUE | 0.11688 | 0.056201 | 0.17756 | pmol/µL [µM]
NIST SRM | Cer d18:1/16:0 | Calibration Curve | 38 | QQQ | 18 | 0.031935 | 0.0046892 | 14.6838 | FALSE | 0.11688 | 0.056201 | 0.17756 | pmol/µL [µM]
NIST SRM | Cer d18:1/16:0 | Calibration Curve | 4 | QQQ | 18 | 0.13248 | 0.0049987 | 3.7732 | TRUE | 0.11688 | 0.056201 | 0.17756 | pmol/µL [µM]
NIST SRM | Cer d18:1/16:0 | Calibration Curve | 7 | Orbitrap | 18 | 0.16143 | 0.027025 | 16.7414 | TRUE | 0.11688 | 0.056201 | 0.17756 | pmol/µL [µM]
NIST SRM | Cer d18:1/16:0 | Single Point | 02a | QQQ | 6 | 0.17206 | 0.0057056 | 3.3161 | TRUE | 0.16545 | 0.077253 | 0.25364 | pmol/µL [µM]
NIST SRM | Cer d18:1/16:0 | Single Point | 02b | QQQ | 18 | 0.18544 | 0.0064773 | 3.4929 | TRUE | 0.16545 | 0.077253 | 0.25364 | pmol/µL [µM]
NIST SRM | Cer d18:1/16:0 | Single Point | 10a | QQQ | 6 | 0.18077 | 0.013271 | 7.3413 | TRUE | 0.16545 | 0.077253 | 0.25364 | pmol/µL [µM]
NIST SRM | Cer d18:1/16:0 | Single Point | 10b | QQQ | 6 | 0.16833 | 0.020798 | 12.3555 | TRUE | 0.16545 | 0.077253 | 0.25364 | pmol/µL [µM]
NIST SRM | Cer d18:1/16:0 | Single Point | 12 | QQQ | 18 | 0.12415 | 0.0051334 | 4.1349 | TRUE | 0.16545 | 0.077253 | 0.25364 | pmol/µL [µM]
NIST SRM | Cer d18:1/16:0 | Single Point | 14 | QQQ | 18 | 0.18366 | 0.0072133 | 3.9275 | TRUE | 0.16545 | 0.077253 | 0.25364 | pmol/µL [µM]
NIST SRM | Cer d18:1/16:0 | Single Point | 15 | QQQ | 18 | 0.16809 | 0.0065605 | 3.9029 | TRUE | 0.16545 | 0.077253 | 0.25364 | pmol/µL [µM]
NIST SRM | Cer d18:1/16:0 | Single Point | 16 | Orbitrap | 18 | 0.18812 | 0.01783 | 9.4779 | TRUE | 0.16545 | 0.077253 | 0.25364 | pmol/µL [µM]
NIST SRM | Cer d18:1/16:0 | Single Point | 17a | TOF | 6 | 0.1587 | 0.0153 | 9.641 | TRUE | 0.16545 | 0.077253 | 0.25364 | pmol/µL [µM]
NIST SRM | Cer d18:1/16:0 | Single Point | 18a | QQQ | 18 | 0.16604 | 0.0052922 | 3.1873 | TRUE | 0.16545 | 0.077253 | 0.25364 | pmol/µL [µM]
NIST SRM | Cer d18:1/16:0 | Single Point | 18b | QQQ | 18 | 0.16757 | 0.0089136 | 5.3192 | TRUE | 0.16545 | 0.077253 | 0.25364 | pmol/µL [µM]
NIST SRM | Cer d18:1/16:0 | Single Point | 19 | QQQ | 6 | 0.12631 | 0.010913 | 8.6404 | TRUE | 0.16545 | 0.077253 | 0.25364 | pmol/µL [µM]
NIST SRM | Cer d18:1/16:0 | Single Point | 22 | QQQ | 18 | 0.16849 | 0.0084801 | 5.033 | TRUE | 0.16545 | 0.077253 | 0.25364 | pmol/µL [µM]
NIST SRM | Cer d18:1/16:0 | Single Point | 23 | QQQ | 6 | 0.17577 | 0.0061912 | 3.5224 | TRUE | 0.16545 | 0.077253 | 0.25364 | pmol/µL [µM]
NIST SRM | Cer d18:1/16:0 | Single Point | 24 | QQQ | 18 | 0.18996 | 0.0097704 | 5.1435 | TRUE | 0.16545 | 0.077253 | 0.25364 | pmol/µL [µM]
NIST SRM | Cer d18:1/16:0 | Single Point | 27 | QQQ | 18 | 0.15521 | 0.0062219 | 4.0086 | TRUE | 0.16545 | 0.077253 | 0.25364 | pmol/µL [µM]
NIST SRM | Cer d18:1/16:0 | Single Point | 28 | QQQ | 18 | 0.19043 | 0.031744 | 16.6693 | TRUE | 0.16545 | 0.077253 | 0.25364 | pmol/µL [µM]
NIST SRM | Cer d18:1/16:0 | Single Point | 29a | TOF | 18 | 0.16221 | 0.0039229 | 2.4184 | TRUE | 0.16545 | 0.077253 | 0.25364 | pmol/µL [µM]
NIST SRM | Cer d18:1/16:0 | Single Point | 3 | QQQ | 6 | 0.18102 | 0.013527 | 7.4723 | TRUE | 0.16545 | 0.077253 | 0.25364 | pmol/µL [µM]
NIST SRM | Cer d18:1/16:0 | Single Point | 30 | QQQ | 18 | 0.1998 | 0.0056732 | 2.8395 | TRUE | 0.16545 | 0.077253 | 0.25364 | pmol/µL [µM]
NIST SRM | Cer d18:1/16:0 | Single Point | 31 | QQQ | 18 | 0.18816 | 0.0053998 | 2.8698 | TRUE | 0.16545 | 0.077253 | 0.25364 | pmol/µL [µM]
NIST SRM | Cer d18:1/16:0 | Single Point | 32 | TOF | 18 | 0.23805 | 0.039478 | 16.5841 | TRUE | 0.16545 | 0.077253 | 0.25364 | pmol/µL [µM]
NIST SRM | Cer d18:1/16:0 | Single Point | 35 | QQQ | 18 | 0.16607 | 0.0067081 | 4.0394 | TRUE | 0.16545 | 0.077253 | 0.25364 | pmol/µL [µM]
NIST SRM | Cer d18:1/16:0 | Single Point | 36 | QQQ | 18 | 0.16652 | 0.0046432 | 2.7883 | TRUE | 0.16545 | 0.077253 | 0.25364 | pmol/µL [µM]
NIST SRM | Cer d18:1/16:0 | Single Point | 37 | QQQ | 18 | 0.037474 | 0.0034145 | 9.1117 | FALSE | 0.16545 | 0.077253 | 0.25364 | pmol/µL [µM]
NIST SRM | Cer d18:1/16:0 | Single Point | 38 | QQQ | 18 | 0.029898 | 0.0045926 | 15.361 | FALSE | 0.16545 | 0.077253 | 0.25364 | pmol/µL [µM]
NIST SRM | Cer d18:1/16:0 | Single Point | 4 | QQQ | 18 | 0.16358 | 0.0061544 | 3.7624 | TRUE | 0.16545 | 0.077253 | 0.25364 | pmol/µL [µM]
NIST SRM | Cer d18:1/16:0 | Single Point | 7 | Orbitrap | 18 | 0.23069 | 0.036432 | 15.7931 | TRUE | 0.16545 | 0.077253 | 0.25364 | pmol/µL [µM]
NIST SRM | Cer d18:1/18:0 | Calibration Curve | 02a | QQQ | 6 | 0.042906 | 0.0014073 | 3.2799 | TRUE | 0.04026 | 0.017494 | 0.063026 | pmol/µL [µM]
NIST SRM | Cer d18:1/18:0 | Calibration Curve | 02b | QQQ | 18 | 0.043729 | 0.0014025 | 3.2072 | TRUE | 0.04026 | 0.017494 | 0.063026 | pmol/µL [µM]
NIST SRM | Cer d18:1/18:0 | Calibration Curve | 10a | QQQ | 6 | 0.025505 | 0.0041136 | 16.1289 | TRUE | 0.04026 | 0.017494 | 0.063026 | pmol/µL [µM]
NIST SRM | Cer d18:1/18:0 | Calibration Curve | 10b | QQQ | 6 | 0.0244 | 0.0038827 | 15.9131 | TRUE | 0.04026 | 0.017494 | 0.063026 | pmol/µL [µM]
NIST SRM | Cer d18:1/18:0 | Calibration Curve | 12 | QQQ | 18 | 0.052427 | 0.0024943 | 4.7576 | TRUE | 0.04026 | 0.017494 | 0.063026 | pmol/µL [µM]
NIST SRM | Cer d18:1/18:0 | Calibration Curve | 14 | QQQ | 18 | 0.040989 | 0.0019533 | 4.7655 | TRUE | 0.04026 | 0.017494 | 0.063026 | pmol/µL [µM]
NIST SRM | Cer d18:1/18:0 | Calibration Curve | 15 | QQQ | 18 | 0.043557 | 0.0034775 | 7.9839 | TRUE | 0.04026 | 0.017494 | 0.063026 | pmol/µL [µM]
NIST SRM | Cer d18:1/18:0 | Calibration Curve | 16 | Orbitrap | 18 | 0.04061 | 0.0047326 | 11.6538 | TRUE | 0.04026 | 0.017494 | 0.063026 | pmol/µL [µM]
NIST SRM | Cer d18:1/18:0 | Calibration Curve | 17a | TOF | 6 | 0.058691 | 0.0068304 | 11.638 | TRUE | 0.04026 | 0.017494 | 0.063026 | pmol/µL [µM]
NIST SRM | Cer d18:1/18:0 | Calibration Curve | 18a | QQQ | 18 | 0.040443 | 0.0015508 | 3.8346 | TRUE | 0.04026 | 0.017494 | 0.063026 | pmol/µL [µM]
NIST SRM | Cer d18:1/18:0 | Calibration Curve | 18b | QQQ | 18 | 0.043304 | 0.0023289 | 5.378 | TRUE | 0.04026 | 0.017494 | 0.063026 | pmol/µL [µM]
NIST SRM | Cer d18:1/18:0 | Calibration Curve | 19 | QQQ | 6 | 0.037268 | 0.0019309 | 5.1811 | TRUE | 0.04026 | 0.017494 | 0.063026 | pmol/µL [µM]
NIST SRM | Cer d18:1/18:0 | Calibration Curve | 22 | QQQ | 18 | 0.04402 | 0.0024825 | 5.6395 | TRUE | 0.04026 | 0.017494 | 0.063026 | pmol/µL [µM]
NIST SRM | Cer d18:1/18:0 | Calibration Curve | 23 | QQQ | 6 | 0.04725 | 0.0018844 | 3.9882 | TRUE | 0.04026 | 0.017494 | 0.063026 | pmol/µL [µM]
NIST SRM | Cer d18:1/18:0 | Calibration Curve | 24 | QQQ | 18 | 0.042417 | 0.0025644 | 6.0456 | TRUE | 0.04026 | 0.017494 | 0.063026 | pmol/µL [µM]
NIST SRM | Cer d18:1/18:0 | Calibration Curve | 27 | QQQ | 18 | 0.024044 | 0.0014939 | 6.2133 | TRUE | 0.04026 | 0.017494 | 0.063026 | pmol/µL [µM]
NIST SRM | Cer d18:1/18:0 | Calibration Curve | 28 | QQQ | 18 | 0.033452 | 0.001179 | 3.5243 | TRUE | 0.04026 | 0.017494 | 0.063026 | pmol/µL [µM]
NIST SRM | Cer d18:1/18:0 | Calibration Curve | 29a | TOF | 18 | 0.040729 | 0.00095418 | 2.3427 | TRUE | 0.04026 | 0.017494 | 0.063026 | pmol/µL [µM]
NIST SRM | Cer d18:1/18:0 | Calibration Curve | 3 | QQQ | 6 | 0.043079 | 0.0032872 | 7.6306 | TRUE | 0.04026 | 0.017494 | 0.063026 | pmol/µL [µM]
NIST SRM | Cer d18:1/18:0 | Calibration Curve | 30 | QQQ | 18 | 0.04218 | 0.0020668 | 4.9 | TRUE | 0.04026 | 0.017494 | 0.063026 | pmol/µL [µM]
NIST SRM | Cer d18:1/18:0 | Calibration Curve | 31 | QQQ | 18 | 0.036885 | 0.0012023 | 3.2597 | TRUE | 0.04026 | 0.017494 | 0.063026 | pmol/µL [µM]
NIST SRM | Cer d18:1/18:0 | Calibration Curve | 32 | TOF | 18 | 0.071378 | 0.015917 | 22.2994 | FALSE | 0.04026 | 0.017494 | 0.063026 | pmol/µL [µM]
NIST SRM | Cer d18:1/18:0 | Calibration Curve | 35 | QQQ | 18 | 0.028074 | 0.0030691 | 10.9322 | TRUE | 0.04026 | 0.017494 | 0.063026 | pmol/µL [µM]
NIST SRM | Cer d18:1/18:0 | Calibration Curve | 36 | QQQ | 18 | 0.041555 | 0.0020626 | 4.9635 | TRUE | 0.04026 | 0.017494 | 0.063026 | pmol/µL [µM]
NIST SRM | Cer d18:1/18:0 | Calibration Curve | 37 | QQQ | 18 | 0.043619 | 0.0031291 | 7.1737 | TRUE | 0.04026 | 0.017494 | 0.063026 | pmol/µL [µM]
NIST SRM | Cer d18:1/18:0 | Calibration Curve | 38 | QQQ | 18 | 0.0096669 | 0.0015247 | 15.7727 | FALSE | 0.04026 | 0.017494 | 0.063026 | pmol/µL [µM]
NIST SRM | Cer d18:1/18:0 | Calibration Curve | 4 | QQQ | 18 | 0.045407 | 0.0015472 | 3.4074 | TRUE | 0.04026 | 0.017494 | 0.063026 | pmol/µL [µM]
NIST SRM | Cer d18:1/18:0 | Calibration Curve | 7 | Orbitrap | 18 | 0.039703 | 0.0060072 | 15.1303 | TRUE | 0.04026 | 0.017494 | 0.063026 | pmol/µL [µM]
NIST SRM | Cer d18:1/18:0 | Single Point | 02a | QQQ | 6 | 0.027866 | 0.0008915 | 3.1993 | TRUE | 0.028325 | 0.011385 | 0.045264 | pmol/µL [µM]
NIST SRM | Cer d18:1/18:0 | Single Point | 02b | QQQ | 18 | 0.028675 | 0.000906 | 3.1596 | TRUE | 0.028325 | 0.011385 | 0.045264 | pmol/µL [µM]
NIST SRM | Cer d18:1/18:0 | Single Point | 10a | QQQ | 6 | 0.027899 | 0.0031299 | 11.2185 | TRUE | 0.028325 | 0.011385 | 0.045264 | pmol/µL [µM]
NIST SRM | Cer d18:1/18:0 | Single Point | 10b | QQQ | 6 | 0.024561 | 0.002965 | 12.0717 | TRUE | 0.028325 | 0.011385 | 0.045264 | pmol/µL [µM]
NIST SRM | Cer d18:1/18:0 | Single Point | 12 | QQQ | 18 | 0.031887 | 0.0015132 | 4.7456 | TRUE | 0.028325 | 0.011385 | 0.045264 | pmol/µL [µM]
NIST SRM | Cer d18:1/18:0 | Single Point | 14 | QQQ | 18 | 0.027315 | 0.0012562 | 4.599 | TRUE | 0.028325 | 0.011385 | 0.045264 | pmol/µL [µM]
NIST SRM | Cer d18:1/18:0 | Single Point | 15 | QQQ | 18 | 0.028625 | 0.0022846 | 7.9812 | TRUE | 0.028325 | 0.011385 | 0.045264 | pmol/µL [µM]
NIST SRM | Cer d18:1/18:0 | Single Point | 16 | Orbitrap | 18 | 0.027005 | 0.0034552 | 12.795 | TRUE | 0.028325 | 0.011385 | 0.045264 | pmol/µL [µM]
NIST SRM | Cer d18:1/18:0 | Single Point | 17a | TOF | 6 | 0.036993 | 0.0042132 | 11.3893 | TRUE | 0.028325 | 0.011385 | 0.045264 | pmol/µL [µM]
NIST SRM | Cer d18:1/18:0 | Single Point | 18a | QQQ | 18 | 0.029613 | 0.0010916 | 3.6862 | TRUE | 0.028325 | 0.011385 | 0.045264 | pmol/µL [µM]
NIST SRM | Cer d18:1/18:0 | Single Point | 18b | QQQ | 18 | 0.024924 | 0.0012638 | 5.0704 | TRUE | 0.028325 | 0.011385 | 0.045264 | pmol/µL [µM]
NIST SRM | Cer d18:1/18:0 | Single Point | 19 | QQQ | 6 | 0.023261 | 0.00099936 | 4.2962 | TRUE | 0.028325 | 0.011385 | 0.045264 | pmol/µL [µM]
NIST SRM | Cer d18:1/18:0 | Single Point | 22 | QQQ | 18 | 0.028012 | 0.0015645 | 5.585 | TRUE | 0.028325 | 0.011385 | 0.045264 | pmol/µL [µM]
NIST SRM | Cer d18:1/18:0 | Single Point | 23 | QQQ | 6 | 0.030688 | 0.0013362 | 4.3542 | TRUE | 0.028325 | 0.011385 | 0.045264 | pmol/µL [µM]
NIST SRM | Cer d18:1/18:0 | Single Point | 24 | QQQ | 18 | 0.033829 | 0.0019625 | 5.8011 | TRUE | 0.028325 | 0.011385 | 0.045264 | pmol/µL [µM]
NIST SRM | Cer d18:1/18:0 | Single Point | 27 | QQQ | 18 | 0.025498 | 0.0010859 | 4.2588 | TRUE | 0.028325 | 0.011385 | 0.045264 | pmol/µL [µM]
NIST SRM | Cer d18:1/18:0 | Single Point | 28 | QQQ | 18 | 0.027998 | 0.00075278 | 2.6887 | TRUE | 0.028325 | 0.011385 | 0.045264 | pmol/µL [µM]
NIST SRM | Cer d18:1/18:0 | Single Point | 29a | TOF | 18 | 0.026078 | 0.00060027 | 2.3018 | TRUE | 0.028325 | 0.011385 | 0.045264 | pmol/µL [µM]
NIST SRM | Cer d18:1/18:0 | Single Point | 3 | QQQ | 6 | 0.0287 | 0.0021669 | 7.5502 | TRUE | 0.028325 | 0.011385 | 0.045264 | pmol/µL [µM]
NIST SRM | Cer d18:1/18:0 | Single Point | 30 | QQQ | 18 | 0.027686 | 0.0013977 | 5.0484 | TRUE | 0.028325 | 0.011385 | 0.045264 | pmol/µL [µM]
NIST SRM | Cer d18:1/18:0 | Single Point | 31 | QQQ | 18 | 0.030065 | 0.00090471 | 3.0092 | TRUE | 0.028325 | 0.011385 | 0.045264 | pmol/µL [µM]
NIST SRM | Cer d18:1/18:0 | Single Point | 32 | TOF | 18 | 0.062149 | 0.013742 | 22.1115 | FALSE | 0.028325 | 0.011385 | 0.045264 | pmol/µL [µM]
NIST SRM | Cer d18:1/18:0 | Single Point | 35 | QQQ | 18 | 0.024389 | 0.0019018 | 7.7979 | TRUE | 0.028325 | 0.011385 | 0.045264 | pmol/µL [µM]
NIST SRM | Cer d18:1/18:0 | Single Point | 36 | QQQ | 18 | 0.026789 | 0.0012897 | 4.8142 | TRUE | 0.028325 | 0.011385 | 0.045264 | pmol/µL [µM]
NIST SRM | Cer d18:1/18:0 | Single Point | 37 | QQQ | 18 | 0.024446 | 0.0016813 | 6.8774 | TRUE | 0.028325 | 0.011385 | 0.045264 | pmol/µL [µM]
NIST SRM | Cer d18:1/18:0 | Single Point | 38 | QQQ | 18 | 0.0047407 | 0.00070416 | 14.8536 | FALSE | 0.028325 | 0.011385 | 0.045264 | pmol/µL [µM]
NIST SRM | Cer d18:1/18:0 | Single Point | 4 | QQQ | 18 | 0.025824 | 0.000878 | 3.4 | TRUE | 0.028325 | 0.011385 | 0.045264 | pmol/µL [µM]
NIST SRM | Cer d18:1/18:0 | Single Point | 7 | Orbitrap | 18 | 0.027579 | 0.0038143 | 13.8308 | TRUE | 0.028325 | 0.011385 | 0.045264 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:0 | Calibration Curve | 02a | QQQ | 6 | 1.3179 | 0.023406 | 1.776 | TRUE | 1.3516 | -0.45899 | 3.1623 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:0 | Calibration Curve | 02b | QQQ | 18 | 1.3106 | 0.033149 | 2.5294 | TRUE | 1.3516 | -0.45899 | 3.1623 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:0 | Calibration Curve | 10a | QQQ | 6 | 0.95171 | 0.18332 | 19.2625 | TRUE | 1.3516 | -0.45899 | 3.1623 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:0 | Calibration Curve | 10b | QQQ | 6 | 0.88768 | 0.01709 | 1.9253 | TRUE | 1.3516 | -0.45899 | 3.1623 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:0 | Calibration Curve | 12 | QQQ | 18 | 1.335 | 0.050554 | 3.7869 | TRUE | 1.3516 | -0.45899 | 3.1623 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:0 | Calibration Curve | 14 | QQQ | 18 | 1.3658 | 0.049098 | 3.5949 | TRUE | 1.3516 | -0.45899 | 3.1623 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:0 | Calibration Curve | 15 | QQQ | 18 | 1.5913 | 0.14404 | 9.052 | TRUE | 1.3516 | -0.45899 | 3.1623 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:0 | Calibration Curve | 16 | Orbitrap | 18 | 1.1966 | 0.11301 | 9.4444 | TRUE | 1.3516 | -0.45899 | 3.1623 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:0 | Calibration Curve | 17a | TOF | 6 | 1.2324 | 0.03602 | 2.9228 | TRUE | 1.3516 | -0.45899 | 3.1623 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:0 | Calibration Curve | 18a | QQQ | 18 | 1.1474 | 0.023274 | 2.0284 | TRUE | 1.3516 | -0.45899 | 3.1623 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:0 | Calibration Curve | 18b | QQQ | 18 | 1.3269 | 0.03045 | 2.2948 | TRUE | 1.3516 | -0.45899 | 3.1623 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:0 | Calibration Curve | 19 | QQQ | 6 | 1.0817 | 0.072139 | 6.6688 | TRUE | 1.3516 | -0.45899 | 3.1623 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:0 | Calibration Curve | 22 | QQQ | 18 | 1.2786 | 0.068092 | 5.3256 | TRUE | 1.3516 | -0.45899 | 3.1623 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:0 | Calibration Curve | 23 | QQQ | 6 | 1.2669 | 0.03856 | 3.0437 | TRUE | 1.3516 | -0.45899 | 3.1623 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:0 | Calibration Curve | 24 | QQQ | 18 | 1.1666 | 0.047042 | 4.0325 | TRUE | 1.3516 | -0.45899 | 3.1623 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:0 | Calibration Curve | 27 | QQQ | 18 | 1.0189 | 0.040209 | 3.9463 | TRUE | 1.3516 | -0.45899 | 3.1623 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:0 | Calibration Curve | 28 | QQQ | 18 | 1.5219 | 0.040664 | 2.672 | TRUE | 1.3516 | -0.45899 | 3.1623 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:0 | Calibration Curve | 29a | TOF | 18 | 1.2257 | 0.03575 | 2.9168 | TRUE | 1.3516 | -0.45899 | 3.1623 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:0 | Calibration Curve | 3 | QQQ | 6 | 1.3324 | 0.077069 | 5.784 | TRUE | 1.3516 | -0.45899 | 3.1623 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:0 | Calibration Curve | 30 | QQQ | 18 | 1.0948 | 0.050212 | 4.5863 | TRUE | 1.3516 | -0.45899 | 3.1623 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:0 | Calibration Curve | 31 | QQQ | 18 | 1.2353 | 0.021608 | 1.7492 | TRUE | 1.3516 | -0.45899 | 3.1623 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:0 | Calibration Curve | 32 | TOF | 18 | 1.1867 | 0.17885 | 15.0708 | TRUE | 1.3516 | -0.45899 | 3.1623 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:0 | Calibration Curve | 35 | QQQ | 18 | 1.0303 | 0.021523 | 2.0889 | TRUE | 1.3516 | -0.45899 | 3.1623 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:0 | Calibration Curve | 36 | QQQ | 18 | 1.2427 | 0.032528 | 2.6177 | TRUE | 1.3516 | -0.45899 | 3.1623 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:0 | Calibration Curve | 37 | QQQ | 18 | 5.8107 | 0.28235 | 4.8591 | FALSE | 1.3516 | -0.45899 | 3.1623 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:0 | Calibration Curve | 38 | QQQ | 18 | 0.26173 | 0.045327 | 17.3186 | TRUE | 1.3516 | -0.45899 | 3.1623 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:0 | Calibration Curve | 4 | QQQ | 18 | 1.2683 | 0.046004 | 3.6271 | TRUE | 1.3516 | -0.45899 | 3.1623 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:0 | Calibration Curve | 7 | Orbitrap | 18 | 1.1594 | 0.16889 | 14.5662 | TRUE | 1.3516 | -0.45899 | 3.1623 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:0 | Single Point | 02a | QQQ | 6 | 1.2796 | 0.022433 | 1.7532 | TRUE | 1.2376 | 0.5907 | 1.8845 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:0 | Single Point | 02b | QQQ | 18 | 1.2753 | 0.031767 | 2.4909 | TRUE | 1.2376 | 0.5907 | 1.8845 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:0 | Single Point | 10a | QQQ | 6 | 1.2564 | 0.2179 | 17.343 | TRUE | 1.2376 | 0.5907 | 1.8845 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:0 | Single Point | 10b | QQQ | 6 | 1.1896 | 0.021163 | 1.7789 | TRUE | 1.2376 | 0.5907 | 1.8845 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:0 | Single Point | 12 | QQQ | 18 | 1.0389 | 0.039396 | 3.7922 | TRUE | 1.2376 | 0.5907 | 1.8845 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:0 | Single Point | 14 | QQQ | 18 | 1.4243 | 0.050319 | 3.5328 | TRUE | 1.2376 | 0.5907 | 1.8845 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:0 | Single Point | 15 | QQQ | 18 | 1.206 | 0.10788 | 8.9452 | TRUE | 1.2376 | 0.5907 | 1.8845 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:0 | Single Point | 16 | Orbitrap | 18 | 1.3203 | 0.12624 | 9.562 | TRUE | 1.2376 | 0.5907 | 1.8845 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:0 | Single Point | 17a | TOF | 6 | 1.5419 | 0.046112 | 2.9906 | TRUE | 1.2376 | 0.5907 | 1.8845 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:0 | Single Point | 18a | QQQ | 18 | 1.276 | 0.025845 | 2.0254 | TRUE | 1.2376 | 0.5907 | 1.8845 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:0 | Single Point | 18b | QQQ | 18 | 1.3008 | 0.028843 | 2.2173 | TRUE | 1.2376 | 0.5907 | 1.8845 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:0 | Single Point | 19 | QQQ | 6 | 1.209 | 0.07989 | 6.6078 | TRUE | 1.2376 | 0.5907 | 1.8845 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:0 | Single Point | 22 | QQQ | 18 | 1.3677 | 0.073011 | 5.3383 | TRUE | 1.2376 | 0.5907 | 1.8845 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:0 | Single Point | 23 | QQQ | 6 | 1.4458 | 0.045533 | 3.1494 | TRUE | 1.2376 | 0.5907 | 1.8845 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:0 | Single Point | 24 | QQQ | 18 | 0.38162 | 0.015245 | 3.9947 | FALSE | 1.2376 | 0.5907 | 1.8845 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:0 | Single Point | 27 | QQQ | 18 | 1.1885 | 0.043965 | 3.6992 | TRUE | 1.2376 | 0.5907 | 1.8845 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:0 | Single Point | 28 | QQQ | 18 | 1.3409 | 0.034204 | 2.5509 | TRUE | 1.2376 | 0.5907 | 1.8845 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:0 | Single Point | 29a | TOF | 18 | 1.3198 | 0.038122 | 2.8886 | TRUE | 1.2376 | 0.5907 | 1.8845 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:0 | Single Point | 3 | QQQ | 6 | 1.5187 | 0.087306 | 5.7486 | TRUE | 1.2376 | 0.5907 | 1.8845 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:0 | Single Point | 30 | QQQ | 18 | 1.0834 | 0.050019 | 4.6167 | TRUE | 1.2376 | 0.5907 | 1.8845 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:0 | Single Point | 31 | QQQ | 18 | 1.9838 | 0.033344 | 1.6808 | FALSE | 1.2376 | 0.5907 | 1.8845 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:0 | Single Point | 32 | TOF | 18 | 1.4087 | 0.21351 | 15.156 | TRUE | 1.2376 | 0.5907 | 1.8845 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:0 | Single Point | 35 | QQQ | 18 | 1.1156 | 0.021393 | 1.9176 | TRUE | 1.2376 | 0.5907 | 1.8845 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:0 | Single Point | 36 | QQQ | 18 | 1.3118 | 0.034122 | 2.6011 | TRUE | 1.2376 | 0.5907 | 1.8845 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:0 | Single Point | 37 | QQQ | 18 | 0.97468 | 0.047354 | 4.8584 | TRUE | 1.2376 | 0.5907 | 1.8845 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:0 | Single Point | 38 | QQQ | 18 | 0.23048 | 0.037547 | 16.2908 | FALSE | 1.2376 | 0.5907 | 1.8845 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:0 | Single Point | 4 | QQQ | 18 | 1.4169 | 0.051495 | 3.6344 | TRUE | 1.2376 | 0.5907 | 1.8845 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:0 | Single Point | 7 | Orbitrap | 18 | 1.2457 | 0.17531 | 14.0735 | TRUE | 1.2376 | 0.5907 | 1.8845 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:1 | Calibration Curve | 02a | QQQ | 6 | 0.48451 | 0.0065083 | 1.3433 | TRUE | 0.42009 | 0.21344 | 0.62673 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:1 | Calibration Curve | 02b | QQQ | 18 | 0.46764 | 0.013338 | 2.8522 | TRUE | 0.42009 | 0.21344 | 0.62673 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:1 | Calibration Curve | 10a | QQQ | 6 | 0.35377 | 0.061332 | 17.337 | TRUE | 0.42009 | 0.21344 | 0.62673 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:1 | Calibration Curve | 10b | QQQ | 6 | 0.25769 | 0.0086023 | 3.3382 | TRUE | 0.42009 | 0.21344 | 0.62673 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:1 | Calibration Curve | 12 | QQQ | 18 | 0.47163 | 0.021172 | 4.4891 | TRUE | 0.42009 | 0.21344 | 0.62673 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:1 | Calibration Curve | 14 | QQQ | 18 | 0.45204 | 0.017358 | 3.8399 | TRUE | 0.42009 | 0.21344 | 0.62673 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:1 | Calibration Curve | 15 | QQQ | 18 | 0.5772 | 0.057135 | 9.8986 | TRUE | 0.42009 | 0.21344 | 0.62673 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:1 | Calibration Curve | 16 | Orbitrap | 18 | 0.46482 | 0.16979 | 36.5278 | TRUE | 0.42009 | 0.21344 | 0.62673 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:1 | Calibration Curve | 17a | TOF | 6 | 0.55484 | 0.011592 | 2.0893 | TRUE | 0.42009 | 0.21344 | 0.62673 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:1 | Calibration Curve | 18a | QQQ | 18 | 0.41132 | 0.0075821 | 1.8434 | TRUE | 0.42009 | 0.21344 | 0.62673 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:1 | Calibration Curve | 18b | QQQ | 18 | 0.4312 | 0.017925 | 4.157 | TRUE | 0.42009 | 0.21344 | 0.62673 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:1 | Calibration Curve | 19 | QQQ | 6 | 0.41927 | 0.039163 | 9.3407 | TRUE | 0.42009 | 0.21344 | 0.62673 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:1 | Calibration Curve | 22 | QQQ | 18 | 0.45103 | 0.022649 | 5.0217 | TRUE | 0.42009 | 0.21344 | 0.62673 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:1 | Calibration Curve | 23 | QQQ | 6 | 0.45969 | 0.014064 | 3.0595 | TRUE | 0.42009 | 0.21344 | 0.62673 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:1 | Calibration Curve | 24 | QQQ | 18 | 0.42019 | 0.022437 | 5.3398 | TRUE | 0.42009 | 0.21344 | 0.62673 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:1 | Calibration Curve | 27 | QQQ | 18 | 0.31102 | 0.010536 | 3.3875 | TRUE | 0.42009 | 0.21344 | 0.62673 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:1 | Calibration Curve | 28 | QQQ | 18 | 0.4872 | 0.012146 | 2.493 | TRUE | 0.42009 | 0.21344 | 0.62673 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:1 | Calibration Curve | 29a | TOF | 18 | 0.42624 | 0.010918 | 2.5615 | TRUE | 0.42009 | 0.21344 | 0.62673 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:1 | Calibration Curve | 3 | QQQ | 6 | 0.43124 | 0.017344 | 4.0218 | TRUE | 0.42009 | 0.21344 | 0.62673 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:1 | Calibration Curve | 30 | QQQ | 18 | 0.40937 | 0.015945 | 3.8951 | TRUE | 0.42009 | 0.21344 | 0.62673 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:1 | Calibration Curve | 31 | QQQ | 18 | 0.42425 | 0.0072395 | 1.7064 | TRUE | 0.42009 | 0.21344 | 0.62673 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:1 | Calibration Curve | 32 | TOF | 18 | 0.62481 | 0.089596 | 14.3398 | TRUE | 0.42009 | 0.21344 | 0.62673 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:1 | Calibration Curve | 35 | QQQ | 18 | 0.31316 | 0.012651 | 4.0397 | TRUE | 0.42009 | 0.21344 | 0.62673 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:1 | Calibration Curve | 36 | QQQ | 18 | 0.29447 | 0.0166 | 5.6372 | TRUE | 0.42009 | 0.21344 | 0.62673 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:1 | Calibration Curve | 37 | QQQ | 18 | 0.39209 | 0.02366 | 6.0344 | TRUE | 0.42009 | 0.21344 | 0.62673 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:1 | Calibration Curve | 38 | QQQ | 18 | 0.086224 | 0.016914 | 19.6164 | FALSE | 0.42009 | 0.21344 | 0.62673 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:1 | Calibration Curve | 4 | QQQ | 18 | 0.46166 | 0.017023 | 3.6873 | TRUE | 0.42009 | 0.21344 | 0.62673 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:1 | Calibration Curve | 7 | Orbitrap | 18 | 0.4239 | 0.060812 | 14.346 | TRUE | 0.42009 | 0.21344 | 0.62673 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:1 | Single Point | 02a | QQQ | 6 | 0.44406 | 0.0058473 | 1.3168 | TRUE | 0.41192 | 0.14877 | 0.67508 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:1 | Single Point | 02b | QQQ | 18 | 0.43286 | 0.012109 | 2.7975 | TRUE | 0.41192 | 0.14877 | 0.67508 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:1 | Single Point | 10a | QQQ | 6 | 0.31095 | 0.044964 | 14.4604 | TRUE | 0.41192 | 0.14877 | 0.67508 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:1 | Single Point | 10b | QQQ | 6 | 0.38683 | 0.010308 | 2.6647 | TRUE | 0.41192 | 0.14877 | 0.67508 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:1 | Single Point | 12 | QQQ | 18 | 0.34403 | 0.015528 | 4.5137 | TRUE | 0.41192 | 0.14877 | 0.67508 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:1 | Single Point | 14 | QQQ | 18 | 0.46043 | 0.017029 | 3.6985 | TRUE | 0.41192 | 0.14877 | 0.67508 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:1 | Single Point | 15 | QQQ | 18 | 0.41187 | 0.039856 | 9.6768 | TRUE | 0.41192 | 0.14877 | 0.67508 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:1 | Single Point | 16 | Orbitrap | 18 | 0.25878 | 0.10233 | 39.5411 | TRUE | 0.41192 | 0.14877 | 0.67508 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:1 | Single Point | 17a | TOF | 6 | 0.71305 | 0.01573 | 2.206 | FALSE | 0.41192 | 0.14877 | 0.67508 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:1 | Single Point | 18a | QQQ | 18 | 0.47356 | 0.0087179 | 1.8409 | TRUE | 0.41192 | 0.14877 | 0.67508 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:1 | Single Point | 18b | QQQ | 18 | 0.46251 | 0.017809 | 3.8506 | TRUE | 0.41192 | 0.14877 | 0.67508 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:1 | Single Point | 19 | QQQ | 6 | 0.37762 | 0.035772 | 9.473 | TRUE | 0.41192 | 0.14877 | 0.67508 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:1 | Single Point | 22 | QQQ | 18 | 0.44785 | 0.022629 | 5.0529 | TRUE | 0.41192 | 0.14877 | 0.67508 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:1 | Single Point | 23 | QQQ | 6 | 0.46635 | 0.015835 | 3.3955 | TRUE | 0.41192 | 0.14877 | 0.67508 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:1 | Single Point | 24 | QQQ | 18 | 0.12335 | 0.0064633 | 5.2397 | FALSE | 0.41192 | 0.14877 | 0.67508 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:1 | Single Point | 27 | QQQ | 18 | 0.38276 | 0.010907 | 2.8495 | TRUE | 0.41192 | 0.14877 | 0.67508 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:1 | Single Point | 28 | QQQ | 18 | 0.42924 | 0.0097907 | 2.281 | TRUE | 0.41192 | 0.14877 | 0.67508 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:1 | Single Point | 29a | TOF | 18 | 0.43975 | 0.011164 | 2.5387 | TRUE | 0.41192 | 0.14877 | 0.67508 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:1 | Single Point | 3 | QQQ | 6 | 0.42441 | 0.017369 | 4.0926 | TRUE | 0.41192 | 0.14877 | 0.67508 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:1 | Single Point | 30 | QQQ | 18 | 0.41826 | 0.016667 | 3.9848 | TRUE | 0.41192 | 0.14877 | 0.67508 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:1 | Single Point | 31 | QQQ | 18 | 0.51868 | 0.0083564 | 1.6111 | TRUE | 0.41192 | 0.14877 | 0.67508 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:1 | Single Point | 32 | TOF | 18 | 0.71288 | 0.10289 | 14.433 | FALSE | 0.41192 | 0.14877 | 0.67508 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:1 | Single Point | 35 | QQQ | 18 | 0.37047 | 0.011539 | 3.1147 | TRUE | 0.41192 | 0.14877 | 0.67508 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:1 | Single Point | 36 | QQQ | 18 | 0.4231 | 0.013353 | 3.1559 | TRUE | 0.41192 | 0.14877 | 0.67508 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:1 | Single Point | 37 | QQQ | 18 | 0.28798 | 0.017178 | 5.9648 | TRUE | 0.41192 | 0.14877 | 0.67508 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:1 | Single Point | 38 | QQQ | 18 | 0.077993 | 0.012669 | 16.2444 | FALSE | 0.41192 | 0.14877 | 0.67508 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:1 | Single Point | 4 | QQQ | 18 | 0.47431 | 0.017767 | 3.746 | TRUE | 0.41192 | 0.14877 | 0.67508 | pmol/µL [µM]
NIST SRM | Cer d18:1/24:1 | Single Point | 7 | Orbitrap | 18 | 0.45998 | 0.060878 | 13.235 | TRUE | 0.41192 | 0.14877 | 0.67508 | pmol/µL [µM]
NIST T1D | Cer d18:1/16:0 | Calibration Curve | 02a | QQQ | 6 | 0.12384 | 0.0027174 | 2.1942 | TRUE | 0.10547 | 0.041653 | 0.16929 | pmol/µL [µM]
NIST T1D | Cer d18:1/16:0 | Calibration Curve | 02b | QQQ | 18 | 0.1241 | 0.003896 | 3.1395 | TRUE | 0.10547 | 0.041653 | 0.16929 | pmol/µL [µM]
NIST T1D | Cer d18:1/16:0 | Calibration Curve | 10a | QQQ | 6 | 0.061799 | 0.0056898 | 9.207 | TRUE | 0.10547 | 0.041653 | 0.16929 | pmol/µL [µM]
NIST T1D | Cer d18:1/16:0 | Calibration Curve | 10b | QQQ | 6 | 0.067275 | 0.0074292 | 11.043 | TRUE | 0.10547 | 0.041653 | 0.16929 | pmol/µL [µM]
NIST T1D | Cer d18:1/16:0 | Calibration Curve | 12 | QQQ | 18 | 0.11712 | 0.011139 | 9.5106 | TRUE | 0.10547 | 0.041653 | 0.16929 | pmol/µL [µM]
NIST T1D | Cer d18:1/16:0 | Calibration Curve | 14 | QQQ | 18 | 0.1047 | 0.0044751 | 4.2741 | TRUE | 0.10547 | 0.041653 | 0.16929 | pmol/µL [µM]
NIST T1D | Cer d18:1/16:0 | Calibration Curve | 15 | QQQ | 18 | 0.11065 | 0.0036023 | 3.2556 | TRUE | 0.10547 | 0.041653 | 0.16929 | pmol/µL [µM]
NIST T1D | Cer d18:1/16:0 | Calibration Curve | 16 | Orbitrap | 18 | 0.10443 | 0.016128 | 15.444 | TRUE | 0.10547 | 0.041653 | 0.16929 | pmol/µL [µM]
NIST T1D | Cer d18:1/16:0 | Calibration Curve | 17a | TOF | 6 | 0.12147 | 0.0041636 | 3.4276 | TRUE | 0.10547 | 0.041653 | 0.16929 | pmol/µL [µM]
NIST T1D | Cer d18:1/16:0 | Calibration Curve | 18a | QQQ | 18 | 0.1054 | 0.0046759 | 4.4366 | TRUE | 0.10547 | 0.041653 | 0.16929 | pmol/µL [µM]
NIST T1D | Cer d18:1/16:0 | Calibration Curve | 18b | QQQ | 18 | 0.10852 | 0.0078886 | 7.2692 | TRUE | 0.10547 | 0.041653 | 0.16929 | pmol/µL [µM]
NIST T1D | Cer d18:1/16:0 | Calibration Curve | 19 | QQQ | 6 | 0.11865 | 0.018602 | 15.6785 | TRUE | 0.10547 | 0.041653 | 0.16929 | pmol/µL [µM]
NIST T1D | Cer d18:1/16:0 | Calibration Curve | 22 | QQQ | 18 | 0.10624 | 0.0035166 | 3.31 | TRUE | 0.10547 | 0.041653 | 0.16929 | pmol/µL [µM]
NIST T1D | Cer d18:1/16:0 | Calibration Curve | 23 | QQQ | 6 | 0.11674 | 0.00436 | 3.7347 | TRUE | 0.10547 | 0.041653 | 0.16929 | pmol/µL [µM]
NIST T1D | Cer d18:1/16:0 | Calibration Curve | 24 | QQQ | 18 | 0.10644 | 0.0043395 | 4.077 | TRUE | 0.10547 | 0.041653 | 0.16929 | pmol/µL [µM]
NIST T1D | Cer d18:1/16:0 | Calibration Curve | 27 | QQQ | 18 | 0.067563 | 0.0033443 | 4.95 | TRUE | 0.10547 | 0.041653 | 0.16929 | pmol/µL [µM]
NIST T1D | Cer d18:1/16:0 | Calibration Curve | 28 | QQQ | 18 | 0.035802 | 0.014123 | 39.4459 | FALSE | 0.10547 | 0.041653 | 0.16929 | pmol/µL [µM]
NIST T1D | Cer d18:1/16:0 | Calibration Curve | 29a | TOF | 18 | 0.10331 | 0.003224 | 3.1208 | TRUE | 0.10547 | 0.041653 | 0.16929 | pmol/µL [µM]
NIST T1D | Cer d18:1/16:0 | Calibration Curve | 3 | QQQ | 6 | 0.098863 | 0.0065712 | 6.6467 | TRUE | 0.10547 | 0.041653 | 0.16929 | pmol/µL [µM]
NIST T1D | Cer d18:1/16:0 | Calibration Curve | 30 | QQQ | 18 | 0.14242 | 0.0077558 | 5.4455 | TRUE | 0.10547 | 0.041653 | 0.16929 | pmol/µL [µM]
NIST T1D | Cer d18:1/16:0 | Calibration Curve | 31 | QQQ | 18 | 0.11187 | 0.0053238 | 4.759 | TRUE | 0.10547 | 0.041653 | 0.16929 | pmol/µL [µM]
NIST T1D | Cer d18:1/16:0 | Calibration Curve | 32 | TOF | 18 | 0.17081 | 0.02723 | 15.9415 | FALSE | 0.10547 | 0.041653 | 0.16929 | pmol/µL [µM]
NIST T1D | Cer d18:1/16:0 | Calibration Curve | 35 | QQQ | 18 | 0.078774 | 0.0073686 | 9.3541 | TRUE | 0.10547 | 0.041653 | 0.16929 | pmol/µL [µM]
NIST T1D | Cer d18:1/16:0 | Calibration Curve | 36 | QQQ | 18 | 0.10631 | 0.0023676 | 2.227 | TRUE | 0.10547 | 0.041653 | 0.16929 | pmol/µL [µM]
NIST T1D | Cer d18:1/16:0 | Calibration Curve | 37 | QQQ | 18 | 0.1244 | 0.021149 | 17.0007 | TRUE | 0.10547 | 0.041653 | 0.16929 | pmol/µL [µM]
NIST T1D | Cer d18:1/16:0 | Calibration Curve | 38 | QQQ | 18 | 0.029783 | 0.005735 | 19.2558 | FALSE | 0.10547 | 0.041653 | 0.16929 | pmol/µL [µM]
NIST T1D | Cer d18:1/16:0 | Calibration Curve | 4 | QQQ | 18 | 0.11969 | 0.0075602 | 6.3166 | TRUE | 0.10547 | 0.041653 | 0.16929 | pmol/µL [µM]
NIST T1D | Cer d18:1/16:0 | Calibration Curve | 7 | Orbitrap | 18 | 0.16623 | 0.016735 | 10.0679 | TRUE | 0.10547 | 0.041653 | 0.16929 | pmol/µL [µM]
NIST T1D | Cer d18:1/16:0 | Single Point | 02a | QQQ | 6 | 0.16662 | 0.0035728 | 2.1442 | TRUE | 0.15052 | 0.06249 | 0.23855 | pmol/µL [µM]
NIST T1D | Cer d18:1/16:0 | Single Point | 02b | QQQ | 18 | 0.17355 | 0.0053928 | 3.1074 | TRUE | 0.15052 | 0.06249 | 0.23855 | pmol/µL [µM]
NIST T1D | Cer d18:1/16:0 | Single Point | 10a | QQQ | 6 | 0.1617 | 0.009832 | 6.0805 | TRUE | 0.15052 | 0.06249 | 0.23855 | pmol/µL [µM]
NIST T1D | Cer d18:1/16:0 | Single Point | 10b | QQQ | 6 | 0.14567 | 0.0115 | 7.8949 | TRUE | 0.15052 | 0.06249 | 0.23855 | pmol/µL [µM]
NIST T1D | Cer d18:1/16:0 | Single Point | 12 | QQQ | 18 | 0.11333 | 0.010788 | 9.5186 | TRUE | 0.15052 | 0.06249 | 0.23855 | pmol/µL [µM]
NIST T1D | Cer d18:1/16:0 | Single Point | 14 | QQQ | 18 | 0.1707 | 0.0071287 | 4.1763 | TRUE | 0.15052 | 0.06249 | 0.23855 | pmol/µL [µM]
NIST T1D | Cer d18:1/16:0 | Single Point | 15 | QQQ | 18 | 0.15168 | 0.0048804 | 3.2176 | TRUE | 0.15052 | 0.06249 | 0.23855 | pmol/µL [µM]
NIST T1D | Cer d18:1/16:0 | Single Point | 16 | Orbitrap | 18 | 0.16122 | 0.02551 | 15.8225 | TRUE | 0.15052 | 0.06249 | 0.23855 | pmol/µL [µM]
NIST T1D | Cer d18:1/16:0 | Single Point | 17a | TOF | 6 | 0.13597 | 0.0045678 | 3.3593 | TRUE | 0.15052 | 0.06249 | 0.23855 | pmol/µL [µM]
NIST T1D | Cer d18:1/16:0 | Single Point | 18a | QQQ | 18 | 0.14941 | 0.0064665 | 4.328 | TRUE | 0.15052 | 0.06249 | 0.23855 | pmol/µL [µM]
NIST T1D | Cer d18:1/16:0 | Single Point | 18b | QQQ | 18 | 0.14924 | 0.010384 | 6.9579 | TRUE | 0.15052 | 0.06249 | 0.23855 | pmol/µL [µM]
NIST T1D | Cer d18:1/16:0 | Single Point | 19 | QQQ | 6 | 0.11307 | 0.016052 | 14.1966 | TRUE | 0.15052 | 0.06249 | 0.23855 | pmol/µL [µM]
NIST T1D | Cer d18:1/16:0 | Single Point | 22 | QQQ | 18 | 0.14406 | 0.0047344 | 3.2864 | TRUE | 0.15052 | 0.06249 | 0.23855 | pmol/µL [µM]
NIST T1D | Cer d18:1/16:0 | Single Point | 23 | QQQ | 6 | 0.15657 | 0.0060345 | 3.8541 | TRUE | 0.15052 | 0.06249 | 0.23855 | pmol/µL [µM]
NIST T1D | Cer d18:1/16:0 | Single Point | 24 | QQQ | 18 | 0.17223 | 0.0068526 | 3.9787 | TRUE | 0.15052 | 0.06249 | 0.23855 | pmol/µL [µM]
NIST T1D | Cer d18:1/16:0 | Single Point | 27 | QQQ | 18 | 0.13341 | 0.004815 | 3.6091 | TRUE | 0.15052 | 0.06249 | 0.23855 | pmol/µL [µM]
NIST T1D | Cer d18:1/16:0 | Single Point | 28 | QQQ | 18 | 0.16735 | 0.018506 | 11.0581 | TRUE | 0.15052 | 0.06249 | 0.23855 | pmol/µL [µM]
NIST T1D | Cer d18:1/16:0 | Single Point | 29a | TOF | 18 | 0.13481 | 0.0041196 | 3.0558 | TRUE | 0.15052 | 0.06249 | 0.23855 | pmol/µL [µM]
NIST T1D | Cer d18:1/16:0 | Single Point | 3 | QQQ | 6 | 0.16489 | 0.01063 | 6.4467 | TRUE | 0.15052 | 0.06249 | 0.23855 | pmol/µL [µM]
NIST T1D | Cer d18:1/16:0 | Single Point | 30 | QQQ | 18 | 0.18293 | 0.0099216 | 5.4237 | TRUE | 0.15052 | 0.06249 | 0.23855 | pmol/µL [µM]
NIST T1D | Cer d18:1/16:0 | Single Point | 31 | QQQ | 18 | 0.18206 | 0.0083605 | 4.5922 | TRUE | 0.15052 | 0.06249 | 0.23855 | pmol/µL [µM]
NIST T1D | Cer d18:1/16:0 | Single Point | 32 | TOF | 18 | 0.24073 | 0.038227 | 15.8796 | FALSE | 0.15052 | 0.06249 | 0.23855 | pmol/µL [µM]
NIST T1D | Cer d18:1/16:0 | Single Point | 35 | QQQ | 18 | 0.14648 | 0.010095 | 6.8916 | TRUE | 0.15052 | 0.06249 | 0.23855 | pmol/µL [µM]
NIST T1D | Cer d18:1/16:0 | Single Point | 36 | QQQ | 18 | 0.14967 | 0.0033057 | 2.2087 | TRUE | 0.15052 | 0.06249 | 0.23855 | pmol/µL [µM]
NIST T1D | Cer d18:1/16:0 | Single Point | 37 | QQQ | 18 | 0.034372 | 0.005655 | 16.4526 | FALSE | 0.15052 | 0.06249 | 0.23855 | pmol/µL [µM]
NIST T1D | Cer d18:1/16:0 | Single Point | 38 | QQQ | 18 | 0.027791 | 0.0056168 | 20.2111 | FALSE | 0.15052 | 0.06249 | 0.23855 | pmol/µL [µM]
NIST T1D | Cer d18:1/16:0 | Single Point | 4 | QQQ | 18 | 0.14783 | 0.0093081 | 6.2966 | TRUE | 0.15052 | 0.06249 | 0.23855 | pmol/µL [µM]
NIST T1D | Cer d18:1/16:0 | Single Point | 7 | Orbitrap | 18 | 0.23715 | 0.022561 | 9.5132 | TRUE | 0.15052 | 0.06249 | 0.23855 | pmol/µL [µM]
NIST T1D | Cer d18:1/18:0 | Calibration Curve | 02a | QQQ | 6 | 0.050908 | 0.0016644 | 3.2694 | TRUE | 0.046565 | 0.020303 | 0.072828 | pmol/µL [µM]
NIST T1D | Cer d18:1/18:0 | Calibration Curve | 02b | QQQ | 18 | 0.05095 | 0.0013363 | 2.6228 | TRUE | 0.046565 | 0.020303 | 0.072828 | pmol/µL [µM]
NIST T1D | Cer d18:1/18:0 | Calibration Curve | 10a | QQQ | 6 | 0.032009 | 0.0038091 | 11.9003 | TRUE | 0.046565 | 0.020303 | 0.072828 | pmol/µL [µM]
NIST T1D | Cer d18:1/18:0 | Calibration Curve | 10b | QQQ | 6 | 0.028263 | 0.0074314 | 26.2936 | TRUE | 0.046565 | 0.020303 | 0.072828 | pmol/µL [µM]
NIST T1D | Cer d18:1/18:0 | Calibration Curve | 12 | QQQ | 18 | 0.059354 | 0.0056781 | 9.5665 | TRUE | 0.046565 | 0.020303 | 0.072828 | pmol/µL [µM]
NIST T1D | Cer d18:1/18:0 | Calibration Curve | 14 | QQQ | 18 | 0.048516 | 0.0024373 | 5.0237 | TRUE | 0.046565 | 0.020303 | 0.072828 | pmol/µL [µM]
NIST T1D | Cer d18:1/18:0 | Calibration Curve | 15 | QQQ | 18 | 0.049746 | 0.0043166 | 8.6773 | TRUE | 0.046565 | 0.020303 | 0.072828 | pmol/µL [µM]
NIST T1D | Cer d18:1/18:0 | Calibration Curve | 16 | Orbitrap | 18 | 0.045605 | 0.0080998 | 17.7609 | TRUE | 0.046565 | 0.020303 | 0.072828 | pmol/µL [µM]
NIST T1D | Cer d18:1/18:0 | Calibration Curve | 17a | TOF | 6 | 0.064206 | 0.0066292 | 10.325 | TRUE | 0.046565 | 0.020303 | 0.072828 | pmol/µL [µM]
NIST T1D | Cer d18:1/18:0 | Calibration Curve | 18a | QQQ | 18 | 0.046675 | 0.0018702 | 4.0068 | TRUE | 0.046565 | 0.020303 | 0.072828 | pmol/µL [µM]
NIST T1D | Cer d18:1/18:0 | Calibration Curve | 18b | QQQ | 18 | 0.046073 | 0.0022078 | 4.792 | TRUE | 0.046565 | 0.020303 | 0.072828 | pmol/µL [µM]
NIST T1D | Cer d18:1/18:0 | Calibration Curve | 19 | QQQ | 6 | 0.045101 | 0.0086433 | 19.1642 | TRUE | 0.046565 | 0.020303 | 0.072828 | pmol/µL [µM]
NIST T1D | Cer d18:1/18:0 | Calibration Curve | 22 | QQQ | 18 | 0.048831 | 0.0022367 | 4.5805 | TRUE | 0.046565 | 0.020303 | 0.072828 | pmol/µL [µM]
NIST T1D | Cer d18:1/18:0 | Calibration Curve | 23 | QQQ | 6 | 0.054058 | 0.0027666 | 5.1178 | TRUE | 0.046565 | 0.020303 | 0.072828 | pmol/µL [µM]
NIST T1D | Cer d18:1/18:0 | Calibration Curve | 24 | QQQ | 18 | 0.049326 | 0.0025912 | 5.2532 | TRUE | 0.046565 | 0.020303 | 0.072828 | pmol/µL [µM]
NIST T1D | Cer d18:1/18:0 | Calibration Curve | 27 | QQQ | 18 | 0.028707 | 0.0016262 | 5.6648 | TRUE | 0.046565 | 0.020303 | 0.072828 | pmol/µL [µM]
NIST T1D | Cer d18:1/18:0 | Calibration Curve | 28 | QQQ | 18 | 0.040479 | 0.0009419 | 2.3269 | TRUE | 0.046565 | 0.020303 | 0.072828 | pmol/µL [µM]
NIST T1D | Cer d18:1/18:0 | Calibration Curve | 29a | TOF | 18 | 0.045039 | 0.0015243 | 3.3843 | TRUE | 0.046565 | 0.020303 | 0.072828 | pmol/µL [µM]
NIST T1D | Cer d18:1/18:0 | Calibration Curve | 3 | QQQ | 6 | 0.047323 | 0.0032002 | 6.7626 | TRUE | 0.046565 | 0.020303 | 0.072828 | pmol/µL [µM]
NIST T1D | Cer d18:1/18:0 | Calibration Curve | 30 | QQQ | 18 | 0.045423 | 0.0021617 | 4.7591 | TRUE | 0.046565 | 0.020303 | 0.072828 | pmol/µL [µM]
NIST T1D | Cer d18:1/18:0 | Calibration Curve | 31 | QQQ | 18 | 0.044723 | 0.0022241 | 4.9731 | TRUE | 0.046565 | 0.020303 | 0.072828 | pmol/µL [µM]
NIST T1D | Cer d18:1/18:0 | Calibration Curve | 32 | TOF | 18 | 0.08781 | 0.011287 | 12.8543 | FALSE | 0.046565 | 0.020303 | 0.072828 | pmol/µL [µM]
NIST T1D | Cer d18:1/18:0 | Calibration Curve | 35 | QQQ | 18 | 0.03375 | 0.0054205 | 16.0608 | TRUE | 0.046565 | 0.020303 | 0.072828 | pmol/µL [µM]
NIST T1D | Cer d18:1/18:0 | Calibration Curve | 36 | QQQ | 18 | 0.046797 | 0.0019943 | 4.2617 | TRUE | 0.046565 | 0.020303 | 0.072828 | pmol/µL [µM]
NIST T1D | Cer d18:1/18:0 | Calibration Curve | 37 | QQQ | 18 | 0.049487 | 0.003698 | 7.4726 | TRUE | 0.046565 | 0.020303 | 0.072828 | pmol/µL [µM]
NIST T1D | Cer d18:1/18:0 | Calibration Curve | 38 | QQQ | 18 | 0.011622 | 0.0024164 | 20.7907 | FALSE | 0.046565 | 0.020303 | 0.072828 | pmol/µL [µM]
NIST T1D | Cer d18:1/18:0 | Calibration Curve | 4 | QQQ | 18 | 0.052572 | 0.0031927 | 6.0731 | TRUE | 0.046565 | 0.020303 | 0.072828 | pmol/µL [µM]
NIST T1D | Cer d18:1/18:0 | Calibration Curve | 7 | Orbitrap | 18 | 0.050479 | 0.0044547 | 8.8248 | TRUE | 0.046565 | 0.020303 | 0.072828 | pmol/µL [µM]
NIST T1D | Cer d18:1/18:0 | Single Point | 02a | QQQ | 6 | 0.032935 | 0.0010544 | 3.2014 | TRUE | 0.03255 | 0.01155 | 0.05355 | pmol/µL [µM]
NIST T1D | Cer d18:1/18:0 | Single Point | 02b | QQQ | 18 | 0.033339 | 0.00086326 | 2.5893 | TRUE | 0.03255 | 0.01155 | 0.05355 | pmol/µL [µM]
NIST T1D | Cer d18:1/18:0 | Single Point | 10a | QQQ | 6 | 0.032848 | 0.0028982 | 8.8231 | TRUE | 0.03255 | 0.01155 | 0.05355 | pmol/µL [µM]
NIST T1D | Cer d18:1/18:0 | Single Point | 10b | QQQ | 6 | 0.027512 | 0.0056749 | 20.627 | TRUE | 0.03255 | 0.01155 | 0.05355 | pmol/µL [µM]
NIST T1D | Cer d18:1/18:0 | Single Point | 12 | QQQ | 18 | 0.036089 | 0.0034447 | 9.5451 | TRUE | 0.03255 | 0.01155 | 0.05355 | pmol/µL [µM]
NIST T1D | Cer d18:1/18:0 | Single Point | 14 | QQQ | 18 | 0.032156 | 0.0015675 | 4.8746 | TRUE | 0.03255 | 0.01155 | 0.05355 | pmol/µL [µM]
NIST T1D | Cer d18:1/18:0 | Single Point | 15 | QQQ | 18 | 0.032691 | 0.0028359 | 8.6748 | TRUE | 0.03255 | 0.01155 | 0.05355 | pmol/µL [µM]
NIST T1D | Cer d18:1/18:0 | Single Point | 16 | Orbitrap | 18 | 0.030651 | 0.0059136 | 19.2933 | TRUE | 0.03255 | 0.01155 | 0.05355 | pmol/µL [µM]
NIST T1D | Cer d18:1/18:0 | Single Point | 17a | TOF | 6 | 0.040394 | 0.0040891 | 10.1229 | TRUE | 0.03255 | 0.01155 | 0.05355 | pmol/µL [µM]
NIST T1D | Cer d18:1/18:0 | Single Point | 18a | QQQ | 18 | 0.034 | 0.0013164 | 3.8718 | TRUE | 0.03255 | 0.01155 | 0.05355 | pmol/µL [µM]
NIST T1D | Cer d18:1/18:0 | Single Point | 18b | QQQ | 18 | 0.026427 | 0.0011981 | 4.5335 | TRUE | 0.03255 | 0.01155 | 0.05355 | pmol/µL [µM]
NIST T1D | Cer d18:1/18:0 | Single Point | 19 | QQQ | 6 | 0.027315 | 0.0044734 | 16.377 | TRUE | 0.03255 | 0.01155 | 0.05355 | pmol/µL [µM]
NIST T1D | Cer d18:1/18:0 | Single Point | 22 | QQQ | 18 | 0.031043 | 0.0014096 | 4.5406 | TRUE | 0.03255 | 0.01155 | 0.05355 | pmol/µL [µM]
NIST T1D | Cer d18:1/18:0 | Single Point | 23 | QQQ | 6 | 0.035516 | 0.0019618 | 5.5237 | TRUE | 0.03255 | 0.01155 | 0.05355 | pmol/µL [µM]
NIST T1D | Cer d18:1/18:0 | Single Point | 24 | QQQ | 18 | 0.039116 | 0.001983 | 5.0695 | TRUE | 0.03255 | 0.01155 | 0.05355 | pmol/µL [µM]
NIST T1D | Cer d18:1/18:0 | Single Point | 27 | QQQ | 18 | 0.028887 | 0.001182 | 4.0919 | TRUE | 0.03255 | 0.01155 | 0.05355 | pmol/µL [µM]
NIST T1D | Cer d18:1/18:0 | Single Point | 28 | QQQ | 18 | 0.032485 | 0.00060141 | 1.8513 | TRUE | 0.03255 | 0.01155 | 0.05355 | pmol/µL [µM]
NIST T1D | Cer d18:1/18:0 | Single Point | 29a | TOF | 18 | 0.028789 | 0.00095891 | 3.3308 | TRUE | 0.03255 | 0.01155 | 0.05355 | pmol/µL [µM]
NIST T1D | Cer d18:1/18:0 | Single Point | 3 | QQQ | 6 | 0.031497 | 0.0021095 | 6.6976 | TRUE | 0.03255 | 0.01155 | 0.05355 | pmol/µL [µM]
NIST T1D | Cer d18:1/18:0 | Single Point | 30 | QQQ | 18 | 0.029879 | 0.0014619 | 4.8926 | TRUE | 0.03255 | 0.01155 | 0.05355 | pmol/µL [µM]
NIST T1D | Cer d18:1/18:0 | Single Point | 31 | QQQ | 18 | 0.035963 | 0.0016736 | 4.6536 | TRUE | 0.03255 | 0.01155 | 0.05355 | pmol/µL [µM]
NIST T1D | Cer d18:1/18:0 | Single Point | 32 | TOF | 18 | 0.076336 | 0.0097452 | 12.7661 | FALSE | 0.03255 | 0.01155 | 0.05355 | pmol/µL [µM]
NIST T1D | Cer d18:1/18:0 | Single Point | 35 | QQQ | 18 | 0.027906 | 0.0033589 | 12.0365 | TRUE | 0.03255 | 0.01155 | 0.05355 | pmol/µL [µM]
NIST T1D | Cer d18:1/18:0 | Single Point | 36 | QQQ | 18 | 0.030066 | 0.001247 | 4.1475 | TRUE | 0.03255 | 0.01155 | 0.05355 | pmol/µL [µM]
NIST T1D | Cer d18:1/18:0 | Single Point | 37 | QQQ | 18 | 0.027599 | 0.001987 | 7.1993 | TRUE | 0.03255 | 0.01155 | 0.05355 | pmol/µL [µM]
NIST T1D | Cer d18:1/18:0 | Single Point | 38 | QQQ | 18 | 0.0056438 | 0.0011159 | 19.773 | FALSE | 0.03255 | 0.01155 | 0.05355 | pmol/µL [µM]
NIST T1D | Cer d18:1/18:0 | Single Point | 4 | QQQ | 18 | 0.02989 | 0.0018118 | 6.0617 | TRUE | 0.03255 | 0.01155 | 0.05355 | pmol/µL [µM]
NIST T1D | Cer d18:1/18:0 | Single Point | 7 | Orbitrap | 18 | 0.034421 | 0.0028285 | 8.2176 | TRUE | 0.03255 | 0.01155 | 0.05355 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:0 | Calibration Curve | 02a | QQQ | 6 | 1.3692 | 0.030532 | 2.2299 | TRUE | 1.3321 | -0.47026 | 3.1345 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:0 | Calibration Curve | 02b | QQQ | 18 | 1.3458 | 0.030799 | 2.2884 | TRUE | 1.3321 | -0.47026 | 3.1345 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:0 | Calibration Curve | 10a | QQQ | 6 | 0.84053 | 0.18482 | 21.989 | TRUE | 1.3321 | -0.47026 | 3.1345 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:0 | Calibration Curve | 10b | QQQ | 6 | 0.84462 | 0.027327 | 3.2355 | TRUE | 1.3321 | -0.47026 | 3.1345 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:0 | Calibration Curve | 12 | QQQ | 18 | 1.3235 | 0.1165 | 8.8026 | TRUE | 1.3321 | -0.47026 | 3.1345 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:0 | Calibration Curve | 14 | QQQ | 18 | 1.3832 | 0.073623 | 5.3226 | TRUE | 1.3321 | -0.47026 | 3.1345 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:0 | Calibration Curve | 15 | QQQ | 18 | 1.5025 | 0.12353 | 8.222 | TRUE | 1.3321 | -0.47026 | 3.1345 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:0 | Calibration Curve | 16 | Orbitrap | 18 | 1.1379 | 0.16292 | 14.3172 | TRUE | 1.3321 | -0.47026 | 3.1345 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:0 | Calibration Curve | 17a | TOF | 6 | 1.1314 | 0.026838 | 2.3721 | TRUE | 1.3321 | -0.47026 | 3.1345 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:0 | Calibration Curve | 18a | QQQ | 18 | 1.1177 | 0.035075 | 3.1381 | TRUE | 1.3321 | -0.47026 | 3.1345 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:0 | Calibration Curve | 18b | QQQ | 18 | 1.3131 | 0.04823 | 3.673 | TRUE | 1.3321 | -0.47026 | 3.1345 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:0 | Calibration Curve | 19 | QQQ | 6 | 1.1561 | 0.16475 | 14.2508 | TRUE | 1.3321 | -0.47026 | 3.1345 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:0 | Calibration Curve | 22 | QQQ | 18 | 1.209 | 0.027859 | 2.3042 | TRUE | 1.3321 | -0.47026 | 3.1345 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:0 | Calibration Curve | 23 | QQQ | 6 | 1.2483 | 0.030264 | 2.4245 | TRUE | 1.3321 | -0.47026 | 3.1345 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:0 | Calibration Curve | 24 | QQQ | 18 | 1.1795 | 0.06971 | 5.9103 | TRUE | 1.3321 | -0.47026 | 3.1345 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:0 | Calibration Curve | 27 | QQQ | 18 | 0.99246 | 0.036663 | 3.6942 | TRUE | 1.3321 | -0.47026 | 3.1345 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:0 | Calibration Curve | 28 | QQQ | 18 | 1.4898 | 0.026932 | 1.8078 | TRUE | 1.3321 | -0.47026 | 3.1345 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:0 | Calibration Curve | 29a | TOF | 18 | 1.175 | 0.031273 | 2.6615 | TRUE | 1.3321 | -0.47026 | 3.1345 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:0 | Calibration Curve | 3 | QQQ | 6 | 1.2498 | 0.069331 | 5.5475 | TRUE | 1.3321 | -0.47026 | 3.1345 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:0 | Calibration Curve | 30 | QQQ | 18 | 1.0392 | 0.04693 | 4.5161 | TRUE | 1.3321 | -0.47026 | 3.1345 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:0 | Calibration Curve | 31 | QQQ | 18 | 1.2496 | 0.041154 | 3.2934 | TRUE | 1.3321 | -0.47026 | 3.1345 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:0 | Calibration Curve | 32 | TOF | 18 | 1.2282 | 0.05836 | 4.7515 | TRUE | 1.3321 | -0.47026 | 3.1345 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:0 | Calibration Curve | 35 | QQQ | 18 | 0.9796 | 0.0377 | 3.8485 | TRUE | 1.3321 | -0.47026 | 3.1345 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:0 | Calibration Curve | 36 | QQQ | 18 | 1.2174 | 0.030279 | 2.4871 | TRUE | 1.3321 | -0.47026 | 3.1345 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:0 | Calibration Curve | 37 | QQQ | 18 | 5.7679 | 0.3688 | 6.3939 | FALSE | 1.3321 | -0.47026 | 3.1345 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:0 | Calibration Curve | 38 | QQQ | 18 | 0.27497 | 0.061036 | 22.1974 | TRUE | 1.3321 | -0.47026 | 3.1345 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:0 | Calibration Curve | 4 | QQQ | 18 | 1.2651 | 0.078249 | 6.1855 | TRUE | 1.3321 | -0.47026 | 3.1345 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:0 | Calibration Curve | 7 | Orbitrap | 18 | 1.2685 | 0.077954 | 6.1453 | TRUE | 1.3321 | -0.47026 | 3.1345 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:0 | Single Point | 02a | QQQ | 6 | 1.3288 | 0.029263 | 2.2022 | TRUE | 1.2175 | 0.57426 | 1.8608 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:0 | Single Point | 02b | QQQ | 18 | 1.3091 | 0.029514 | 2.2545 | TRUE | 1.2175 | 0.57426 | 1.8608 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:0 | Single Point | 10a | QQQ | 6 | 1.1243 | 0.21968 | 19.5401 | TRUE | 1.2175 | 0.57426 | 1.8608 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:0 | Single Point | 10b | QQQ | 6 | 1.1363 | 0.033839 | 2.9779 | TRUE | 1.2175 | 0.57426 | 1.8608 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:0 | Single Point | 12 | QQQ | 18 | 1.0299 | 0.090786 | 8.8149 | TRUE | 1.2175 | 0.57426 | 1.8608 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:0 | Single Point | 14 | QQQ | 18 | 1.4422 | 0.075454 | 5.2317 | TRUE | 1.2175 | 0.57426 | 1.8608 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:0 | Single Point | 15 | QQQ | 18 | 1.1394 | 0.092516 | 8.1194 | TRUE | 1.2175 | 0.57426 | 1.8608 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:0 | Single Point | 16 | Orbitrap | 18 | 1.2547 | 0.182 | 14.5048 | TRUE | 1.2175 | 0.57426 | 1.8608 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:0 | Single Point | 17a | TOF | 6 | 1.4127 | 0.034358 | 2.4321 | TRUE | 1.2175 | 0.57426 | 1.8608 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:0 | Single Point | 18a | QQQ | 18 | 1.2431 | 0.03895 | 3.1334 | TRUE | 1.2175 | 0.57426 | 1.8608 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:0 | Single Point | 18b | QQQ | 18 | 1.2878 | 0.045685 | 3.5477 | TRUE | 1.2175 | 0.57426 | 1.8608 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:0 | Single Point | 19 | QQQ | 6 | 1.2914 | 0.18245 | 14.1287 | TRUE | 1.2175 | 0.57426 | 1.8608 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:0 | Single Point | 22 | QQQ | 18 | 1.2931 | 0.029871 | 2.31 | TRUE | 1.2175 | 0.57426 | 1.8608 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:0 | Single Point | 23 | QQQ | 6 | 1.4238 | 0.035738 | 2.51 | TRUE | 1.2175 | 0.57426 | 1.8608 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:0 | Single Point | 24 | QQQ | 18 | 0.3858 | 0.022591 | 5.8556 | FALSE | 1.2175 | 0.57426 | 1.8608 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:0 | Single Point | 27 | QQQ | 18 | 1.1596 | 0.040088 | 3.4571 | TRUE | 1.2175 | 0.57426 | 1.8608 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:0 | Single Point | 28 | QQQ | 18 | 1.3139 | 0.022653 | 1.7242 | TRUE | 1.2175 | 0.57426 | 1.8608 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:0 | Single Point | 29a | TOF | 18 | 1.2658 | 0.033348 | 2.6346 | TRUE | 1.2175 | 0.57426 | 1.8608 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:0 | Single Point | 3 | QQQ | 6 | 1.4251 | 0.078541 | 5.5114 | TRUE | 1.2175 | 0.57426 | 1.8608 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:0 | Single Point | 30 | QQQ | 18 | 1.028 | 0.046749 | 4.5476 | TRUE | 1.2175 | 0.57426 | 1.8608 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:0 | Single Point | 31 | QQQ | 18 | 2.0058 | 0.063506 | 3.1661 | FALSE | 1.2175 | 0.57426 | 1.8608 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:0 | Single Point | 32 | TOF | 18 | 1.4583 | 0.06967 | 4.7774 | TRUE | 1.2175 | 0.57426 | 1.8608 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:0 | Single Point | 35 | QQQ | 18 | 1.0652 | 0.037472 | 3.5179 | TRUE | 1.2175 | 0.57426 | 1.8608 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:0 | Single Point | 36 | QQQ | 18 | 1.2854 | 0.031763 | 2.4711 | TRUE | 1.2175 | 0.57426 | 1.8608 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:0 | Single Point | 37 | QQQ | 18 | 0.96751 | 0.061853 | 6.393 | TRUE | 1.2175 | 0.57426 | 1.8608 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:0 | Single Point | 38 | QQQ | 18 | 0.24145 | 0.05056 | 20.9399 | FALSE | 1.2175 | 0.57426 | 1.8608 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:0 | Single Point | 4 | QQQ | 18 | 1.4132 | 0.08759 | 6.198 | TRUE | 1.2175 | 0.57426 | 1.8608 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:0 | Single Point | 7 | Orbitrap | 18 | 1.3589 | 0.080921 | 5.9548 | TRUE | 1.2175 | 0.57426 | 1.8608 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:1 | Calibration Curve | 02a | QQQ | 6 | 0.49039 | 0.011303 | 2.3049 | TRUE | 0.41461 | 0.20343 | 0.62579 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:1 | Calibration Curve | 02b | QQQ | 18 | 0.48476 | 0.011301 | 2.3312 | TRUE | 0.41461 | 0.20343 | 0.62579 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:1 | Calibration Curve | 10a | QQQ | 6 | 0.30823 | 0.073919 | 23.9814 | TRUE | 0.41461 | 0.20343 | 0.62579 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:1 | Calibration Curve | 10b | QQQ | 6 | 0.21876 | 0.01565 | 7.154 | TRUE | 0.41461 | 0.20343 | 0.62579 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:1 | Calibration Curve | 12 | QQQ | 18 | 0.47506 | 0.041822 | 8.8035 | TRUE | 0.41461 | 0.20343 | 0.62579 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:1 | Calibration Curve | 14 | QQQ | 18 | 0.46322 | 0.029422 | 6.3516 | TRUE | 0.41461 | 0.20343 | 0.62579 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:1 | Calibration Curve | 15 | QQQ | 18 | 0.54209 | 0.040693 | 7.5066 | TRUE | 0.41461 | 0.20343 | 0.62579 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:1 | Calibration Curve | 16 | Orbitrap | 18 | 0.44485 | 0.16727 | 37.6026 | TRUE | 0.41461 | 0.20343 | 0.62579 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:1 | Calibration Curve | 17a | TOF | 6 | 0.49068 | 0.011717 | 2.3879 | TRUE | 0.41461 | 0.20343 | 0.62579 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:1 | Calibration Curve | 18a | QQQ | 18 | 0.41419 | 0.011333 | 2.7361 | TRUE | 0.41461 | 0.20343 | 0.62579 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:1 | Calibration Curve | 18b | QQQ | 18 | 0.43604 | 0.01729 | 3.9653 | TRUE | 0.41461 | 0.20343 | 0.62579 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:1 | Calibration Curve | 19 | QQQ | 6 | 0.38303 | 0.036305 | 9.4785 | TRUE | 0.41461 | 0.20343 | 0.62579 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:1 | Calibration Curve | 22 | QQQ | 18 | 0.41971 | 0.010734 | 2.5575 | TRUE | 0.41461 | 0.20343 | 0.62579 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:1 | Calibration Curve | 23 | QQQ | 6 | 0.52429 | 0.017504 | 3.3386 | TRUE | 0.41461 | 0.20343 | 0.62579 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:1 | Calibration Curve | 24 | QQQ | 18 | 0.42136 | 0.020845 | 4.9472 | TRUE | 0.41461 | 0.20343 | 0.62579 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:1 | Calibration Curve | 27 | QQQ | 18 | 0.29581 | 0.010705 | 3.6187 | TRUE | 0.41461 | 0.20343 | 0.62579 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:1 | Calibration Curve | 28 | QQQ | 18 | 0.48358 | 0.009172 | 1.8967 | TRUE | 0.41461 | 0.20343 | 0.62579 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:1 | Calibration Curve | 29a | TOF | 18 | 0.41578 | 0.01906 | 4.5841 | TRUE | 0.41461 | 0.20343 | 0.62579 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:1 | Calibration Curve | 3 | QQQ | 6 | 0.44158 | 0.02315 | 5.2424 | TRUE | 0.41461 | 0.20343 | 0.62579 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:1 | Calibration Curve | 30 | QQQ | 18 | 0.41657 | 0.022138 | 5.3145 | TRUE | 0.41461 | 0.20343 | 0.62579 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:1 | Calibration Curve | 31 | QQQ | 18 | 0.4294 | 0.017532 | 4.0829 | TRUE | 0.41461 | 0.20343 | 0.62579 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:1 | Calibration Curve | 32 | TOF | 18 | 0.61473 | 0.050521 | 8.2185 | TRUE | 0.41461 | 0.20343 | 0.62579 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:1 | Calibration Curve | 35 | QQQ | 18 | 0.29877 | 0.010995 | 3.6802 | TRUE | 0.41461 | 0.20343 | 0.62579 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:1 | Calibration Curve | 36 | QQQ | 18 | 0.29588 | 0.018042 | 6.0976 | TRUE | 0.41461 | 0.20343 | 0.62579 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:1 | Calibration Curve | 37 | QQQ | 18 | 0.38475 | 0.02764 | 7.1837 | TRUE | 0.41461 | 0.20343 | 0.62579 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:1 | Calibration Curve | 38 | QQQ | 18 | 0.090563 | 0.021922 | 24.2058 | FALSE | 0.41461 | 0.20343 | 0.62579 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:1 | Calibration Curve | 4 | QQQ | 18 | 0.45031 | 0.029811 | 6.62 | TRUE | 0.41461 | 0.20343 | 0.62579 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:1 | Calibration Curve | 7 | Orbitrap | 18 | 0.47472 | 0.020986 | 4.4207 | TRUE | 0.41461 | 0.20343 | 0.62579 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:1 | Single Point | 02a | QQQ | 6 | 0.44934 | 0.010155 | 2.26 | TRUE | 0.40676 | 0.14599 | 0.66753 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:1 | Single Point | 02b | QQQ | 18 | 0.4484 | 0.01026 | 2.288 | TRUE | 0.40676 | 0.14599 | 0.66753 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:1 | Single Point | 10a | QQQ | 6 | 0.27757 | 0.054192 | 19.5237 | TRUE | 0.40676 | 0.14599 | 0.66753 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:1 | Single Point | 10b | QQQ | 6 | 0.34018 | 0.018753 | 5.5127 | TRUE | 0.40676 | 0.14599 | 0.66753 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:1 | Single Point | 12 | QQQ | 18 | 0.34654 | 0.030674 | 8.8513 | TRUE | 0.40676 | 0.14599 | 0.66753 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:1 | Single Point | 14 | QQQ | 18 | 0.47139 | 0.028864 | 6.1231 | TRUE | 0.40676 | 0.14599 | 0.66753 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:1 | Single Point | 15 | QQQ | 18 | 0.38738 | 0.028386 | 7.3278 | TRUE | 0.40676 | 0.14599 | 0.66753 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:1 | Single Point | 16 | Orbitrap | 18 | 0.24674 | 0.10081 | 40.8559 | TRUE | 0.40676 | 0.14599 | 0.66753 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:1 | Single Point | 17a | TOF | 6 | 0.62599 | 0.015899 | 2.5398 | TRUE | 0.40676 | 0.14599 | 0.66753 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:1 | Single Point | 18a | QQQ | 18 | 0.47686 | 0.01303 | 2.7325 | TRUE | 0.40676 | 0.14599 | 0.66753 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:1 | Single Point | 18b | QQQ | 18 | 0.46732 | 0.017179 | 3.676 | TRUE | 0.40676 | 0.14599 | 0.66753 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:1 | Single Point | 19 | QQQ | 6 | 0.34452 | 0.033162 | 9.6255 | TRUE | 0.40676 | 0.14599 | 0.66753 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:1 | Single Point | 22 | QQQ | 18 | 0.41656 | 0.010725 | 2.5746 | TRUE | 0.40676 | 0.14599 | 0.66753 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:1 | Single Point | 23 | QQQ | 6 | 0.53908 | 0.019708 | 3.6558 | TRUE | 0.40676 | 0.14599 | 0.66753 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:1 | Single Point | 24 | QQQ | 18 | 0.12369 | 0.0060048 | 4.8547 | FALSE | 0.40676 | 0.14599 | 0.66753 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:1 | Single Point | 27 | QQQ | 18 | 0.36701 | 0.011081 | 3.0194 | TRUE | 0.40676 | 0.14599 | 0.66753 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:1 | Single Point | 28 | QQQ | 18 | 0.42632 | 0.0073934 | 1.7342 | TRUE | 0.40676 | 0.14599 | 0.66753 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:1 | Single Point | 29a | TOF | 18 | 0.42905 | 0.01949 | 4.5424 | TRUE | 0.40676 | 0.14599 | 0.66753 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:1 | Single Point | 3 | QQQ | 6 | 0.43477 | 0.023184 | 5.3325 | TRUE | 0.40676 | 0.14599 | 0.66753 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:1 | Single Point | 30 | QQQ | 18 | 0.42578 | 0.02314 | 5.4347 | TRUE | 0.40676 | 0.14599 | 0.66753 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:1 | Single Point | 31 | QQQ | 18 | 0.52463 | 0.020237 | 3.8574 | TRUE | 0.40676 | 0.14599 | 0.66753 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:1 | Single Point | 32 | TOF | 18 | 0.7013 | 0.058017 | 8.2728 | FALSE | 0.40676 | 0.14599 | 0.66753 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:1 | Single Point | 35 | QQQ | 18 | 0.35734 | 0.010029 | 2.8066 | TRUE | 0.40676 | 0.14599 | 0.66753 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:1 | Single Point | 36 | QQQ | 18 | 0.42424 | 0.014512 | 3.4208 | TRUE | 0.40676 | 0.14599 | 0.66753 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:1 | Single Point | 37 | QQQ | 18 | 0.28266 | 0.020067 | 7.0993 | TRUE | 0.40676 | 0.14599 | 0.66753 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:1 | Single Point | 38 | QQQ | 18 | 0.081243 | 0.01642 | 20.2114 | FALSE | 0.40676 | 0.14599 | 0.66753 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:1 | Single Point | 4 | QQQ | 18 | 0.46247 | 0.031115 | 6.728 | TRUE | 0.40676 | 0.14599 | 0.66753 | pmol/µL [µM]
NIST T1D | Cer d18:1/24:1 | Single Point | 7 | Orbitrap | 18 | 0.51086 | 0.021009 | 4.1125 | TRUE | 0.40676 | 0.14599 | 0.66753 | pmol/µL [µM]
NIST YAA | Cer d18:1/16:0 | Calibration Curve | 02a | QQQ | 6 | 0.08617 | 0.0049364 | 5.7286 | TRUE | 0.071458 | 0.017038 | 0.12588 | pmol/µL [µM]
NIST YAA | Cer d18:1/16:0 | Calibration Curve | 02b | QQQ | 18 | 0.088457 | 0.0031264 | 3.5344 | TRUE | 0.071458 | 0.017038 | 0.12588 | pmol/µL [µM]
NIST YAA | Cer d18:1/16:0 | Calibration Curve | 10a | QQQ | 6 | 0.0361 | 0.008617 | 23.8698 | TRUE | 0.071458 | 0.017038 | 0.12588 | pmol/µL [µM]
NIST YAA | Cer d18:1/16:0 | Calibration Curve | 10b | QQQ | 6 | 0.039138 | 0.0083 | 21.2073 | TRUE | 0.071458 | 0.017038 | 0.12588 | pmol/µL [µM]
NIST YAA | Cer d18:1/16:0 | Calibration Curve | 12 | QQQ | 18 | 0.085349 | 0.0027733 | 3.2494 | TRUE | 0.071458 | 0.017038 | 0.12588 | pmol/µL [µM]
NIST YAA | Cer d18:1/16:0 | Calibration Curve | 14 | QQQ | 18 | 0.070539 | 0.0025972 | 3.682 | TRUE | 0.071458 | 0.017038 | 0.12588 | pmol/µL [µM]
NIST YAA | Cer d18:1/16:0 | Calibration Curve | 15 | QQQ | 18 | 0.076404 | 0.0036647 | 4.7964 | TRUE | 0.071458 | 0.017038 | 0.12588 | pmol/µL [µM]
NIST YAA | Cer d18:1/16:0 | Calibration Curve | 16 | Orbitrap | 18 | 0.070472 | 0.0068136 | 9.6685 | TRUE | 0.071458 | 0.017038 | 0.12588 | pmol/µL [µM]
NIST YAA | Cer d18:1/16:0 | Calibration Curve | 17a | TOF | 6 | 0.081913 | 0.018667 | 22.7892 | TRUE | 0.071458 | 0.017038 | 0.12588 | pmol/µL [µM]
NIST YAA | Cer d18:1/16:0 | Calibration Curve | 18a | QQQ | 18 | 0.073411 | 0.0021788 | 2.9679 | TRUE | 0.071458 | 0.017038 | 0.12588 | pmol/µL [µM]
NIST YAA | Cer d18:1/16:0 | Calibration Curve | 18b | QQQ | 18 | 0.076251 | 0.0062148 | 8.1505 | TRUE | 0.071458 | 0.017038 | 0.12588 | pmol/µL [µM]
NIST YAA | Cer d18:1/16:0 | Calibration Curve | 19 | QQQ | 6 | 0.072927 | 0.01278 | 17.5243 | TRUE | 0.071458 | 0.017038 | 0.12588 | pmol/µL [µM]
NIST YAA | Cer d18:1/16:0 | Calibration Curve | 22 | QQQ | 18 | 0.077033 | 0.0055962 | 7.2647 | TRUE | 0.071458 | 0.017038 | 0.12588 | pmol/µL [µM]
NIST YAA | Cer d18:1/16:0 | Calibration Curve | 23 | QQQ | 6 | 0.084852 | 0.0053477 | 6.3024 | TRUE | 0.071458 | 0.017038 | 0.12588 | pmol/µL [µM]
NIST YAA | Cer d18:1/16:0 | Calibration Curve | 24 | QQQ | 18 | 0.075992 | 0.0047169 | 6.2071 | TRUE | 0.071458 | 0.017038 | 0.12588 | pmol/µL [µM]
NIST YAA | Cer d18:1/16:0 | Calibration Curve | 27 | QQQ | 18 | 0.042906 | 0.0059929 | 13.9674 | TRUE | 0.071458 | 0.017038 | 0.12588 | pmol/µL [µM]
NIST YAA | Cer d18:1/16:0 | Calibration Curve | 28 | QQQ | 18 | -0.00022559 | 0.016815 | -7453.4635 | FALSE | 0.071458 | 0.017038 | 0.12588 | pmol/µL [µM]
NIST YAA | Cer d18:1/16:0 | Calibration Curve | 29a | TOF | 18 | 0.07375 | 0.001804 | 2.4461 | TRUE | 0.071458 | 0.017038 | 0.12588 | pmol/µL [µM]
NIST YAA | Cer d18:1/16:0 | Calibration Curve | 3 | QQQ | 6 | 0.053191 | 0.011092 | 20.8523 | TRUE | 0.071458 | 0.017038 | 0.12588 | pmol/µL [µM]
NIST YAA | Cer d18:1/16:0 | Calibration Curve | 30 | QQQ | 18 | 0.13672 | 0.0040672 | 2.9748 | FALSE | 0.071458 | 0.017038 | 0.12588 | pmol/µL [µM]
NIST YAA | Cer d18:1/16:0 | Calibration Curve | 31 | QQQ | 18 | 0.071548 | 0.0038136 | 5.3301 | TRUE | 0.071458 | 0.017038 | 0.12588 | pmol/µL [µM]
NIST YAA | Cer d18:1/16:0 | Calibration Curve | 32 | TOF | 18 | 0.10916 | 0.025038 | 22.9362 | TRUE | 0.071458 | 0.017038 | 0.12588 | pmol/µL [µM]
NIST YAA | Cer d18:1/16:0 | Calibration Curve | 35 | QQQ | 18 | 0.04812 | 0.0076998 | 16.0013 | TRUE | 0.071458 | 0.017038 | 0.12588 | pmol/µL [µM]
NIST YAA | Cer d18:1/16:0 | Calibration Curve | 36 | QQQ | 18 | 0.074149 | 0.002334 | 3.1477 | TRUE | 0.071458 | 0.017038 | 0.12588 | pmol/µL [µM]
NIST YAA | Cer d18:1/16:0 | Calibration Curve | 37 | QQQ | 18 | 0.083401 | 0.012522 | 15.0148 | TRUE | 0.071458 | 0.017038 | 0.12588 | pmol/µL [µM]
NIST YAA | Cer d18:1/16:0 | Calibration Curve | 38 | QQQ | 18 | 0.02325 | 0.0074724 | 32.1394 | TRUE | 0.071458 | 0.017038 | 0.12588 | pmol/µL [µM]
NIST YAA | Cer d18:1/16:0 | Calibration Curve | 4 | QQQ | 18 | 0.078968 | 0.0020967 | 2.6551 | TRUE | 0.071458 | 0.017038 | 0.12588 | pmol/µL [µM]
NIST YAA | Cer d18:1/16:0 | Calibration Curve | 7 | Orbitrap | 18 | 0.11086 | 0.014624 | 13.1914 | TRUE | 0.071458 | 0.017038 | 0.12588 | pmol/µL [µM]
NIST YAA | Cer d18:1/16:0 | Single Point | 02a | QQQ | 6 | 0.11709 | 0.0064902 | 5.5429 | TRUE | 0.10574 | 0.041328 | 0.17015 | pmol/µL [µM]
NIST YAA | Cer d18:1/16:0 | Single Point | 02b | QQQ | 18 | 0.12421 | 0.0043275 | 3.4839 | TRUE | 0.10574 | 0.041328 | 0.17015 | pmol/µL [µM]
NIST YAA | Cer d18:1/16:0 | Single Point | 10a | QQQ | 6 | 0.11729 | 0.01489 | 12.6953 | TRUE | 0.10574 | 0.041328 | 0.17015 | pmol/µL [µM]
NIST YAA | Cer d18:1/16:0 | Single Point | 10b | QQQ | 6 | 0.10211 | 0.012848 | 12.5826 | TRUE | 0.10574 | 0.041328 | 0.17015 | pmol/µL [µM]
NIST YAA | Cer d18:1/16:0 | Single Point | 12 | QQQ | 18 | 0.082564 | 0.0026859 | 3.2531 | TRUE | 0.10574 | 0.041328 | 0.17015 | pmol/µL [µM]
NIST YAA | Cer d18:1/16:0 | Single Point | 14 | QQQ | 18 | 0.11628 | 0.0041373 | 3.5582 | TRUE | 0.10574 | 0.041328 | 0.17015 | pmol/µL [µM]
NIST YAA | Cer d18:1/16:0 | Single Point | 15 | QQQ | 18 | 0.10528 | 0.0049649 | 4.7158 | TRUE | 0.10574 | 0.041328 | 0.17015 | pmol/µL [µM]
NIST YAA | Cer d18:1/16:0 | Single Point | 16 | Orbitrap | 18 | 0.10751 | 0.010777 | 10.0239 | TRUE | 0.10574 | 0.041328 | 0.17015 | pmol/µL [µM]
NIST YAA | Cer d18:1/16:0 | Single Point | 17a | TOF | 6 | 0.092576 | 0.02048 | 22.1219 | TRUE | 0.10574 | 0.041328 | 0.17015 | pmol/µL [µM]
NIST YAA | Cer d18:1/16:0 | Single Point | 18a | QQQ | 18 | 0.10518 | 0.0030131 | 2.8647 | TRUE | 0.10574 | 0.041328 | 0.17015 | pmol/µL [µM]
NIST YAA | Cer d18:1/16:0 | Single Point | 18b | QQQ | 18 | 0.10676 | 0.0081807 | 7.6626 | TRUE | 0.10574 | 0.041328 | 0.17015 | pmol/µL [µM]
NIST YAA | Cer d18:1/16:0 | Single Point | 19 | QQQ | 6 | 0.073619 | 0.011028 | 14.9802 | TRUE | 0.10574 | 0.041328 | 0.17015 | pmol/µL [µM]
NIST YAA | Cer d18:1/16:0 | Single Point | 22 | QQQ | 18 | 0.10474 | 0.0075341 | 7.1934 | TRUE | 0.10574 | 0.041328 | 0.17015 | pmol/µL [µM]
NIST YAA | Cer d18:1/16:0 | Single Point | 23 | QQQ | 6 | 0.11244 | 0.0074016 | 6.583 | TRUE | 0.10574 | 0.041328 | 0.17015 | pmol/µL [µM]
NIST YAA | Cer d18:1/16:0 | Single Point | 24 | QQQ | 18 | 0.12416 | 0.0074486 | 5.9994 | TRUE | 0.10574 | 0.041328 | 0.17015 | pmol/µL [µM]
NIST YAA | Cer d18:1/16:0 | Single Point | 27 | QQQ | 18 | 0.097911 | 0.0086282 | 8.8123 | TRUE | 0.10574 | 0.041328 | 0.17015 | pmol/µL [µM]
NIST YAA | Cer d18:1/16:0 | Single Point | 28 | QQQ | 18 | 0.12014 | 0.022033 | 18.3395 | TRUE | 0.10574 | 0.041328 | 0.17015 | pmol/µL [µM]
NIST YAA | Cer d18:1/16:0 | Single Point | 29a | TOF | 18 | 0.097047 | 0.0023051 | 2.3753 | TRUE | 0.10574 | 0.041328 | 0.17015 | pmol/µL [µM]
NIST YAA | Cer d18:1/16:0 | Single Point | 3 | QQQ | 6 | 0.091007 | 0.017942 | 19.7154 | TRUE | 0.10574 | 0.041328 | 0.17015 | pmol/µL [µM]
NIST YAA | Cer d18:1/16:0 | Single Point | 30 | QQQ | 18 | 0.17564 | 0.005203 | 2.9624 | FALSE | 0.10574 | 0.041328 | 0.17015 | pmol/µL [µM]
NIST YAA | Cer d18:1/16:0 | Single Point | 31 | QQQ | 18 | 0.11874 | 0.0059888 | 5.0436 | TRUE | 0.10574 | 0.041328 | 0.17015 | pmol/µL [µM]
NIST YAA | Cer d18:1/16:0 | Single Point | 32 | TOF | 18 | 0.15419 | 0.03515 | 22.7971 | TRUE | 0.10574 | 0.041328 | 0.17015 | pmol/µL [µM]
NIST YAA | Cer d18:1/16:0 | Single Point | 35 | QQQ | 18 | 0.10449 | 0.010549 | 10.0958 | TRUE | 0.10574 | 0.041328 | 0.17015 | pmol/µL [µM]
NIST YAA | Cer d18:1/16:0 | Single Point | 36 | QQQ | 18 | 0.10476 | 0.0032588 | 3.1106 | TRUE | 0.10574 | 0.041328 | 0.17015 | pmol/µL [µM]
NIST YAA | Cer d18:1/16:0 | Single Point | 37 | QQQ | 18 | 0.023409 | 0.0033484 | 14.304 | FALSE | 0.10574 | 0.041328 | 0.17015 | pmol/µL [µM]
NIST YAA | Cer d18:1/16:0 | Single Point | 38 | QQQ | 18 | 0.021392 | 0.0073184 | 34.2108 | FALSE | 0.10574 | 0.041328 | 0.17015 | pmol/µL [µM]
NIST YAA | Cer d18:1/16:0 | Single Point | 4 | QQQ | 18 | 0.097692 | 0.0025814 | 2.6424 | TRUE | 0.10574 | 0.041328 | 0.17015 | pmol/µL [µM]
NIST YAA | Cer d18:1/16:0 | Single Point | 7 | Orbitrap | 18 | 0.16251 | 0.019714 | 12.1307 | TRUE | 0.10574 | 0.041328 | 0.17015 | pmol/µL [µM]
NIST YAA | Cer d18:1/18:0 | Calibration Curve | 02a | QQQ | 6 | 0.026245 | 0.0004667 | 1.7782 | TRUE | 0.023837 | 0.0053582 | 0.042316 | pmol/µL [µM]
NIST YAA | Cer d18:1/18:0 | Calibration Curve | 02b | QQQ | 18 | 0.027975 | 0.0012008 | 4.2925 | TRUE | 0.023837 | 0.0053582 | 0.042316 | pmol/µL [µM]
NIST YAA | Cer d18:1/18:0 | Calibration Curve | 10a | QQQ | 6 | 0.012212 | 0.0047273 | 38.709 | TRUE | 0.023837 | 0.0053582 | 0.042316 | pmol/µL [µM]
NIST YAA | Cer d18:1/18:0 | Calibration Curve | 10b | QQQ | 6 | 0.00999 | 0.0031048 | 31.0794 | TRUE | 0.023837 | 0.0053582 | 0.042316 | pmol/µL [µM]
NIST YAA | Cer d18:1/18:0 | Calibration Curve | 12 | QQQ | 18 | 0.031793 | 0.0018418 | 5.7931 | TRUE | 0.023837 | 0.0053582 | 0.042316 | pmol/µL [µM]
NIST YAA | Cer d18:1/18:0 | Calibration Curve | 14 | QQQ | 18 | 0.023618 | 0.0010446 | 4.4229 | TRUE | 0.023837 | 0.0053582 | 0.042316 | pmol/µL [µM]
NIST YAA | Cer d18:1/18:0 | Calibration Curve | 15 | QQQ | 18 | 0.025133 | 0.0016463 | 6.5504 | TRUE | 0.023837 | 0.0053582 | 0.042316 | pmol/µL [µM]
NIST YAA | Cer d18:1/18:0 | Calibration Curve | 16 | Orbitrap | 18 | 0.023528 | 0.0024856 | 10.5644 | TRUE | 0.023837 | 0.0053582 | 0.042316 | pmol/µL [µM]
NIST YAA | Cer d18:1/18:0 | Calibration Curve | 17a | TOF | 6 | 0.033392 | 0.0037412 | 11.204 | TRUE | 0.023837 | 0.0053582 | 0.042316 | pmol/µL [µM]
NIST YAA | Cer d18:1/18:0 | Calibration Curve | 18a | QQQ | 18 | 0.022926 | 0.00073298 | 3.1972 | TRUE | 0.023837 | 0.0053582 | 0.042316 | pmol/µL [µM]
NIST YAA | Cer d18:1/18:0 | Calibration Curve | 18b | QQQ | 18 | 0.02191 | 0.0012628 | 5.7633 | TRUE | 0.023837 | 0.0053582 | 0.042316 | pmol/µL [µM]
NIST YAA | Cer d18:1/18:0 | Calibration Curve | 19 | QQQ | 6 | 0.022128 | 0.0049085 | 22.1825 | TRUE | 0.023837 | 0.0053582 | 0.042316 | pmol/µL [µM]
NIST YAA | Cer d18:1/18:0 | Calibration Curve | 22 | QQQ | 18 | 0.024965 | 0.001567 | 6.2769 | TRUE | 0.023837 | 0.0053582 | 0.042316 | pmol/µL [µM]
NIST YAA | Cer d18:1/18:0 | Calibration Curve | 23 | QQQ | 6 | 0.030357 | 0.0020374 | 6.7117 | TRUE | 0.023837 | 0.0053582 | 0.042316 | pmol/µL [µM]
NIST YAA | Cer d18:1/18:0 | Calibration Curve | 24 | QQQ | 18 | 0.024978 | 0.0015645 | 6.2637 | TRUE | 0.023837 | 0.0053582 | 0.042316 | pmol/µL [µM]
NIST YAA | Cer d18:1/18:0 | Calibration Curve | 27 | QQQ | 18 | 0.0094917 | 0.0019788 | 20.8474 | TRUE | 0.023837 | 0.0053582 | 0.042316 | pmol/µL [µM]
NIST YAA | Cer d18:1/18:0 | Calibration Curve | 28 | QQQ | 18 | 0.016202 | 0.00045281 | 2.7948 | TRUE | 0.023837 | 0.0053582 | 0.042316 | pmol/µL [µM]
NIST YAA | Cer d18:1/18:0 | Calibration Curve | 29a | TOF | 18 | 0.022698 | 0.00080731 | 3.5567 | TRUE | 0.023837 | 0.0053582 | 0.042316 | pmol/µL [µM]
NIST YAA | Cer d18:1/18:0 | Calibration Curve | 3 | QQQ | 6 | 0.051339 | 0.0028126 | 5.4784 | FALSE | 0.023837 | 0.0053582 | 0.042316 | pmol/µL [µM]
NIST YAA | Cer d18:1/18:0 | Calibration Curve | 30 | QQQ | 18 | 0.028151 | 0.0019095 | 6.7831 | TRUE | 0.023837 | 0.0053582 | 0.042316 | pmol/µL [µM]
NIST YAA | Cer d18:1/18:0 | Calibration Curve | 31 | QQQ | 18 | 0.021247 | 0.0011732 | 5.5218 | TRUE | 0.023837 | 0.0053582 | 0.042316 | pmol/µL [µM]
NIST YAA | Cer d18:1/18:0 | Calibration Curve | 32 | TOF | 18 | 0.038295 | 0.0099268 | 25.9219 | TRUE | 0.023837 | 0.0053582 | 0.042316 | pmol/µL [µM]
NIST YAA | Cer d18:1/18:0 | Calibration Curve | 35 | QQQ | 18 | 0.011942 | 0.0030196 | 25.2855 | TRUE | 0.023837 | 0.0053582 | 0.042316 | pmol/µL [µM]
NIST YAA | Cer d18:1/18:0 | Calibration Curve | 36 | QQQ | 18 | 0.023779 | 0.00075702 | 3.1836 | TRUE | 0.023837 | 0.0053582 | 0.042316 | pmol/µL [µM]
NIST YAA | Cer d18:1/18:0 | Calibration Curve | 37 | QQQ | 18 | 0.029529 | 0.0094291 | 31.9315 | TRUE | 0.023837 | 0.0053582 | 0.042316 | pmol/µL [µM]
NIST YAA | Cer d18:1/18:0 | Calibration Curve | 38 | QQQ | 18 | 0.0062316 | 0.0022005 | 35.3125 | TRUE | 0.023837 | 0.0053582 | 0.042316 | pmol/µL [µM]
NIST YAA | Cer d18:1/18:0 | Calibration Curve | 4 | QQQ | 18 | 0.025456 | 0.00079292 | 3.1148 | TRUE | 0.023837 | 0.0053582 | 0.042316 | pmol/µL [µM]
NIST YAA | Cer d18:1/18:0 | Calibration Curve | 7 | Orbitrap | 18 | 0.021928 | 0.0030501 | 13.9096 | TRUE | 0.023837 | 0.0053582 | 0.042316 | pmol/µL [µM]
NIST YAA | Cer d18:1/18:0 | Single Point | 02a | QQQ | 6 | 0.017311 | 0.00029565 | 1.7079 | TRUE | 0.017431 | 0.0061287 | 0.028734 | pmol/µL [µM]
NIST YAA | Cer d18:1/18:0 | Single Point | 02b | QQQ | 18 | 0.018497 | 0.00077572 | 4.1936 | TRUE | 0.017431 | 0.0061287 | 0.028734 | pmol/µL [µM]
NIST YAA | Cer d18:1/18:0 | Single Point | 10a | QQQ | 6 | 0.017786 | 0.0035968 | 20.223 | TRUE | 0.017431 | 0.0061287 | 0.028734 | pmol/µL [µM]
NIST YAA | Cer d18:1/18:0 | Single Point | 10b | QQQ | 6 | 0.013558 | 0.002371 | 17.4877 | TRUE | 0.017431 | 0.0061287 | 0.028734 | pmol/µL [µM]
NIST YAA | Cer d18:1/18:0 | Single Point | 12 | QQQ | 18 | 0.019369 | 0.0011174 | 5.7689 | TRUE | 0.017431 | 0.0061287 | 0.028734 | pmol/µL [µM]
NIST YAA | Cer d18:1/18:0 | Single Point | 14 | QQQ | 18 | 0.016143 | 0.00067179 | 4.1614 | TRUE | 0.017431 | 0.0061287 | 0.028734 | pmol/µL [µM]
NIST YAA | Cer d18:1/18:0 | Single Point | 15 | QQQ | 18 | 0.016521 | 0.0010816 | 6.5466 | TRUE | 0.017431 | 0.0061287 | 0.028734 | pmol/µL [µM]
NIST YAA | Cer d18:1/18:0 | Single Point | 16 | Orbitrap | 18 | 0.014533 | 0.0018147 | 12.4867 | TRUE | 0.017431 | 0.0061287 | 0.028734 | pmol/µL [µM]
NIST YAA | Cer d18:1/18:0 | Single Point | 17a | TOF | 6 | 0.021387 | 0.0023077 | 10.7899 | TRUE | 0.017431 | 0.0061287 | 0.028734 | pmol/µL [µM]
NIST YAA | Cer d18:1/18:0 | Single Point | 18a | QQQ | 18 | 0.017283 | 0.00051595 | 2.9853 | TRUE | 0.017431 | 0.0061287 | 0.028734 | pmol/µL [µM]
NIST YAA | Cer d18:1/18:0 | Single Point | 18b | QQQ | 18 | 0.013315 | 0.00068524 | 5.1463 | TRUE | 0.017431 | 0.0061287 | 0.028734 | pmol/µL [µM]
NIST YAA | Cer d18:1/18:0 | Single Point | 19 | QQQ | 6 | 0.015425 | 0.0025404 | 16.4694 | TRUE | 0.017431 | 0.0061287 | 0.028734 | pmol/µL [µM]
NIST YAA | Cer d18:1/18:0 | Single Point | 22 | QQQ | 18 | 0.016003 | 0.00098751 | 6.1708 | TRUE | 0.017431 | 0.0061287 | 0.028734 | pmol/µL [µM]
NIST YAA | Cer d18:1/18:0 | Single Point | 23 | QQQ | 6 | 0.018709 | 0.0014447 | 7.7221 | TRUE | 0.017431 | 0.0061287 | 0.028734 | pmol/µL [µM]
NIST YAA | Cer d18:1/18:0 | Single Point | 24 | QQQ | 18 | 0.020484 | 0.0011973 | 5.8452 | TRUE | 0.017431 | 0.0061287 | 0.028734 | pmol/µL [µM]
NIST YAA | Cer d18:1/18:0 | Single Point | 27 | QQQ | 18 | 0.01492 | 0.0014383 | 9.6401 | TRUE | 0.017431 | 0.0061287 | 0.028734 | pmol/µL [µM]
NIST YAA | Cer d18:1/18:0 | Single Point | 28 | QQQ | 18 | 0.016984 | 0.00028912 | 1.7023 | TRUE | 0.017431 | 0.0061287 | 0.028734 | pmol/µL [µM]
NIST YAA | Cer d18:1/18:0 | Single Point | 29a | TOF | 18 | 0.014735 | 0.00050788 | 3.4468 | TRUE | 0.017431 | 0.0061287 | 0.028734 | pmol/µL [µM]
NIST YAA | Cer d18:1/18:0 | Single Point | 3 | QQQ | 6 | 0.034144 | 0.001854 | 5.4299 | FALSE | 0.017431 | 0.0061287 | 0.028734 | pmol/µL [µM]
NIST YAA | Cer d18:1/18:0 | Single Point | 30 | QQQ | 18 | 0.018199 | 0.0012913 | 7.0956 | TRUE | 0.017431 | 0.0061287 | 0.028734 | pmol/µL [µM]
NIST YAA | Cer d18:1/18:0 | Single Point | 31 | QQQ | 18 | 0.018298 | 0.00088279 | 4.8246 | TRUE | 0.017431 | 0.0061287 | 0.028734 | pmol/µL [µM]
NIST YAA | Cer d18:1/18:0 | Single Point | 32 | TOF | 18 | 0.033587 | 0.0085705 | 25.5177 | FALSE | 0.017431 | 0.0061287 | 0.028734 | pmol/µL [µM]
NIST YAA | Cer d18:1/18:0 | Single Point | 35 | QQQ | 18 | 0.014392 | 0.0018711 | 13.0009 | TRUE | 0.017431 | 0.0061287 | 0.028734 | pmol/µL [µM]
NIST YAA | Cer d18:1/18:0 | Single Point | 36 | QQQ | 18 | 0.015674 | 0.00047334 | 3.02 | TRUE | 0.017431 | 0.0061287 | 0.028734 | pmol/µL [µM]
NIST YAA | Cer d18:1/18:0 | Single Point | 37 | QQQ | 18 | 0.016876 | 0.0050663 | 30.0211 | TRUE | 0.017431 | 0.0061287 | 0.028734 | pmol/µL [µM]
NIST YAA | Cer d18:1/18:0 | Single Point | 38 | QQQ | 18 | 0.0031542 | 0.0010163 | 32.2196 | FALSE | 0.017431 | 0.0061287 | 0.028734 | pmol/µL [µM]
NIST YAA | Cer d18:1/18:0 | Single Point | 4 | QQQ | 18 | 0.014502 | 0.00044997 | 3.1028 | TRUE | 0.017431 | 0.0061287 | 0.028734 | pmol/µL [µM]
NIST YAA | Cer d18:1/18:0 | Single Point | 7 | Orbitrap | 18 | 0.016292 | 0.0019367 | 11.8873 | TRUE | 0.017431 | 0.0061287 | 0.028734 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:0 | Calibration Curve | 02a | QQQ | 6 | 0.92193 | 0.027389 | 2.9709 | TRUE | 0.95382 | -0.35492 | 2.2626 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:0 | Calibration Curve | 02b | QQQ | 18 | 1.04 | 0.035165 | 3.3812 | TRUE | 0.95382 | -0.35492 | 2.2626 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:0 | Calibration Curve | 10a | QQQ | 6 | 0.52831 | 0.088123 | 16.68 | TRUE | 0.95382 | -0.35492 | 2.2626 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:0 | Calibration Curve | 10b | QQQ | 6 | 0.57015 | 0.059371 | 10.4133 | TRUE | 0.95382 | -0.35492 | 2.2626 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:0 | Calibration Curve | 12 | QQQ | 18 | 0.93621 | 0.038444 | 4.1064 | TRUE | 0.95382 | -0.35492 | 2.2626 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:0 | Calibration Curve | 14 | QQQ | 18 | 0.93341 | 0.038723 | 4.1485 | TRUE | 0.95382 | -0.35492 | 2.2626 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:0 | Calibration Curve | 15 | QQQ | 18 | 1.0774 | 0.088547 | 8.2187 | TRUE | 0.95382 | -0.35492 | 2.2626 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:0 | Calibration Curve | 16 | Orbitrap | 18 | 0.76905 | 0.041147 | 5.3503 | TRUE | 0.95382 | -0.35492 | 2.2626 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:0 | Calibration Curve | 17a | TOF | 6 | 0.69386 | 0.010888 | 1.5693 | TRUE | 0.95382 | -0.35492 | 2.2626 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:0 | Calibration Curve | 18a | QQQ | 18 | 0.77492 | 0.018485 | 2.3854 | TRUE | 0.95382 | -0.35492 | 2.2626 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:0 | Calibration Curve | 18b | QQQ | 18 | 0.92337 | 0.015508 | 1.6795 | TRUE | 0.95382 | -0.35492 | 2.2626 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:0 | Calibration Curve | 19 | QQQ | 6 | 0.76327 | 0.065155 | 8.5363 | TRUE | 0.95382 | -0.35492 | 2.2626 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:0 | Calibration Curve | 22 | QQQ | 18 | 0.84319 | 0.061168 | 7.2544 | TRUE | 0.95382 | -0.35492 | 2.2626 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:0 | Calibration Curve | 23 | QQQ | 6 | 0.85922 | 0.020104 | 2.3399 | TRUE | 0.95382 | -0.35492 | 2.2626 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:0 | Calibration Curve | 24 | QQQ | 18 | 0.8422 | 0.071975 | 8.5461 | TRUE | 0.95382 | -0.35492 | 2.2626 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:0 | Calibration Curve | 27 | QQQ | 18 | 0.66979 | 0.074678 | 11.1494 | TRUE | 0.95382 | -0.35492 | 2.2626 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:0 | Calibration Curve | 28 | QQQ | 18 | 0.98014 | 0.015366 | 1.5677 | TRUE | 0.95382 | -0.35492 | 2.2626 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:0 | Calibration Curve | 29a | TOF | 18 | 0.805 | 0.017021 | 2.1144 | TRUE | 0.95382 | -0.35492 | 2.2626 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:0 | Calibration Curve | 3 | QQQ | 6 | 2.0248 | 0.099961 | 4.9368 | TRUE | 0.95382 | -0.35492 | 2.2626 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:0 | Calibration Curve | 30 | QQQ | 18 | 0.82182 | 0.02354 | 2.8644 | TRUE | 0.95382 | -0.35492 | 2.2626 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:0 | Calibration Curve | 31 | QQQ | 18 | 0.83171 | 0.020058 | 2.4117 | TRUE | 0.95382 | -0.35492 | 2.2626 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:0 | Calibration Curve | 32 | TOF | 18 | 0.79228 | 0.08607 | 10.8636 | TRUE | 0.95382 | -0.35492 | 2.2626 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:0 | Calibration Curve | 35 | QQQ | 18 | 0.66448 | 0.075616 | 11.3797 | TRUE | 0.95382 | -0.35492 | 2.2626 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:0 | Calibration Curve | 36 | QQQ | 18 | 0.85796 | 0.031435 | 3.6639 | TRUE | 0.95382 | -0.35492 | 2.2626 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:0 | Calibration Curve | 37 | QQQ | 18 | 3.9515 | 0.24289 | 6.1469 | FALSE | 0.95382 | -0.35492 | 2.2626 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:0 | Calibration Curve | 38 | QQQ | 18 | 0.19298 | 0.073383 | 38.0268 | TRUE | 0.95382 | -0.35492 | 2.2626 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:0 | Calibration Curve | 4 | QQQ | 18 | 0.84171 | 0.022682 | 2.6948 | TRUE | 0.95382 | -0.35492 | 2.2626 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:0 | Calibration Curve | 7 | Orbitrap | 18 | 0.7964 | 0.093421 | 11.7305 | TRUE | 0.95382 | -0.35492 | 2.2626 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:0 | Single Point | 02a | QQQ | 6 | 0.90007 | 0.026251 | 2.9166 | TRUE | 0.88399 | 0.18568 | 1.5823 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:0 | Single Point | 02b | QQQ | 18 | 1.0161 | 0.033699 | 3.3166 | TRUE | 0.88399 | 0.18568 | 1.5823 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:0 | Single Point | 10a | QQQ | 6 | 0.75316 | 0.10474 | 13.9072 | TRUE | 0.88399 | 0.18568 | 1.5823 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:0 | Single Point | 10b | QQQ | 6 | 0.79645 | 0.073519 | 9.2308 | TRUE | 0.88399 | 0.18568 | 1.5823 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:0 | Single Point | 12 | QQQ | 18 | 0.72813 | 0.029959 | 4.1145 | TRUE | 0.88399 | 0.18568 | 1.5823 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:0 | Single Point | 14 | QQQ | 18 | 0.98124 | 0.039686 | 4.0445 | TRUE | 0.88399 | 0.18568 | 1.5823 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:0 | Single Point | 15 | QQQ | 18 | 0.82107 | 0.066314 | 8.0764 | TRUE | 0.88399 | 0.18568 | 1.5823 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:0 | Single Point | 16 | Orbitrap | 18 | 0.84267 | 0.045965 | 5.4547 | TRUE | 0.88399 | 0.18568 | 1.5823 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:0 | Single Point | 17a | TOF | 6 | 0.85253 | 0.013939 | 1.635 | TRUE | 0.88399 | 0.18568 | 1.5823 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:0 | Single Point | 18a | QQQ | 18 | 0.86238 | 0.020527 | 2.3803 | TRUE | 0.88399 | 0.18568 | 1.5823 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:0 | Single Point | 18b | QQQ | 18 | 0.91858 | 0.01469 | 1.5992 | TRUE | 0.88399 | 0.18568 | 1.5823 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:0 | Single Point | 19 | QQQ | 6 | 0.85634 | 0.072155 | 8.426 | TRUE | 0.88399 | 0.18568 | 1.5823 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:0 | Single Point | 22 | QQQ | 18 | 0.90084 | 0.065586 | 7.2806 | TRUE | 0.88399 | 0.18568 | 1.5823 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:0 | Single Point | 23 | QQQ | 6 | 0.96438 | 0.02374 | 2.4617 | TRUE | 0.88399 | 0.18568 | 1.5823 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:0 | Single Point | 24 | QQQ | 18 | 0.2765 | 0.023325 | 8.4355 | TRUE | 0.88399 | 0.18568 | 1.5823 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:0 | Single Point | 27 | QQQ | 18 | 0.80677 | 0.081653 | 10.121 | TRUE | 0.88399 | 0.18568 | 1.5823 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:0 | Single Point | 28 | QQQ | 18 | 0.88521 | 0.012924 | 1.46 | TRUE | 0.88399 | 0.18568 | 1.5823 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:0 | Single Point | 29a | TOF | 18 | 0.87117 | 0.018151 | 2.0835 | TRUE | 0.88399 | 0.18568 | 1.5823 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:0 | Single Point | 3 | QQQ | 6 | 2.303 | 0.11324 | 4.9169 | FALSE | 0.88399 | 0.18568 | 1.5823 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:0 | Single Point | 30 | QQQ | 18 | 0.81148 | 0.023449 | 2.8897 | TRUE | 0.88399 | 0.18568 | 1.5823 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:0 | Single Point | 31 | QQQ | 18 | 1.361 | 0.030952 | 2.2743 | TRUE | 0.88399 | 0.18568 | 1.5823 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:0 | Single Point | 32 | TOF | 18 | 0.93786 | 0.10275 | 10.9559 | TRUE | 0.88399 | 0.18568 | 1.5823 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:0 | Single Point | 35 | QQQ | 18 | 0.75198 | 0.075158 | 9.9948 | TRUE | 0.88399 | 0.18568 | 1.5823 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:0 | Single Point | 36 | QQQ | 18 | 0.90829 | 0.032975 | 3.6305 | TRUE | 0.88399 | 0.18568 | 1.5823 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:0 | Single Point | 37 | QQQ | 18 | 0.66286 | 0.040737 | 6.1457 | TRUE | 0.88399 | 0.18568 | 1.5823 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:0 | Single Point | 38 | QQQ | 18 | 0.17353 | 0.060787 | 35.0294 | FALSE | 0.88399 | 0.18568 | 1.5823 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:0 | Single Point | 4 | QQQ | 18 | 0.93933 | 0.02539 | 2.703 | TRUE | 0.88399 | 0.18568 | 1.5823 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:0 | Single Point | 7 | Orbitrap | 18 | 0.86884 | 0.096976 | 11.1616 | TRUE | 0.88399 | 0.18568 | 1.5823 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:1 | Calibration Curve | 02a | QQQ | 6 | 0.30084 | 0.0057135 | 1.8992 | TRUE | 0.27447 | 0.049912 | 0.49902 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:1 | Calibration Curve | 02b | QQQ | 18 | 0.30189 | 0.0066915 | 2.2165 | TRUE | 0.27447 | 0.049912 | 0.49902 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:1 | Calibration Curve | 10a | QQQ | 6 | 0.19504 | 0.047642 | 24.4269 | TRUE | 0.27447 | 0.049912 | 0.49902 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:1 | Calibration Curve | 10b | QQQ | 6 | 0.13006 | 0.025858 | 19.8813 | TRUE | 0.27447 | 0.049912 | 0.49902 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:1 | Calibration Curve | 12 | QQQ | 18 | 0.3128 | 0.013978 | 4.4687 | TRUE | 0.27447 | 0.049912 | 0.49902 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:1 | Calibration Curve | 14 | QQQ | 18 | 0.28609 | 0.013835 | 4.8359 | TRUE | 0.27447 | 0.049912 | 0.49902 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:1 | Calibration Curve | 15 | QQQ | 18 | 0.34474 | 0.027634 | 8.0159 | TRUE | 0.27447 | 0.049912 | 0.49902 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:1 | Calibration Curve | 16 | Orbitrap | 18 | 0.29411 | 0.096698 | 32.8785 | TRUE | 0.27447 | 0.049912 | 0.49902 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:1 | Calibration Curve | 17a | TOF | 6 | 0.35972 | 0.017112 | 4.7569 | TRUE | 0.27447 | 0.049912 | 0.49902 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:1 | Calibration Curve | 18a | QQQ | 18 | 0.26296 | 0.0047327 | 1.7998 | TRUE | 0.27447 | 0.049912 | 0.49902 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:1 | Calibration Curve | 18b | QQQ | 18 | 0.26771 | 0.0097756 | 3.6516 | TRUE | 0.27447 | 0.049912 | 0.49902 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:1 | Calibration Curve | 19 | QQQ | 6 | 0.27891 | 0.024101 | 8.6412 | TRUE | 0.27447 | 0.049912 | 0.49902 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:1 | Calibration Curve | 22 | QQQ | 18 | 0.27301 | 0.01725 | 6.3184 | TRUE | 0.27447 | 0.049912 | 0.49902 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:1 | Calibration Curve | 23 | QQQ | 6 | 0.31846 | 0.013613 | 4.2746 | TRUE | 0.27447 | 0.049912 | 0.49902 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:1 | Calibration Curve | 24 | QQQ | 18 | 0.26753 | 0.023806 | 8.8985 | TRUE | 0.27447 | 0.049912 | 0.49902 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:1 | Calibration Curve | 27 | QQQ | 18 | 0.17348 | 0.026212 | 15.1097 | TRUE | 0.27447 | 0.049912 | 0.49902 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:1 | Calibration Curve | 28 | QQQ | 18 | 0.2924 | 0.0056277 | 1.9247 | TRUE | 0.27447 | 0.049912 | 0.49902 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:1 | Calibration Curve | 29a | TOF | 18 | 0.25968 | 0.0083507 | 3.2158 | TRUE | 0.27447 | 0.049912 | 0.49902 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:1 | Calibration Curve | 3 | QQQ | 6 | 0.69209 | 0.037788 | 5.4599 | FALSE | 0.27447 | 0.049912 | 0.49902 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:1 | Calibration Curve | 30 | QQQ | 18 | 0.31316 | 0.0072644 | 2.3197 | TRUE | 0.27447 | 0.049912 | 0.49902 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:1 | Calibration Curve | 31 | QQQ | 18 | 0.24973 | 0.007208 | 2.8863 | TRUE | 0.27447 | 0.049912 | 0.49902 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:1 | Calibration Curve | 32 | TOF | 18 | 0.40626 | 0.038717 | 9.5301 | TRUE | 0.27447 | 0.049912 | 0.49902 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:1 | Calibration Curve | 35 | QQQ | 18 | 0.15955 | 0.024679 | 15.4682 | TRUE | 0.27447 | 0.049912 | 0.49902 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:1 | Calibration Curve | 36 | QQQ | 18 | 0.10748 | 0.0090816 | 8.4495 | TRUE | 0.27447 | 0.049912 | 0.49902 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:1 | Calibration Curve | 37 | QQQ | 18 | 0.24418 | 0.024703 | 10.1166 | TRUE | 0.27447 | 0.049912 | 0.49902 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:1 | Calibration Curve | 38 | QQQ | 18 | 0.053843 | 0.024635 | 45.7534 | TRUE | 0.27447 | 0.049912 | 0.49902 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:1 | Calibration Curve | 4 | QQQ | 18 | 0.27323 | 0.0125 | 4.5749 | TRUE | 0.27447 | 0.049912 | 0.49902 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:1 | Calibration Curve | 7 | Orbitrap | 18 | 0.26615 | 0.027567 | 10.3575 | TRUE | 0.27447 | 0.049912 | 0.49902 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:1 | Single Point | 02a | QQQ | 6 | 0.27903 | 0.0051332 | 1.8397 | TRUE | 0.27713 | 0.044229 | 0.51003 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:1 | Single Point | 02b | QQQ | 18 | 0.28238 | 0.006075 | 2.1514 | TRUE | 0.27713 | 0.044229 | 0.51003 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:1 | Single Point | 10a | QQQ | 6 | 0.19458 | 0.034928 | 17.95 | TRUE | 0.27713 | 0.044229 | 0.51003 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:1 | Single Point | 10b | QQQ | 6 | 0.23389 | 0.030985 | 13.2475 | TRUE | 0.27713 | 0.044229 | 0.51003 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:1 | Single Point | 12 | QQQ | 18 | 0.22754 | 0.010252 | 4.5056 | TRUE | 0.27713 | 0.044229 | 0.51003 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:1 | Single Point | 14 | QQQ | 18 | 0.29763 | 0.013573 | 4.5603 | TRUE | 0.27713 | 0.044229 | 0.51003 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:1 | Single Point | 15 | QQQ | 18 | 0.24971 | 0.019277 | 7.7197 | TRUE | 0.27713 | 0.044229 | 0.51003 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:1 | Single Point | 16 | Orbitrap | 18 | 0.1559 | 0.058276 | 37.3808 | TRUE | 0.27713 | 0.044229 | 0.51003 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:1 | Single Point | 17a | TOF | 6 | 0.44829 | 0.023219 | 5.1794 | TRUE | 0.27713 | 0.044229 | 0.51003 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:1 | Single Point | 18a | QQQ | 18 | 0.30297 | 0.0054417 | 1.7961 | TRUE | 0.27713 | 0.044229 | 0.51003 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:1 | Single Point | 18b | QQQ | 18 | 0.30007 | 0.0097124 | 3.2367 | TRUE | 0.27713 | 0.044229 | 0.51003 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:1 | Single Point | 19 | QQQ | 6 | 0.24941 | 0.022014 | 8.8264 | TRUE | 0.27713 | 0.044229 | 0.51003 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:1 | Single Point | 22 | QQQ | 18 | 0.26999 | 0.017235 | 6.3835 | TRUE | 0.27713 | 0.044229 | 0.51003 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:1 | Single Point | 23 | QQQ | 6 | 0.30734 | 0.015327 | 4.987 | TRUE | 0.27713 | 0.044229 | 0.51003 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:1 | Single Point | 24 | QQQ | 18 | 0.079377 | 0.0068577 | 8.6394 | TRUE | 0.27713 | 0.044229 | 0.51003 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:1 | Single Point | 27 | QQQ | 18 | 0.24038 | 0.027135 | 11.2887 | TRUE | 0.27713 | 0.044229 | 0.51003 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:1 | Single Point | 28 | QQQ | 18 | 0.27221 | 0.0045364 | 1.6665 | TRUE | 0.27713 | 0.044229 | 0.51003 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:1 | Single Point | 29a | TOF | 18 | 0.26943 | 0.0085389 | 3.1692 | TRUE | 0.27713 | 0.044229 | 0.51003 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:1 | Single Point | 3 | QQQ | 6 | 0.68565 | 0.037844 | 5.5194 | FALSE | 0.27713 | 0.044229 | 0.51003 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:1 | Single Point | 30 | QQQ | 18 | 0.3177 | 0.0075932 | 2.39 | TRUE | 0.27713 | 0.044229 | 0.51003 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:1 | Single Point | 31 | QQQ | 18 | 0.31724 | 0.0083201 | 2.6226 | TRUE | 0.27713 | 0.044229 | 0.51003 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:1 | Single Point | 32 | TOF | 18 | 0.46191 | 0.044462 | 9.6257 | TRUE | 0.27713 | 0.044229 | 0.51003 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:1 | Single Point | 35 | QQQ | 18 | 0.23035 | 0.022511 | 9.7721 | TRUE | 0.27713 | 0.044229 | 0.51003 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:1 | Single Point | 36 | QQQ | 18 | 0.27269 | 0.0073052 | 2.6789 | TRUE | 0.27713 | 0.044229 | 0.51003 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:1 | Single Point | 37 | QQQ | 18 | 0.1806 | 0.017935 | 9.9305 | TRUE | 0.27713 | 0.044229 | 0.51003 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:1 | Single Point | 38 | QQQ | 18 | 0.053737 | 0.018453 | 34.3388 | TRUE | 0.27713 | 0.044229 | 0.51003 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:1 | Single Point | 4 | QQQ | 18 | 0.27763 | 0.013047 | 4.6993 | TRUE | 0.27713 | 0.044229 | 0.51003 | pmol/µL [µM]
NIST YAA | Cer d18:1/24:1 | Single Point | 7 | Orbitrap | 18 | 0.30206 | 0.027597 | 9.1361 | TRUE | 0.27713 | 0.044229 | 0.51003 | pmol/µL [µM]