## Participant Information
Participant Information
Participant Information |  | Group Name
|  | Lab code | 4
| Institution | Institution name
Street
City
Postal code
Country
| Contact person | Name
Address
Email
Phone

## Summary
Summary
Summary | This protocol describes the extraction and LC-MRM procedure for absolute quantification of Cer d18:1/16:0, Cer d18:1/18:0, Cer d18:1/24:0 and Cer d18:1/24:1 in human plasma according to Kauhanen et al, Anal Bioanal Chem (2016) 408:3475–3483.
A calibration curve is built by injecting several dilutions of the non-labelled standards mixed with a fixed concentration of labelled standards. Each calibration point is prepared and analysed in triplicates (n=3).
To establish the endogenous concentration of the four selected ceramides in the study samples (NIST SRM1950, NIST high TAG, NIST T1D, and NIST young A-A) each sample is extracted and analyzed in six replicates (n =6).
Prepare the pooled QC by mixing 100 mL each of NIST SRM1950, NIST high TAG, NIST T1D, and NIST young AA (total 400 mL).
We suggest proceeding first to measure the ceramides concentration in the four study samples and then continue with the intra-assay validation to avoid wasting material.

## Extraction - Overview
Extraction Overview
Extraction Overview | Table 1 - Extraction - Materials
Chemicals, solvents and materials
Ethyl acetate (EtAc)
2-propanol (IPA)
5% Bovine Serum Albumin (BSA)
96-well plates/tubes
Sealing foil
2 mL DeepWell plates/tubes/LC vials
Table 2 - Extraction - Sample Types
| Sample type | Components
| Pooled Quality Control samples (QC) | 10 µL pooled QC |  |  | 20 µL Labelled IS |  | 570 µL EtAc:IPA 2:8 (v/v)
| Matrix Blank (MB) | 10 µL Matrix (pooled QC) only |  |  | - |  | 590 µL EtAc:IPA 2:8
| Study Samples (S) | 10 µL Human plasma (NIST SRM1950, high TAG, T1D, young AA) |  |  | 20 µL Labelled IS |  | 570 µL EtAc:IPA 2:8
| Total Blank (TB) | 10 µL 5% BSA |  |  | - |  | 590 µL EtAc:IPA 2:8
| Calibration line standards (STD) | 10 µL 5% BSA + 20 µL non labelled STD mixture (Table 5) |  |  | 20 µL Labelled IS |  | 550 µL EtAc:IPA 2:8
Table 3 - Extraction - Labelled Internal Standard Mixture Solution (provided)
| Labelled IS mixture in EtAc:IPA 2:8 (v/v) | Concentration |  |  |  |  | Unit | Vendor
| D7-Cer d18:1/16:0 | 0.125 |  |  |  |  | pmol/µL [µM] | Avanti Polar Lipids
| D7-Cer d18:1/18:0 | 0.050
| D7-Cer d18:1/24:0 | 1500
| D7-Cer d18:1/24:1 | 0.500
Table 4 - Extraction - Non-Labelled Standards - for STD1 Mixture Solution (provided)
| Non-labelled standards in EtAc:IPA 2:8 (v/v) | Concentration |  |  |  |  | Unit | Vendor
| Cer d18:1/16:0 | 2.0 |  |  |  |  | pmol/µL [µM] | Avanti Polar Lipids
| Cer d18:1/18:0 | 2.0
| Cer d18:1/24:0 | 20.0
| Cer d18:1/24:1 | 20.0
Table 5 - Serial dilutions of Non-Labelled Standards from STD1 Mixture Solution for calibration line
| Add 20 µL of | STD1 Conc. | STD2 Conc. | STD3 Conc. | STD4 Conc. | STD5 Conc. | STD6 Conc. | Unit
| Cer d18:1/16:0 | 2.0 | 1.0 | 0.10 | 0.02 | 0.01 | 0.008 | pmol/µL [µM]
| Cer d18:1/18:0 | 2.0 | 1.0 | 0.10 | 0.02 | 0.01 | 0.008
| Cer d18:1/24:0 | 20.0 | 10.0 | 1.0 | 0.2 | 0.1 | 0.08
| Cer d18:1/24:1 | 20.0 | 10.0 | 1.0 | 0.2 | 0.1 | 0.08

## Extraction - Procedure
Extraction Procedure
Extraction Procedure | Extraction Procedure
| Step | Description | Reference Table
| 1 | Prepare different dilutions of STD1 by diluting the provided mixture (Extraction - Overview, Table 4) with EtAc:IPA 2:8 as suggested in table | Extraction - Overview'!B34
| 2 | Thaw plasma on ice (+4°C) and bring to room temperature prior to extraction. Mix by pipetting before use.
| 3 | Prepare samples in plates/tubes manually or with a robot, according to specifications in column C of table | Extraction - Overview'!C14
| 4 | Add 20 µL of labelled internal standard mixture to samples according to column F of table | Extraction - Overview'!F14
| 5 | For the calibration line, add 20 µL of different dilutions of the non-labelled mixture (STD1-6).
| 6 | Add 590 µL of ethyl acetate:isopropanol (2:8, vol/vol) to Total Blank (TB) and Matrix Blank (MB); 550 µL to calibration line (STD); 570 µL to study samples and QCs according to column H of table | Extraction - Overview'!H14
| 7 | Mix samples 10 min by pipetting or vortexing
| 8 | Centrifuge samples for 10 min at 3000 x gat room temperature
| 9 | Transfer 50 µL of the clear supernatant to MS-vials or MS plates, close/seal, store samples at -20 °C prior to LC-MRM analysis.

## LC-MRM - Overview
LC-MRM Overview
LC-MRM Overview | Table 6 Chemicals and solvents
Ammonium formate
Formic acid
Ultra-pure water
Acetonitrile
Table 7 LC-MRM LC-MS Details
| Property | Specifications
| Mass spectrometry | QQQ
| LC | UPLC
| Injection volume (indicative) | 5 µL
| Weak wash | Water
| Strong wash | Acetonitrile
| Pre-column | Waters Acquity BEH C18, 1.7 µm VanGuard Pre-Column
| Column | Waters Acquity BEH C18, 2.1 × 50 mm id. 1.7 µm
| Column temperature | 60 °C
| Mobile phase A | 10 mM ammonium acetate in water with 0.1% formic acid
| Mobile phase B | 10 mM ammonium acetate in acetonitrile:2-propanol (4:3, vol/vol) with 0.1% formic acid
Table 8 LC-MRM LC gradient
| Time (minutes) | A% | B% |  | Flow (µL/min)
| 0.00 | 15 | 85 |  | 500
| 0.50 | 15 | 85 |  | 500
| 1.50 | 0 | 100 |  | 500
| 4.00 | 0 | 100 |  | 500
| 4.10 | 15 | 85 |  | 500
| 5.00 | 15 | 85 |  | 500
Table 9 LC-MRM MS conditions (EXAMPLE: this only applies to Sciex QTRAP, please adapt)
| Polarity mode | Positive
| Ion spray | 5000 V
| Curtain gas | 25 psi
| Source temperature | 300 °C
| Gas 1 | 50 psi
| Gas 2 | 35 psi
| DP | 30 V
| EP | 10 V
| Collision cell exit potential | 20 V
| Collision energy (CE) | 40 eV
Table 10 LC-MRM Mass transitions used for absolute quantification. On different systems, RT and CE might be different!
| Analyte | Mass transitions | IS | Mass transitions | RT (min) | CE
| Cer d18:1/16:0 | 538.5 – 264.3 | D7- Cer d18:1/16:0 | 545.6 - 271.3 | 1.9-2.0 | 40
C34H68NO3 – C18H34N
| Cer d18:1/18:0 | 566.6 – 264.3 | D7- Cer d18:1/18:0 | 573.6 – 271.3 | 2.1-2.2 | 40
C36H72NO3 – C18H34N
| Cer d18:1/24:0 | 650.6 – 264.3 | D7- Cer d18:1/24:0 | 657.7 – 271.3 | 2.4-2.5 | 40
C42H84NO3 – C18H34N
| Cer d18:1/24:1 | 648.6 – 264.3 | D7- Cer d18:1/24:1 | 655.7 – 271.3 | 2.3-2.4 | 40
C42H82NO3 – C18H34N
Table 11 Data processing
| Processing software | Sciex OS 1.4

## LC-MRM - Sample Sequence
Sample Sequence (Example)
Sample Sequence (Example) | Sample Sequence 1
| Number | Sample Name | Remarks
| 1 | Total Blank 1
| 2 | Matrix blank
| 3 | Calibration line 1 from STD6 to 1 (n=3)
| 4 | […]
| 5 | Pooled QC
| 6 | Pooled QC
| 7 | SRM 1950-1
| 8 | SRM 1950-2
| 9 | SRM 1950-3
| 10 | SRM 1950-4
| 11 | SRM 1950-5
| 12 | SRM 1950-6
| 13 | Pooled QC
| 14 | Pooled QC
| 15 | T1D-1
| 16 | T1D-2
| 17 | T1D-3
| 18 | T1D-4
| 19 | T1D-5
| 20 | T1D-6
| 21 | Total Blank 2
| 22 | Pooled QC
| 23 | Pooled QC
| 24 | Young AA-1
| 25 | Young AA-2
| 26 | Young AA-3
| 27 | Young AA-4
| 28 | Young AA-5
| 29 | Young AA-6
| 30 | Pooled QC
| 31 | Pooled QC
| 32 | hTAG-1
| 33 | hTAG-2
| 34 | hTAG-3
| 35 | hTAG-4
| 36 | hTAG-5
| 37 | hTAG-6
| 38 | Pooled QC
| 39 | Pooled QC
| 40 | Calibration line 2 from STD1 to 6 (n=3)
| 41 | […]
| 42 | Matrix blank
| 43 | Total Blank 3

## LC-MRM - Intra Assay QC Proc
Intra Assay Procedure
Intra Assay Procedure | Quality Controls for Intra-Assay Variation
| Step | Description | Reference Table
| 1 | Prepare the Pooled NIST QC sample containing equal volumes of the four NIST reference plasma you received.
| 2 | The pooled QC representing middle quality control (MQC), low QC (LQC), and lower limit of quantification QC (LLOQ) is prepared by diluting MQC (e.g. two- and three-fold) with water.
| 3 | High QC (HQC) and upper limit of quantification (ULOQ) QC are prepared by spiking in 10 µL and 20 µL respectively of STD3 endogenous standards from table | Extraction - Overview'!B34
| 4 | Internal standards are added to all QC samples.
Intra-Assay Variation
| 6 | The precision and accuracy of the assay is determined for each of the four ceramides measured in LLOQ, LQC, MQC, HQC, and HLOQ.
| 7 | Intra-assay variation, precision, and accuracy is calculated for each of the ceramide in replicates of six (n = 6) at each QC concentration independently on three separate occasions.
| 8 | The intra-assay precision (percentage coefficient variance, %CV) and accuracy (percentage accuracy, %Accuracy) are calculated from the nominal concentrations according to the formulas reported in corresponding sheets and tables of this document.
| 9 | Recovery estimation is not necessary for the published method but only for different ones. Stability estimation is not necessary for the published method but only for different ones. Calibration line is prepared in 5% BSA.

## LC-MRM - Intra Assay QC Res
INTRA ASSAY - QC RESULTS
INTRA ASSAY - QC RESULTS | Table 12 LC-MRM Calibration Line 1 Data (Beginning). Please report the values of each of the three replicates separately for each compound and standard.
| Calibration Line 1 | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| STD 1 | 11690000 | 13160000 | 13220000 | 13120000 | 14600000 | 14300000 | 25570000 | 26660000 | 25580000 | 16110000 | 17280000 | 16330000 | 554700 | 615500 | 655300 | 538500 | 580300 | 624200 | 1652000 | 1734000 | 1727000 | 362900 | 387500 | 384600
| STD 2 | 5885000 | 6425000 | 6634000 | 6580000 | 7198000 | 7131000 | 12350000 | 12900000 | 12670000 | 7704000 | 8017000 | 7901000 | 602800 | 650800 | 617400 | 569500 | 622600 | 606200 | 1692000 | 1700000 | 1609000 | 375200 | 378200 | 363500
| STD 3 | 543900 | 644900 | 646100 | 606000 | 734900 | 726000 | 1079000 | 1218000 | 1172000 | 681400 | 766400 | 738300 | 665500 | 652900 | 654900 | 635100 | 619500 | 618900 | 1748000 | 1597000 | 1557000 | 397600 | 381200 | 353300
| STD 4 | 109800 | 127100 | 136400 | 119700 | 139000 | 146400 | 223600 | 228600 | 231900 | 135500 | 139800 | 140000 | 577700 | 643700 | 611200 | 566600 | 620600 | 585100 | 1573000 | 1593000 | 1465000 | 348400 | 363000 | 336400
| STD 5 | 53840 | 69530 | 60330 | 60940 | 72140 | 65770 | 102100 | 118700 | 106100 | 66370 | 72150 | 65910 | 573600 | 678900 | 645500 | 559500 | 657200 | 639700 | 1559000 | 1628000 | 1545000 | 356200 | 380700 | 354700
| STD 6 | 46150 | 56890 | 52160 | 53450 | 62280 | 58740 | 101300 | 96020 | 84880 | 61590 | 55220 | 53200 | 563600 | 661600 | 654300 | 571600 | 634400 | 640200 | 1678000 | 1540000 | 1491000 | 391400 | 353200 | 348500
Table 13 LC-MRM Calibration Line 2 Data (End). Please report the values of each of the three replicates separately for each compound and standard.
| Calibration Line 2 | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| STD 1 | 12330000 | 13260000 | 13030000 | 13930000 | 14620000 | 14290000 | 25870000 | 26320000 | 25240000 | 16350000 | 16850000 | 16120000 | 612100 | 631000 | 648200 | 647500 | 619800 | 631600 | 1672000 | 1666000 | 1694000 | 373100 | 370300 | 385700
| STD 2 | 6416000 | 6459000 | 6394000 | 7054000 | 7189000 | 7137000 | 12770000 | 12750000 | 12390000 | 8012000 | 7954000 | 7871000 | 653000 | 658500 | 636000 | 630100 | 640400 | 626800 | 1749000 | 1659000 | 1601000 | 382000 | 383600 | 362700
| STD 3 | 610100 | 647500 | 623500 | 667400 | 726700 | 710400 | 1134000 | 1187000 | 1193000 | 702300 | 749300 | 729900 | 735500 | 672100 | 641800 | 619800 | 625100 | 585600 | 1794000 | 1565000 | 1546000 | 391900 | 358400 | 339100
| STD 4 | 125600 | 128100 | 128700 | 137400 | 140100 | 140100 | 222200 | 227000 | 233100 | 135800 | 136800 | 135100 | 624800 | 654900 | 612700 | 700200 | 632300 | 622100 | 1585000 | 1587000 | 1460000 | 352800 | 356100 | 330100
| STD 5 | 63140 | 68150 | 61210 | 65350 | 66750 | 65750 | 105500 | 112800 | 116400 | 63000 | 66470 | 65070 | 655400 | 673800 | 660600 | 621300 | 608000 | 595300 | 1597000 | 1611000 | 1557000 | 356700 | 353300 | 350400
| STD 6 | 70990 | 51780 | 54320 | 60000 | 56460 | 58740 | 90280 | 96640 | 89550 | 57490 | 57390 | 52930 | 655100 | 663800 | 657600 | 559700 | 581500 | 610200 | 1550000 | 1524000 | 1548000 | 350700 | 341800 | 339400
Table 14 LC-MRM Sample Data. Please report the values of each of the three replicates separately for each compound and standard.
| NIST SRM | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| SRM 1 | 760300 | 755500 | 748100 | 301500 | 295100 | 297100 | 1427000 | 1415000 | 1381000 | 342100 | 323500 | 327600 | 633700 | 623900 | 592400 | 613500 | 610100 | 584000 | 1631000 | 1584000 | 1555000 | 378200 | 363500 | 358800
| SRM 2 | 814700 | 844100 | 824700 | 325200 | 321700 | 308700 | 1569000 | 1541000 | 1516000 | 370500 | 361300 | 350300 | 593700 | 619500 | 608000 | 619200 | 604500 | 587200 | 1644000 | 1579000 | 1575000 | 387200 | 371300 | 374400
| SRM 3 | 824500 | 817400 | 791100 | 321400 | 311200 | 311300 | 1588000 | 1581000 | 1533000 | 363800 | 377100 | 365600 | 600400 | 617100 | 581600 | 597500 | 606100 | 560100 | 1601000 | 1600000 | 1557000 | 371200 | 381800 | 367900
| SRM 4 | 847700 | 829800 | 817500 | 333900 | 328700 | 319800 | 1625000 | 1601000 | 1604000 | 382100 | 372100 | 380900 | 649800 | 641100 | 629800 | 653900 | 640500 | 617700 | 1729000 | 1725000 | 1679000 | 407400 | 416500 | 405200
| SRM 5 | 831300 | 849600 | 804400 | 333000 | 324200 | 315600 | 1600000 | 1589000 | 1565000 | 387000 | 367800 | 390100 | 642700 | 650500 | 625300 | 636200 | 632200 | 620600 | 1697000 | 1726000 | 1677000 | 408000 | 401700 | 401100
| SRM 6 | 831000 | 824300 | 805800 | 319400 | 331100 | 302700 | 1586000 | 1605000 | 1557000 | 368700 | 385400 | 374700 | 626700 | 629200 | 613500 | 642300 | 618600 | 601100 | 1689000 | 1644000 | 1633000 | 388200 | 380700 | 386200
| NIST hTAG | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| hTAG 1 | 937100 | 915000 | 885800 | 368400 | 364500 | 362100 | 2394000 | 2365000 | 2304000 | 511900 | 497800 | 462500 | 620800 | 600700 | 585500 | 620100 | 599300 | 572500 | 1668000 | 1646000 | 1599000 | 378700 | 377300 | 360600
| hTAG 2 | 968800 | 928400 | 909300 | 392200 | 373200 | 360700 | 2481000 | 2383000 | 2351000 | 524900 | 516200 | 485100 | 630300 | 605600 | 602700 | 642100 | 610200 | 604500 | 1729000 | 1692000 | 1632000 | 392400 | 386400 | 381600
| hTAG 3 | 959900 | 920000 | 924000 | 393500 | 374100 | 374900 | 2493000 | 2467000 | 2397000 | 529200 | 518900 | 510000 | 642200 | 630000 | 602200 | 639200 | 615100 | 603900 | 1731000 | 1704000 | 1711000 | 405700 | 395300 | 384800
| hTAG 4 | 978400 | 951100 | 902600 | 392700 | 382500 | 369800 | 2575000 | 2495000 | 2419000 | 540200 | 518800 | 505100 | 631700 | 616400 | 600900 | 642000 | 609900 | 592300 | 1759000 | 1733000 | 1693000 | 416500 | 394300 | 384000
| hTAG 5 | 998000 | 960200 | 907600 | 394900 | 383400 | 371800 | 2541000 | 2501000 | 2449000 | 539700 | 531500 | 513800 | 624800 | 624100 | 606700 | 654600 | 618800 | 618600 | 1814000 | 1770000 | 1711000 | 413200 | 403600 | 398400
| hTAG 6 | 962100 | 926700 | 907500 | 395100 | 371100 | 364300 | 2461000 | 2437000 | 2364000 | 530600 | 514400 | 496900 | 666900 | 631600 | 616700 | 660900 | 625300 | 613200 | 1789000 | 1731000 | 1717000 | 412700 | 412700 | 401500
| NIST T1D | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| T1D 1 | 734900 | 714800 | 701400 | 362300 | 352700 | 345900 | 1505000 | 1452000 | 1454000 | 336100 | 336400 | 328100 | 681400 | 647500 | 659300 | 665600 | 647200 | 631800 | 1762000 | 1660000 | 1683000 | 415500 | 393200 | 394800
| T1D 2 | 750500 | 754500 | 719100 | 376400 | 371400 | 355900 | 1567000 | 1520000 | 1510000 | 363100 | 355900 | 363000 | 651800 | 644300 | 617400 | 648000 | 638500 | 610600 | 1698000 | 1676000 | 1641000 | 409700 | 402700 | 389800
| T1D 3 | 727800 | 729300 | 706200 | 366200 | 356600 | 361800 | 1485000 | 1544000 | 1492000 | 358900 | 357800 | 348400 | 657400 | 642000 | 635000 | 649300 | 635700 | 606000 | 1668000 | 1731000 | 1671000 | 395400 | 401200 | 397300
| T1D 4 | 759000 | 753700 | 726700 | 373800 | 380100 | 370100 | 1609000 | 1597000 | 1553000 | 363500 | 365700 | 372400 | 649500 | 629600 | 600700 | 630700 | 610600 | 602700 | 1634000 | 1660000 | 1653000 | 391000 | 397700 | 382200
| T1D 5 | 749600 | 761100 | 721000 | 375500 | 378800 | 352900 | 1604000 | 1588000 | 1546000 | 368300 | 367400 | 362900 | 603400 | 592800 | 568300 | 600200 | 583800 | 556900 | 1565000 | 1555000 | 1534000 | 364800 | 370000 | 370300
| T1D 6 | 819600 | 796700 | 786500 | 411400 | 394000 | 388600 | 1704000 | 1707000 | 1624000 | 407400 | 386800 | 389500 | 644200 | 629000 | 610000 | 638100 | 622100 | 606600 | 1709000 | 1693000 | 1632000 | 399100 | 403000 | 393900
| NIST Young AA | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| yA-A 1 | 516300 | 499100 | 505300 | 179900 | 180400 | 173000 | 1008000 | 968700 | 954800 | 208600 | 202800 | 198300 | 664700 | 671700 | 651100 | 661100 | 623400 | 619100 | 1690000 | 1614000 | 1561000 | 400400 | 368300 | 386000
| yA-A 2 | 521500 | 522300 | 504800 | 181900 | 187500 | 185200 | 1064000 | 1019000 | 1022000 | 212100 | 217900 | 218400 | 638500 | 650200 | 623700 | 616600 | 620100 | 603900 | 1611000 | 1584000 | 1575000 | 372400 | 375200 | 380300
| yA-A 3 | 547600 | 541200 | 510900 | 188300 | 192100 | 183200 | 1090000 | 1053000 | 1057000 | 225200 | 224700 | 230500 | 683000 | 684600 | 676700 | 661000 | 644700 | 627900 | 1719000 | 1684000 | 1683000 | 403600 | 391100 | 396800
| yA-A 4 | 519400 | 508200 | 503200 | 188800 | 178400 | 177400 | 1038000 | 1035000 | 1007000 | 207700 | 220700 | 221000 | 666800 | 677600 | 651600 | 655200 | 629300 | 612100 | 1720000 | 1631000 | 1594000 | 405100 | 388700 | 378700
| yA-A 5 | 523300 | 544900 | 522600 | 189400 | 187600 | 184300 | 1072000 | 1069000 | 1058000 | 214600 | 229000 | 217400 | 687300 | 689500 | 679100 | 671800 | 653100 | 638500 | 1747000 | 1703000 | 1709000 | 412200 | 414000 | 414100
| yA-A 6 | 535600 | 526200 | 523100 | 185900 | 195300 | 184000 | 1088000 | 1078000 | 1049000 | 227400 | 221600 | 233900 | 682200 | 664300 | 657600 | 657600 | 642400 | 623500 | 1705000 | 1701000 | 1688000 | 405200 | 400200 | 394000
Table 15 LC-MRM Intra Assay QC Data. Please report the values of each of the three replicates separately for each compound and standard.
| Lowest Level QC | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| LLQC 1 | 225800 | 217500 | 234200 | 96290 | 91450 | 87510 | 560800 | 525400 | 476000 | 125100 | 113100 | 101100 | 554100 | 539600 | 563200 | 534300 | 520300 | 533800 | 1627000 | 1543000 | 1387000 | 382400 | 352600 | 303100
| LLQC 2 | 247200 | 240400 | 248700 | 98900 | 93380 | 91930 | 531200 | 514800 | 498400 | 112200 | 106200 | 110600 | 579500 | 545600 | 581800 | 542700 | 528200 | 523800 | 1480000 | 1452000 | 1415000 | 337100 | 337100 | 324500
| LLQC 3 | 263000 | 264200 | 258900 | 103000 | 96100 | 102500 | 558700 | 525300 | 515800 | 120200 | 110900 | 109000 | 619100 | 583000 | 622600 | 583400 | 557600 | 584100 | 1569000 | 1510000 | 1493000 | 361100 | 353300 | 326400
| LLQC 4 | 257200 | 248800 | 261700 | 100000 | 99570 | 101200 | 549600 | 522900 | 539600 | 118100 | 109400 | 118900 | 612300 | 588200 | 616700 | 591000 | 557200 | 582000 | 1556000 | 1524000 | 1559000 | 362400 | 347900 | 360600
| LLQC 5 | 252900 | 241500 | 255600 | 99040 | 96840 | 101400 | 514500 | 508300 | 522400 | 112500 | 103400 | 110400 | 610800 | 600300 | 621600 | 592300 | 570800 | 583500 | 1568000 | 1533000 | 1556000 | 359200 | 351900 | 362700
| LLQC 6 | 230400 | 221100 | 234400 | 96320 | 90370 | 86200 | 472500 | 446100 | 475200 | 96910 | 93570 | 98270 | 636500 | 610500 | 643400 | 592500 | 583400 | 591000 | 1584000 | 1535000 | 1606000 | 357200 | 345800 | 370600
| Low QC | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| LQC 1 | 360200 | 352600 | 358500 | 148400 | 148900 | 142100 | 770000 | 763600 | 759100 | 162700 | 161000 | 168200 | 607300 | 581800 | 608100 | 556500 | 554700 | 561300 | 1531000 | 1515000 | 1504000 | 349500 | 350700 | 351600
| LQC 2 | 386600 | 383100 | 393900 | 162600 | 153800 | 152200 | 837900 | 819500 | 818700 | 179400 | 174300 | 180700 | 622700 | 605800 | 618200 | 605500 | 571000 | 584700 | 1613000 | 1542000 | 1584000 | 370400 | 369700 | 371300
| LQC 3 | 398900 | 382000 | 400300 | 157500 | 153800 | 153500 | 852700 | 819000 | 827200 | 181700 | 172100 | 179000 | 642300 | 619000 | 637200 | 605900 | 594100 | 588600 | 1619000 | 1586000 | 1618000 | 376900 | 365300 | 375400
| LQC 4 | 384100 | 374800 | 373500 | 155200 | 151400 | 152300 | 809700 | 787000 | 821200 | 174300 | 163700 | 176100 | 632300 | 616800 | 622200 | 601700 | 585300 | 601900 | 1622000 | 1598000 | 1618000 | 369500 | 360300 | 373300
| LQC 5 | 402300 | 378300 | 385800 | 157900 | 152000 | 153800 | 846400 | 814900 | 844700 | 181700 | 170900 | 179600 | 639000 | 616900 | 627900 | 590400 | 577200 | 611400 | 1609000 | 1588000 | 1668000 | 369300 | 356600 | 378000
| LQC 6 | 387400 | 372400 | 376200 | 157400 | 154700 | 152000 | 817700 | 775200 | 802400 | 172700 | 164900 | 172700 | 608700 | 590000 | 612300 | 575600 | 559000 | 579600 | 1575000 | 1539000 | 1565000 | 365300 | 346400 | 377100
| Middle QC | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| MQC 1 | 721700 | 692800 | 708800 | 300000 | 280900 | 290800 | 1599000 | 1547000 | 1581000 | 350600 | 346400 | 341900 | 586800 | 570800 | 592900 | 556700 | 552200 | 559000 | 1574000 | 1537000 | 1564000 | 371700 | 354100 | 375400
| MQC 2 | 730800 | 705100 | 717500 | 310200 | 297800 | 292700 | 1662000 | 1579000 | 1630000 | 356500 | 338200 | 362800 | 606700 | 588400 | 588300 | 596100 | 582700 | 564800 | 1679000 | 1586000 | 1631000 | 392000 | 364400 | 379200
| MQC 3 | 688800 | 661400 | 684400 | 278600 | 265800 | 273700 | 1518000 | 1453000 | 1479000 | 338900 | 319200 | 330200 | 629400 | 605500 | 601600 | 597000 | 584000 | 587700 | 1670000 | 1565000 | 1650000 | 408800 | 376300 | 392100
| MQC 4 | 729200 | 704700 | 706800 | 298300 | 290500 | 299600 | 1628000 | 1558000 | 1630000 | 357000 | 345000 | 349400 | 595000 | 578700 | 582800 | 588700 | 563800 | 574400 | 1623000 | 1572000 | 1593000 | 377500 | 368200 | 378600
| MQC 5 | 770600 | 767500 | 770000 | 326100 | 304200 | 322900 | 1768000 | 1694000 | 1735000 | 395300 | 359200 | 367500 | 515700 | 512100 | 508900 | 514500 | 500300 | 497700 | 1435000 | 1377000 | 1414000 | 342400 | 321900 | 327500
| MQC 6 | 763400 | 721900 | 739200 | 312600 | 303500 | 316500 | 1708000 | 1638000 | 1661000 | 389100 | 360000 | 367400 | 572300 | 555100 | 561200 | 578900 | 538100 | 563200 | 1587000 | 1485000 | 1561000 | 368900 | 340300 | 369400
| High QC | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| HQC 1 | 947600 | 906100 | 916700 | 550700 | 548900 | 550500 | 2066000 | 1984000 | 1955000 | 677500 | 631300 | 610300 | 588600 | 571100 | 560700 | 561800 | 561600 | 557600 | 1607000 | 1506000 | 1524000 | 378600 | 336200 | 355500
| HQC 2 | 945500 | 922200 | 924600 | 555100 | 547500 | 551200 | 2123000 | 1933000 | 1999000 | 684600 | 620100 | 648700 | 582100 | 574500 | 572300 | 566700 | 544200 | 566200 | 1642000 | 1500000 | 1525000 | 369200 | 344000 | 359000
| HQC 3 | 938200 | 911300 | 923600 | 565700 | 543300 | 549600 | 2136000 | 1977000 | 2015000 | 683400 | 644900 | 637700 | 595100 | 562100 | 573400 | 576600 | 560800 | 559000 | 1623000 | 1524000 | 1538000 | 385800 | 348700 | 373200
| HQC 4 | 946900 | 908800 | 930100 | 557900 | 552400 | 554300 | 2107000 | 1993000 | 2009000 | 682400 | 622800 | 655300 | 586100 | 582900 | 574100 | 561300 | 556200 | 563400 | 1629000 | 1497000 | 1559000 | 387200 | 353200 | 358000
| HQC 5 | 943000 | 931000 | 924500 | 559500 | 550900 | 548700 | 2118000 | 1990000 | 2028000 | 700200 | 643400 | 648500 | 587700 | 564800 | 565800 | 584200 | 551200 | 566900 | 1650000 | 1521000 | 1588000 | 387100 | 360600 | 367000
| HQC 6 | 947000 | 911000 | 939300 | 560500 | 549500 | 561300 | 2112000 | 2021000 | 2044000 | 664100 | 639000 | 639400 | 589700 | 571400 | 591200 | 571900 | 556000 | 562400 | 1635000 | 1578000 | 1571000 | 392100 | 367000 | 368300
| Highest Level QC | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| HLQC 1 | 1153000 | 1113000 | 1124000 | 799300 | 790000 | 786100 | 2531000 | 2411000 | 2444000 | 972600 | 912700 | 932000 | 606900 | 589200 | 577900 | 573900 | 576500 | 560900 | 1627000 | 1577000 | 1617000 | 370500 | 366800 | 373500
| HLQC 2 | 1133000 | 1097000 | 1114000 | 781100 | 770300 | 782800 | 2489000 | 2349000 | 2391000 | 951200 | 929700 | 932800 | 594000 | 583400 | 589500 | 571900 | 563100 | 571800 | 1640000 | 1566000 | 1583000 | 389800 | 365400 | 378200
| HLQC 3 | 1145000 | 1085000 | 1128000 | 824700 | 767700 | 773700 | 2486000 | 2333000 | 2427000 | 983100 | 919900 | 942200 | 585300 | 573300 | 571800 | 585400 | 552400 | 555100 | 1629000 | 1534000 | 1574000 | 384900 | 358100 | 374800
| HLQC 4 | 1146000 | 1115000 | 1135000 | 801600 | 774300 | 785000 | 2476000 | 2417000 | 2463000 | 986200 | 934200 | 957600 | 610600 | 582100 | 612800 | 592900 | 563000 | 592600 | 1625000 | 1554000 | 1611000 | 383200 | 367500 | 369900
| HLQC 5 | 1140000 | 1136000 | 1112000 | 796900 | 771800 | 784500 | 2470000 | 2400000 | 2446000 | 975300 | 925800 | 949500 | 598800 | 580500 | 586700 | 584300 | 566300 | 572600 | 1660000 | 1568000 | 1593000 | 378200 | 357800 | 376800
| HLQC 6 | 1156000 | 1100000 | 1118000 | 788600 | 793200 | 790000 | 2563000 | 2413000 | 2428000 | 974900 | 915700 | 950900 | 597900 | 587800 | 580400 | 589500 | 569600 | 568300 | 1655000 | 1566000 | 1573000 | 385900 | 361900 | 368900
| Blank QC* | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| Blank | 644200 | 632900 | 666600 | 291800 | 294500 | 258300 | 2140000 | 1778000 | 910600 | 479700 | 396600 | 213900 | 741.2 | 598 | 97.66 | 55.23 | 361.5 | 769.7 | 11810 | 7966 | 5314 | 34.44 | N/A | N/A
| Blank | 665700 | 617900 | 663800 | 268900 | 264500 | 258200 | 1448000 | 1142000 | 990800 | 327600 | 257400 | 236200 | 1128 | 59.83 | 490.9 | 945.2 | 27.93 | 693.7 | 6691 | 5820 | 5131 | N/A | N/A | 33.7
| Blank | 644600 | 630800 | 638600 | 289300 | 259600 | 262700 | 1717000 | 1123000 | 952400 | 401700 | 267400 | 244800 | 703 | 178.2 | 16.27 | 234.1 | 115.1 | 494.5 | 9204 | 5279 | 5532 | 314.6 | 367.4 | 26.88
| * | Blank QC are the matrix blank without any Internal Standard
Note on F18: These appear to have been flipped in the submitted sheet.

## Raw Peak Areas
Replicate 1 (n=1)
Sample_Name | Cer d18:1/16:0 [M>264] | Cer d18:1/16:0-d7 [M>264] | Cer d18:1/18:0 [M>264] | Cer d18:1/18:0-d7 [M>264] | Cer d18:1/24:0 [M>264] | Cer d18:1/24:0-d7 [M>264] | Cer d18:1/24:1 [M>264] | Cer d18:1/24:1-d7 [M>264]
Total Blank 1 | 190.6 | 15.3 | 224.9 | 409.5 | 597 | 262.3 | 212.4 | N/A
Matrix blank | 678400 | 1500 | 311400 | 374.8 | 1920000 | 9641 | 456800 | N/A
Calibration line 1_STD6 | 46150 | 563600 | 53450 | 571600 | 101300 | 1678000 | 61590 | 391400
Calibration line 1_STD5 | 53840 | 573600 | 60940 | 559500 | 102100 | 1559000 | 66370 | 356200
Calibration line 1_STD4 | 109800 | 577700 | 119700 | 566600 | 223600 | 1573000 | 135500 | 348400
Calibration line 1_STD3 | 543900 | 665500 | 606000 | 635100 | 1079000 | 1748000 | 681400 | 397600
Calibration line 1_STD2 | 5885000 | 602800 | 6580000 | 569500 | 12350000 | 1692000 | 7704000 | 375200
Calibration line 1_STD1 | 11690000 | 554700 | 13120000 | 538500 | 25570000 | 1652000 | 16110000 | 362900
Pooled QC | 763700 | 622300 | 311100 | 626100 | 1611000 | 1638000 | 358700 | 387600
Pooled QC | 722800 | 577600 | 290900 | 568800 | 1457000 | 1444000 | 334900 | 344700
SRM 1950-1 | 760300 | 633700 | 301500 | 613500 | 1427000 | 1631000 | 342100 | 378200
SRM 1950-2 | 814700 | 593700 | 325200 | 619200 | 1569000 | 1644000 | 370500 | 387200
SRM 1950-3 | 824500 | 600400 | 321400 | 597500 | 1588000 | 1601000 | 363800 | 371200
SRM 1950-4 | 847700 | 649800 | 333900 | 653900 | 1625000 | 1729000 | 382100 | 407400
SRM 1950-5 | 831300 | 642700 | 333000 | 636200 | 1600000 | 1697000 | 387000 | 408000
SRM 1950-6 | 831000 | 626700 | 319400 | 642300 | 1586000 | 1689000 | 368700 | 388200
Pooled QC | 790900 | 641000 | 325800 | 638600 | 1555000 | 1561000 | 352700 | 371900
Pooled QC | 736900 | 588500 | 301600 | 574600 | 1430000 | 1409000 | 325500 | 323400
T1D-1 | 734900 | 681400 | 362300 | 665600 | 1505000 | 1762000 | 336100 | 415500
T1D-2 | 750500 | 651800 | 376400 | 648000 | 1567000 | 1698000 | 363100 | 409700
T1D-3 | 727800 | 657400 | 366200 | 649300 | 1485000 | 1668000 | 358900 | 395400
T1D-4 | 759000 | 649500 | 373800 | 630700 | 1609000 | 1634000 | 363500 | 391000
T1D-5 | 749600 | 603400 | 375500 | 600200 | 1604000 | 1565000 | 368300 | 364800
T1D-6 | 819600 | 644200 | 411400 | 638100 | 1704000 | 1709000 | 407400 | 399100
Total Blank 2 | 21570 | 202.2 | 4576 | 917.8 | 9568 | 630.6 | 1676 | N/A
Pooled QC | 793400 | 632100 | 315700 | 622500 | 1497000 | 1544000 | 341100 | 370700
Pooled QC | 739600 | 583100 | 307700 | 573000 | 1415000 | 1366000 | 316200 | 338300
Young AA-1 | 516300 | 664700 | 179900 | 661100 | 1008000 | 1690000 | 208600 | 400400
Young AA-2 | 521500 | 638500 | 181900 | 616600 | 1064000 | 1611000 | 212100 | 372400
Young AA-3 | 547600 | 683000 | 188300 | 661000 | 1090000 | 1719000 | 225200 | 403600
Young AA-4 | 519400 | 666800 | 188800 | 655200 | 1038000 | 1720000 | 207700 | 405100
Young AA-5 | 523300 | 687300 | 189400 | 671800 | 1072000 | 1747000 | 214600 | 412200
Young AA-6 | 535600 | 682200 | 185900 | 657600 | 1088000 | 1705000 | 227400 | 405200
Pooled QC | 785100 | 643500 | 314700 | 641600 | 1517000 | 1517000 | 345900 | 367200
Pooled QC | 743200 | 594000 | 304100 | 565400 | 1360000 | 1344000 | 313600 | 334700
hTAG-1 | 937100 | 620800 | 368400 | 620100 | 2394000 | 1668000 | 511900 | 378700
hTAG-2 | 968800 | 630300 | 392200 | 642100 | 2481000 | 1729000 | 524900 | 392400
hTAG-3 | 959900 | 642200 | 393500 | 639200 | 2493000 | 1731000 | 529200 | 405700
hTAG-4 | 978400 | 631700 | 392700 | 642000 | 2575000 | 1759000 | 540200 | 416500
hTAG-5 | 998000 | 624800 | 394900 | 654600 | 2541000 | 1814000 | 539700 | 413200
hTAG-6 | 962100 | 666900 | 395100 | 660900 | 2461000 | 1789000 | 530600 | 412700
Pooled QC | 779200 | 626200 | 326200 | 643400 | 1495000 | 1481000 | 330800 | 367600
Pooled QC | 717500 | 579700 | 303000 | 567000 | 1346000 | 1314000 | 305100 | 322700
Calibration line 2_STD6 | 70990 | 655100 | 60000 | 647500 | 90280 | 1550000 | 57490 | 350700
Calibration line 2_STD5 | 63140 | 655400 | 65350 | 630100 | 105500 | 1597000 | 63000 | 356700
Calibration line 2_STD4 | 125600 | 624800 | 137400 | 619800 | 222200 | 1585000 | 135800 | 352800
Calibration line 2_STD3 | 610100 | 735500 | 667400 | 700200 | 1134000 | 1794000 | 702300 | 391900
Calibration line 2_STD2 | 6416000 | 653000 | 7054000 | 621300 | 12770000 | 1749000 | 8012000 | 382000
Calibration line 2_STD1 | 12330000 | 612100 | 13930000 | 559700 | 25870000 | 1672000 | 16350000 | 373100
Matrix blank | 741900 | 264.2 | 306100 | 416.8 | 1552000 | 8617 | 357900 | 342.8
Total Blank 3 | 15590 | 1481 | 5041 | 1294 | 7628 | 649 | 1668 | 448.9
Replicate 2 (n=2)
Sample_Name | Cer d18:1/16:0 [M>264] | Cer d18:1/16:0-d7 [M>264] | Cer d18:1/18:0 [M>264] | Cer d18:1/18:0-d7 [M>264] | Cer d18:1/24:0 [M>264] | Cer d18:1/24:0-d7 [M>264] | Cer d18:1/24:1 [M>264] | Cer d18:1/24:1-d7 [M>264]
Total Blank 1 | 22400 | 66.47 | 4456 | 525.3 | 8408 | 387.7 | 1674 | 118.5
Matrix blank | 742400 | 295.3 | 314500 | 150.3 | 1403000 | 8435 | 323300 | 78
Calibration line 1_STD6 | 56890 | 661600 | 62280 | 634400 | 96020 | 1540000 | 55220 | 353200
Calibration line 1_STD5 | 69530 | 678900 | 72140 | 657200 | 118700 | 1628000 | 72150 | 380700
Calibration line 1_STD4 | 127100 | 643700 | 139000 | 620600 | 228600 | 1593000 | 139800 | 363000
Calibration line 1_STD3 | 644900 | 652900 | 734900 | 619500 | 1218000 | 1597000 | 766400 | 381200
Calibration line 1_STD2 | 6425000 | 650800 | 7198000 | 622600 | 12900000 | 1700000 | 8017000 | 378200
Calibration line 1_STD1 | 13160000 | 615500 | 14600000 | 580300 | 26660000 | 1734000 | 17280000 | 387500
Pooled QC | 810300 | 649000 | 319000 | 624300 | 1493000 | 1463000 | 344200 | 383600
Pooled QC | 736900 | 581800 | 292900 | 548100 | 1338000 | 1323000 | 304400 | 328900
SRM 1950-1 | 755500 | 623900 | 295100 | 610100 | 1415000 | 1584000 | 323500 | 363500
SRM 1950-2 | 844100 | 619500 | 321700 | 604500 | 1541000 | 1579000 | 361300 | 371300
SRM 1950-3 | 817400 | 617100 | 311200 | 606100 | 1581000 | 1600000 | 377100 | 381800
SRM 1950-4 | 829800 | 641100 | 328700 | 640500 | 1601000 | 1725000 | 372100 | 416500
SRM 1950-5 | 849600 | 650500 | 324200 | 632200 | 1589000 | 1726000 | 367800 | 401700
SRM 1950-6 | 824300 | 629200 | 331100 | 618600 | 1605000 | 1644000 | 385400 | 380700
Pooled QC | 786300 | 632800 | 302000 | 631000 | 1433000 | 1446000 | 330100 | 359600
Pooled QC | 733000 | 573500 | 310100 | 565600 | 1313000 | 1285000 | 301200 | 315500
T1D-1 | 714800 | 647500 | 352700 | 647200 | 1452000 | 1660000 | 336400 | 393200
T1D-2 | 754500 | 644300 | 371400 | 638500 | 1520000 | 1676000 | 355900 | 402700
T1D-3 | 729300 | 642000 | 356600 | 635700 | 1544000 | 1731000 | 357800 | 401200
T1D-4 | 753700 | 629600 | 380100 | 610600 | 1597000 | 1660000 | 365700 | 397700
T1D-5 | 761100 | 592800 | 378800 | 583800 | 1588000 | 1555000 | 367400 | 370000
T1D-6 | 796700 | 629000 | 394000 | 622100 | 1707000 | 1693000 | 386800 | 403000
Total Blank 2 | 25420 | 129.2 | 6489 | 317.6 | 10910 | 555.9 | 2497 | 58.11
Pooled QC | 793500 | 630000 | 315400 | 617300 | 1298000 | 1312000 | 301000 | 317100
Pooled QC | 705600 | 570000 | 303200 | 555300 | 1247000 | 1227000 | 283500 | 300800
Young AA-1 | 499100 | 671700 | 180400 | 623400 | 968700 | 1614000 | 202800 | 368300
Young AA-2 | 522300 | 650200 | 187500 | 620100 | 1019000 | 1584000 | 217900 | 375200
Young AA-3 | 541200 | 684600 | 192100 | 644700 | 1053000 | 1684000 | 224700 | 391100
Young AA-4 | 508200 | 677600 | 178400 | 629300 | 1035000 | 1631000 | 220700 | 388700
Young AA-5 | 544900 | 689500 | 187600 | 653100 | 1069000 | 1703000 | 229000 | 414000
Young AA-6 | 526200 | 664300 | 195300 | 642400 | 1078000 | 1701000 | 221600 | 400200
Pooled QC | 773800 | 632500 | 313000 | 604000 | 1365000 | 1382000 | 323000 | 347100
Pooled QC | 722200 | 591700 | 287900 | 553300 | 1280000 | 1268000 | 287600 | 300300
hTAG-1 | 915000 | 600700 | 364500 | 599300 | 2365000 | 1646000 | 497800 | 377300
hTAG-2 | 928400 | 605600 | 373200 | 610200 | 2383000 | 1692000 | 516200 | 386400
hTAG-3 | 920000 | 630000 | 374100 | 615100 | 2467000 | 1704000 | 518900 | 395300
hTAG-4 | 951100 | 616400 | 382500 | 609900 | 2495000 | 1733000 | 518800 | 394300
hTAG-5 | 960200 | 624100 | 383400 | 618800 | 2501000 | 1770000 | 531500 | 403600
hTAG-6 | 926700 | 631600 | 371100 | 625300 | 2437000 | 1731000 | 514400 | 412700
Pooled QC | 785500 | 623800 | 309500 | 591900 | 1378000 | 1357000 | 328100 | 344500
Pooled QC | 724800 | 575100 | 294500 | 540200 | 1277000 | 1247000 | 287800 | 304200
Calibration line 2_STD6 | 51780 | 663800 | 56460 | 619800 | 96640 | 1524000 | 57390 | 341800
Calibration line 2_STD5 | 68150 | 673800 | 66750 | 640400 | 112800 | 1611000 | 66470 | 353300
Calibration line 2_STD4 | 128100 | 654900 | 140100 | 625100 | 227000 | 1587000 | 136800 | 356100
Calibration line 2_STD3 | 647500 | 672100 | 726700 | 632300 | 1187000 | 1565000 | 749300 | 358400
Calibration line 2_STD2 | 6459000 | 658500 | 7189000 | 608000 | 12750000 | 1659000 | 7954000 | 383600
Calibration line 2_STD1 | 13260000 | 631000 | 14620000 | 581500 | 26320000 | 1666000 | 16850000 | 370300
Matrix blank | 761200 | 11.33 | 306500 | 317.9 | 1444000 | 7497 | 330600 | N/A
Total Blank 3 | 31870 | 1468 | 8207 | 964.4 | 9582 | 632.4 | 2521 | N/A
Replicate 3 (n=3)
Sample_Name | Cer d18:1/16:0 [M>264] | Cer d18:1/16:0-d7 [M>264] | Cer d18:1/18:0 [M>264] | Cer d18:1/18:0-d7 [M>264] | Cer d18:1/24:0 [M>264] | Cer d18:1/24:0-d7 [M>264] | Cer d18:1/24:1 [M>264] | Cer d18:1/24:1-d7 [M>264]
Total Blank 1 | 11460 | 801 | 4214 | 13.68 | 6443 | 192.6 | 2281 | N/A
Matrix blank | 739600 | 110.4 | 303200 | 500.7 | 1260000 | 6073 | 287700 | 36.98
Calibration line 1_STD6 | 52160 | 654300 | 58740 | 640200 | 84880 | 1491000 | 53200 | 348500
Calibration line 1_STD5 | 60330 | 645500 | 65770 | 639700 | 106100 | 1545000 | 65910 | 354700
Calibration line 1_STD4 | 136400 | 611200 | 146400 | 585100 | 231900 | 1465000 | 140000 | 336400
Calibration line 1_STD3 | 646100 | 654900 | 726000 | 618900 | 1172000 | 1557000 | 738300 | 353300
Calibration line 1_STD2 | 6634000 | 617400 | 7131000 | 606200 | 12670000 | 1609000 | 7901000 | 363500
Calibration line 1_STD1 | 13220000 | 655300 | 14300000 | 624200 | 25580000 | 1727000 | 16330000 | 384600
Pooled QC | 766300 | 627500 | 305000 | 606800 | 1363000 | 1372000 | 329100 | 343200
Pooled QC | 717900 | 579200 | 291600 | 556400 | 1264000 | 1254000 | 300800 | 310900
SRM 1950-1 | 748100 | 592400 | 297100 | 584000 | 1381000 | 1555000 | 327600 | 358800
SRM 1950-2 | 824700 | 608000 | 308700 | 587200 | 1516000 | 1575000 | 350300 | 374400
SRM 1950-3 | 791100 | 581600 | 311300 | 560100 | 1533000 | 1557000 | 365600 | 367900
SRM 1950-4 | 817500 | 629800 | 319800 | 617700 | 1604000 | 1679000 | 380900 | 405200
SRM 1950-5 | 804400 | 625300 | 315600 | 620600 | 1565000 | 1677000 | 390100 | 401100
SRM 1950-6 | 805800 | 613500 | 302700 | 601100 | 1557000 | 1633000 | 374700 | 386200
Pooled QC | 764500 | 607900 | 306700 | 591300 | 1308000 | 1319000 | 308900 | 332500
Pooled QC | 722000 | 555300 | 288700 | 539000 | 1244000 | 1182000 | 278500 | 291100
T1D-1 | 701400 | 659300 | 345900 | 631800 | 1454000 | 1683000 | 328100 | 394800
T1D-2 | 719100 | 617400 | 355900 | 610600 | 1510000 | 1641000 | 363000 | 389800
T1D-3 | 706200 | 635000 | 361800 | 606000 | 1492000 | 1671000 | 348400 | 397300
T1D-4 | 726700 | 600700 | 370100 | 602700 | 1553000 | 1653000 | 372400 | 382200
T1D-5 | 721000 | 568300 | 352900 | 556900 | 1546000 | 1534000 | 362900 | 370300
T1D-6 | 786500 | 610000 | 388600 | 606600 | 1624000 | 1632000 | 389500 | 393900
Total Blank 2 | 6741 | 119.4 | 3450 | 681.2 | 5363 | 244 | 605.5 | 200.9
Pooled QC | 764100 | 598500 | 309400 | 589900 | 1312000 | 1331000 | 307500 | 334600
Pooled QC | 716700 | 566900 | 294300 | 551300 | 1230000 | 1205000 | 280400 | 290400
Young AA-1 | 505300 | 651100 | 173000 | 619100 | 954800 | 1561000 | 198300 | 386000
Young AA-2 | 504800 | 623700 | 185200 | 603900 | 1022000 | 1575000 | 218400 | 380300
Young AA-3 | 510900 | 676700 | 183200 | 627900 | 1057000 | 1683000 | 230500 | 396800
Young AA-4 | 503200 | 651600 | 177400 | 612100 | 1007000 | 1594000 | 221000 | 378700
Young AA-5 | 522600 | 679100 | 184300 | 638500 | 1058000 | 1709000 | 217400 | 414100
Young AA-6 | 523100 | 657600 | 184000 | 623500 | 1049000 | 1688000 | 233900 | 394000
Pooled QC | 781000 | 609800 | 310900 | 582200 | 1309000 | 1321000 | 297100 | 328100
Pooled QC | 747200 | 554900 | 293100 | 549300 | 1229000 | 1183000 | 273900 | 291100
hTAG-1 | 885800 | 585500 | 362100 | 572500 | 2304000 | 1599000 | 462500 | 360600
hTAG-2 | 909300 | 602700 | 360700 | 604500 | 2351000 | 1632000 | 485100 | 381600
hTAG-3 | 924000 | 602200 | 374900 | 603900 | 2397000 | 1711000 | 510000 | 384800
hTAG-4 | 902600 | 600900 | 369800 | 592300 | 2419000 | 1693000 | 505100 | 384000
hTAG-5 | 907600 | 606700 | 371800 | 618600 | 2449000 | 1711000 | 513800 | 398400
hTAG-6 | 907500 | 616700 | 364300 | 613200 | 2364000 | 1717000 | 496900 | 401500
Pooled QC | 771200 | 593600 | 303700 | 560600 | 1299000 | 1258000 | 300400 | 308900
Pooled QC | 745200 | 558700 | 288200 | 528900 | 1203000 | 1155000 | 263100 | 284100
Calibration line 2_STD6 | 54320 | 657600 | 58740 | 631600 | 89550 | 1548000 | 52930 | 339400
Calibration line 2_STD5 | 61210 | 660600 | 65750 | 626800 | 116400 | 1557000 | 65070 | 350400
Calibration line 2_STD4 | 128700 | 612700 | 140100 | 585600 | 233100 | 1460000 | 135100 | 330100
Calibration line 2_STD3 | 623500 | 641800 | 710400 | 622100 | 1193000 | 1546000 | 729900 | 339100
Calibration line 2_STD2 | 6394000 | 636000 | 7137000 | 595300 | 12390000 | 1601000 | 7871000 | 362700
Calibration line 2_STD1 | 13030000 | 648200 | 14290000 | 610200 | 25240000 | 1694000 | 16120000 | 385700
Matrix blank | 708500 | 313.6 | 300600 | 541.2 | 1341000 | 6938 | 304600 | 20.93
Total Blank 3 | 3255 | 87.46 | 753.5 | 42.63 | 7510 | 293.9 | 88.89 | 292