## Participant Information
Participant Information
Participant Information |  | Group Name
|  | Lab code | #36
| Institution | Institution name
Street
City
Postal code
Country
| Contact person | Name
Address
Email
Phone

## Summary
Summary
Summary | This protocol describes the extraction and LC-MRM procedure for absolute quantification of Cer d18:1/16:0, Cer d18:1/18:0, Cer d18:1/24:0 and Cer d18:1/24:1 in human plasma according to Kauhanen et al, Anal Bioanal Chem (2016) 408:3475–3483.
A calibration curve is built by injecting several dilutions of the non-labelled standards mixed with a fixed concentration of labelled standards. Each calibration point is prepared and analysed in triplicates (n=3).
To establish the endogenous concentration of the four selected ceramides in the study samples (NIST SRM1950, NIST high TAG, NIST T1D, and NIST young A-A) each sample is extracted and analyzed in six replicates (n =6).
Prepare the pooled QC by mixing 100 mL each of NIST SRM1950, NIST high TAG, NIST T1D, and NIST young AA (total 400 mL).
We suggest proceeding first to measure the ceramides concentration in the four study samples and then continue with the intra-assay validation to avoid wasting material.

## Extraction - Overview
Extraction Overview
Extraction Overview | Table 1 - Extraction - Materials
Chemicals, solvents and materials
Ethyl acetate (EtAc)
2-propanol (IPA)
5% Bovine Serum Albumin (BSA)
1.5ml-Eppendorf tubes
LC Vials
Table 2 - Extraction - Sample Types
| Sample type | Components
| Pooled Quality Control samples (QC) | 10 µL pooled QC |  |  | 20 µL Labelled IS |  | 570 µL EtAc:IPA 2:8 (v/v)
| Matrix Blank (MB) | 10 µL Matrix (pooled QC) only |  |  | - |  | 590 µL EtAc:IPA 2:8
| Study Samples (S) | 10 µL Human plasma (NIST SRM1950, high TAG, T1D, young AA) |  |  | 20 µL Labelled IS |  | 570 µL EtAc:IPA 2:8
| Total Blank (TB) | 10 µL 5% BSA |  |  | - |  | 590 µL EtAc:IPA 2:8
| Calibration line standards (STD) | 10 µL 5% BSA + 20 µL non labelled STD mixture (Table 5) |  |  | 20 µL Labelled IS |  | 550 µL EtAc:IPA 2:8
Table 3 - Extraction - Labelled Internal Standard Mixture Solution (provided)
| Labelled IS mixture in EtAc:IPA 2:8 (v/v) | Concentration |  |  |  |  | Unit | Vendor
| D7-Cer d18:1/16:0 | 0.125 |  |  |  |  | pmol/µL [µM] | Avanti Polar Lipids
| D7-Cer d18:1/18:0 | 0.050
| D7-Cer d18:1/24:0 | 1.5
| D7-Cer d18:1/24:1 | 0.500
Table 4 - Extraction - Non-Labelled Standards - for STD1 Mixture Solution (provided)
| Non-labelled standards in EtAc:IPA 2:8 (v/v) | Concentration |  |  |  |  | Unit | Vendor
| Cer d18:1/16:0 | 2.0 |  |  |  |  | pmol/µL [µM] | Avanti Polar Lipids
| Cer d18:1/18:0 | 2.0
| Cer d18:1/24:0 | 20.0
| Cer d18:1/24:1 | 20.0
Table 5 - Serial dilutions of Non-Labelled Standards from STD1 Mixture Solution for calibration line
| Add 20 µL of | STD1 Conc. | STD2 Conc. | STD3 Conc. | STD4 Conc. | STD5 Conc. | STD6 Conc. | Unit
| Cer d18:1/16:0 | 2.0 | 1.0 | 0.10 | 0.02 | 0.01 | 0.008 | pmol/µL [µM]
| Cer d18:1/18:0 | 2.0 | 1.0 | 0.10 | 0.02 | 0.01 | 0.008
| Cer d18:1/24:0 | 20.0 | 10.0 | 1.0 | 0.2 | 0.1 | 0.08
| Cer d18:1/24:1 | 20.0 | 10.0 | 1.0 | 0.2 | 0.1 | 0.08

## Extraction - Procedure
Extraction Procedure
Extraction Procedure | Extraction Procedure
| Step | Description | Reference Table
| 1 | Prepare different dilutions of STD1 by diluting the provided mixture (Extraction - Overview, Table 4) with EtAc:IPA 2:8 as suggested in table | Extraction - Overview'!B34
| 2 | Thaw plasma on ice (+4°C) and bring to room temperature prior to extraction. Mix by pipetting before use.
| 3 | Prepare samples in plates/tubes manually or with a robot, according to specifications in column C of table | Extraction - Overview'!C14
| 4 | Add 20 µL of labelled internal standard mixture to samples according to column F of table | Extraction - Overview'!F14
| 5 | For the calibration line, add 20 µL of different dilutions of the non-labelled mixture (STD1-6).
| 6 | Add 590 µL of ethyl acetate:isopropanol (2:8, vol/vol) to Total Blank (TB) and Matrix Blank (MB); 550 µL to calibration line (STD); 570 µL to study samples and QCs according to column H of table | Extraction - Overview'!H14
| 7 | Mix samples 10 min by pipetting or vortexing
| 8 | Centrifuge samples for 10 min at 3000 x gat room temperature
| 9 | Transfer 50 µL of the clear supernatant to MS-vials or MS plates, close/seal, store samples at -20 °C prior to LC-MRM analysis.

## LC-MRM - Overview
LC-MRM Overview
LC-MRM Overview | Table 6 Chemicals and solvents
Ammonium acetate
Formic acid
Ultra-pure water
Acetonitrile
Table 7 LC-MRM LC-MS Details
| Property | Specifications
| Mass spectrometry | Sciex Qtrap 6500+
| LC | Shimadzu Nexera X2
| Injection volume (indicative) | 3 µL
| Weak wash | Water
| Strong wash | Acetonitrile
| Pre-column | Waters Acquity BEH C18, 1.7 µm VanGuard Pre-Column
| Column | Waters Acquity BEH C18, 2.1 × 50 mm id. 1.7 µm
| Column temperature | 60 °C
| Mobile phase A | 10 mM ammonium acetate in water with 0.1% formic acid
| Mobile phase B | 10 mM ammonium acetate in acetonitrile:2-propanol (4:3, vol/vol) with 0.1% formic acid
Table 8 LC-MRM LC gradient
| Time (minutes) | A% | B% |  | Flow (µL/min)
| 0.00 | 15 | 85 |  | 500
| 0.50 | 15 | 85 |  | 500
| 1.50 | 0 | 100 |  | 500
| 4.00 | 0 | 100 |  | 500
| 4.10 | 15 | 85 |  | 500
| 5.00 | 15 | 85 |  | 500
Table 9 LC-MRM MS conditions (EXAMPLE: this only applies to Sciex QTRAP, please adapt)
| Polarity mode | Positive
| Ion spray | 5000 V
| Curtain gas | 25 psi
| Source temperature | 300 °C
| Gas 1 | 50 psi
| Gas 2 | 35 psi
| DP | 30 V
| EP | 10 V
| Collision cell exit potential | 20 V
| Collision energy (CE) | 40 eV
Table 10 LC-MRM Mass transitions used for absolute quantification. On different systems, RT and CE might be different!
| Analyte | Mass transitions | IS | Mass transitions | RT (min) | CE
| Cer d18:1/16:0 | 538.5 – 264.3 | D7- Cer d18:1/16:0 | 545.6 - 271.3 | 1.9-2.0 | 40
C34H68NO3 – C18H34N
| Cer d18:1/18:0 | 566.6 – 264.3 | D7- Cer d18:1/18:0 | 573.6 – 271.3 | 2.1-2.2 | 40
C36H72NO3 – C18H34N
| Cer d18:1/24:0 | 650.6 – 264.3 | D7- Cer d18:1/24:0 | 657.7 – 271.3 | 2.4-2.5 | 40
C42H84NO3 – C18H34N
| Cer d18:1/24:1 | 648.6 – 264.3 | D7- Cer d18:1/24:1 | 655.7 – 271.3 | 2.3-2.4 | 40
C42H82NO3 – C18H34N
Table 11 Data processing
| Processing software | Analyst 1.6.3

## LC-MRM - Sample Sequence
Sample Sequence (Example)
Sample Sequence (Example) | Sample Sequence 1
| Number | Sample Name | Remarks
| 1 | Total Blank 1
| 2 | Matrix blank | only two replicates in total
| 3 | Calibration line 1 from STD6 to 1 (n=3)
| 4 | […]
| 5 | Pooled QC
| 6 | Pooled QC
| 7 | SRM 1950-1
| 8 | SRM 1950-2
| 9 | SRM 1950-3
| 10 | SRM 1950-4
| 11 | SRM 1950-5
| 12 | SRM 1950-6
| 13 | Pooled QC
| 14 | Pooled QC
| 15 | T1D-1
| 16 | T1D-2
| 17 | T1D-3
| 18 | T1D-4
| 19 | T1D-5
| 20 | T1D-6
| 21 | Total Blank 2
| 22 | Pooled QC
| 23 | Pooled QC
| 24 | Young AA-1
| 25 | Young AA-2
| 26 | Young AA-3
| 27 | Young AA-4
| 28 | Young AA-5
| 29 | Young AA-6 | Young AA-6 Rep2 failed
| 30 | Pooled QC
| 31 | Pooled QC
| 32 | hTAG-1
| 33 | hTAG-2
| 34 | hTAG-3
| 35 | hTAG-4
| 36 | hTAG-5
| 37 | hTAG-6
| 38 | Pooled QC
| 39 | Pooled QC
| 40 | Calibration line 2 from STD1 to 6 (n=3)
| 41 | Intra-assay QCs: LLQC, LQC, MQC, HQC | LLQC: 40µL pooled QC + 80µL H2O -> 10µL taken for each extraction (no further dilutions tested) LQC: 40µL pooled QC + 40µL H2O -> 10µL taken for each extraction
| 42 | Matrix blank | only two replicates in total
| 43 | Total Blank 3

## LC-MRM - Intra Assay QC Proc
Intra Assay Procedure
Intra Assay Procedure | Quality Controls for Intra-Assay Variation
| Step | Description | Reference Table
| 1 | Prepare the Pooled NIST QC sample containing equal volumes of the four NIST reference plasma you received.
| 2 | The pooled QC representing middle quality control (MQC), low QC (LQC), and lower limit of quantification QC (LLOQ) is prepared by diluting MQC (e.g. two- and three-fold) with water.
| 3 | High QC (HQC) and upper limit of quantification (ULOQ) QC are prepared by spiking in 10 µL and 20 µL respectively of STD3 endogenous standards from table | Extraction - Overview'!B34
| 4 | Internal standards are added to all QC samples.
Intra-Assay Variation
| 6 | The precision and accuracy of the assay is determined for each of the four ceramides measured in LLOQ, LQC, MQC, HQC, and HLOQ. |  | HLOQ = HLQC
| 7 | Intra-assay variation, precision, and accuracy is calculated for each of the ceramide in replicates of six (n = 6) at each QC concentration.
| 8 | The intra-assay precision (percentage coefficient variance, %CV) and accuracy (percentage accuracy, %Accuracy) are calculated from the nominal concentrations according to the formulas reported in corresponding sheets and tables of this document.
| 9 | Recovery estimation is not necessary for the published method but only for different ones. Stability estimation is not necessary for the published method but only for different ones. Calibration line is prepared in 5% BSA.

## LC-MRM - Intra Assay QC Res
INTRA ASSAY - QC RESULTS
INTRA ASSAY - QC RESULTS | Table 12 LC-MRM Calibration Line 1 Data (Beginning). Please report the values of each of the three replicates separately for each compound and standard. Each STD prepared in triplicate, injected once from each vial
| Calibration Line 1 | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| STD 1 | 4320000 | 4330000 | 4210000 | 4590000 | 4750000 | 4790000 | 68100000 | 66500000 | 67800000 | 54600000 | 55800000 | 56600000 | 194000 | 197000 | 200000 | 194000 | 191000 | 195000 | 4910000 | 4760000 | 5060000 | 1400000 | 1420000 | 1450000
| STD 2 | 2130000 | 2180000 | 2160000 | 2340000 | 2380000 | 2370000 | 30400000 | 30100000 | 31200000 | 25100000 | 25800000 | 25300000 | 193000 | 195000 | 194000 | 187000 | 186000 | 186000 | 4340000 | 4330000 | 4380000 | 1270000 | 1250000 | 1280000
| STD 3 | 223000 | 218000 | 221000 | 246000 | 236000 | 241000 | 2770000 | 2860000 | 2770000 | 2480000 | 2580000 | 2510000 | 201000 | 182000 | 198000 | 195000 | 191000 | 189000 | 3860000 | 3930000 | 3930000 | 1190000 | 1260000 | 1240000
| STD 4 | 49500 | 40400 | 44200 | 52800 | 47700 | 50200 | 553000 | 519000 | 569000 | 496000 | 452000 | 501000 | 199000 | 203000 | 198000 | 177000 | 191000 | 202000 | 3760000 | 3890000 | 3920000 | 1240000 | 1280000 | 1230000
| STD 5 | 23500 | 24100 | 20500 | 24500 | 25300 | 24000 | 297000 | 291000 | 289000 | 241000 | 244000 | 232000 | 193000 | 190000 | 197000 | 195000 | 190000 | 186000 | 4010000 | 3960000 | 3880000 | 1240000 | 1220000 | 1240000
| STD 6 | 18700 | 19200 | 19200 | 21900 | 21600 | 20700 | 235000 | 229000 | 229000 | 195000 | 206000 | 192000 | 188000 | 200000 | 196000 | 176000 | 186000 | 192000 | 3790000 | 3860000 | 3770000 | 1200000 | 1230000 | 1250000
Table 13 LC-MRM Calibration Line 2 Data (End). Please report the values of each of the three replicates separately for each compound and standard. Each STD prepared in triplicate, injected once from each vial
| Calibration Line 2 | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| STD 1 | 4330000 | 4300000 | 4310000 | 4580000 | 4780000 | 4860000 | 70400000 | 69600000 | 67500000 | 57300000 | 57300000 | 59300000 | 204000 | 200000 | 200000 | 193000 | 198000 | 186000 | 5040000 | 5030000 | 5010000 | 1470000 | 1520000 | 1480000
| STD 2 | 2110000 | 2200000 | 2110000 | 2180000 | 2390000 | 2360000 | 31000000 | 32000000 | 32200000 | 26400000 | 26800000 | 27100000 | 199000 | 201000 | 203000 | 187000 | 198000 | 200000 | 4460000 | 4520000 | 4580000 | 1320000 | 1310000 | 1310000
| STD 3 | 231000 | 224000 | 219000 | 255000 | 249000 | 252000 | 2950000 | 2840000 | 3030000 | 2680000 | 2610000 | 2670000 | 205000 | 209000 | 198000 | 196000 | 202000 | 197000 | 4260000 | 4210000 | 4160000 | 1300000 | 1320000 | 1330000
| STD 4 | 50600 | 46100 | 84600 | 46400 | 50900 | 74700 | 616000 | 566000 | 710000 | 531000 | 521000 | 512000 | 212000 | 195000 | 210000 | 199000 | 204000 | 181000 | 4140000 | 4070000 | 4110000 | 1310000 | 1330000 | 1270000
| STD 5 | 24200 | 4940 | 27100 | 23700 | 7780 | 27400 | 299000 | 73500 | 297000 | 254000 | 55900 | 257000 | 207000 | 41700 | 199000 | 195000 | 36600 | 190000 | 4010000 | 880000 | 4080000 | 209000 | 213000 | 213000
| STD 6 | 20400 | 18600 | 19900 | 20600 | 22100 | 18900 | 250000 | 248000 | 250000 | 1310000 | 280000 | 1270000 | 211000 | 204000 | 211000 | 187000 | 188000 | 201000 | 3940000 | 4230000 | 4010000 | 1320000 | 1280000 | 1310000
Table 14 LC-MRM Sample Data. Please report the values of each of the three replicates separately for each compound and standard. Each vial injected 3 times
| NIST SRM | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| SRM 1 | 279000 | 266000 | 270000 | 125000 | 129000 | 123000 | 3720000 | 3470000 | 3620000 | 1170000 | 1140000 | 1130000 | 195000 | 202000 | 196000 | 232000 | 235000 | 227000 | 4240000 | 4170000 | 4030000 | 1350000 | 1320000 | 1350000
| SRM 2 | 277000 | 265000 | 270000 | 125000 | 124000 | 122000 | 3690000 | 3590000 | 3670000 | 1180000 | 1110000 | 1110000 | 207000 | 197000 | 204000 | 231000 | 233000 | 229000 | 4240000 | 4000000 | 4090000 | 1360000 | 1320000 | 1290000
| SRM 3 | 273000 | 266000 | 263000 | 126000 | 122000 | 126000 | 3780000 | 3670000 | 3590000 | 1220000 | 1170000 | 1140000 | 199000 | 205000 | 203000 | 239000 | 236000 | 231000 | 4360000 | 4110000 | 4080000 | 1370000 | 1360000 | 1360000
| SRM 4 | 267000 | 269000 | 278000 | 129000 | 123000 | 127000 | 3550000 | 3630000 | 3490000 | 1170000 | 1080000 | 1110000 | 202000 | 201000 | 204000 | 235000 | 226000 | 227000 | 4120000 | 4140000 | 4200000 | 1360000 | 1350000 | 1370000
| SRM 5 | 274000 | 275000 | 270000 | 121000 | 127000 | 109000 | 3780000 | 3710000 | 3560000 | 1210000 | 1130000 | 1110000 | 211000 | 206000 | 201000 | 238000 | 233000 | 241000 | 4420000 | 4070000 | 4020000 | 1380000 | 1370000 | 1320000
| SRM 6 | 271000 | 275000 | 268000 | 128000 | 126000 | 119000 | 3750000 | 3750000 | 3600000 | 1160000 | 1190000 | 1090000 | 212000 | 211000 | 206000 | 224000 | 225000 | 225000 | 4350000 | 4400000 | 4030000 | 1350000 | 1410000 | 1380000
| NIST hTAG | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| hTAG 1 | 312000 | 304000 | 295000 | 154000 | 148000 | 149000 | 5920000 | 5990000 | 5900000 | 1830000 | 1650000 | 1680000 | 179000 | 182000 | 183000 | 238000 | 233000 | 210000 | 4400000 | 4290000 | 4160000 | 1380000 | 1270000 | 1350000
| hTAG 2 | 293000 | 304000 | 314000 | 154000 | 152000 | 152000 | 5870000 | 6010000 | 5870000 | 1720000 | 1740000 | 1730000 | 194000 | 193000 | 184000 | 234000 | 230000 | 241000 | 4310000 | 4250000 | 4190000 | 1420000 | 1360000 | 1350000
| hTAG 3 | 296000 | 304000 | 296000 | 160000 | 146000 | 144000 | 5990000 | 5680000 | 5900000 | 1720000 | 1670000 | 1620000 | 197000 | 187000 | 189000 | 239000 | 236000 | 224000 | 4500000 | 4040000 | 4150000 | 1380000 | 1340000 | 1360000
| hTAG 4 | 296000 | 296000 | 294000 | 152000 | 148000 | 138000 | 6120000 | 5850000 | 5710000 | 1650000 | 1650000 | 1630000 | 193000 | 189000 | 184000 | 220000 | 226000 | 224000 | 4500000 | 4020000 | 3960000 | 1320000 | 1340000 | 1330000
| hTAG 5 | 300000 | 305000 | 311000 | 150000 | 143000 | 144000 | 5790000 | 5820000 | 5860000 | 1660000 | 1610000 | 1620000 | 188000 | 188000 | 186000 | 230000 | 227000 | 231000 | 4360000 | 4200000 | 3950000 | 1320000 | 1320000 | 1300000
| hTAG 6 | 294000 | 300000 | 290000 | 157000 | 150000 | 147000 | 6090000 | 5700000 | 5600000 | 1770000 | 1640000 | 1670000 | 192000 | 195000 | 183000 | 234000 | 226000 | 231000 | 4380000 | 4180000 | 4100000 | 1430000 | 1340000 | 1340000
| NIST T1D | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| T1D 1 | 253000 | 245000 | 244000 | 143000 | 148000 | 146000 | 3760000 | 3590000 | 3570000 | 1180000 | 1120000 | 1170000 | 214000 | 208000 | 197000 | 226000 | 232000 | 241000 | 4500000 | 4290000 | 4270000 | 1280000 | 1370000 | 1330000
| T1D 2 | 256000 | 243000 | 246000 | 140000 | 142000 | 148000 | 3900000 | 3860000 | 3720000 | 1190000 | 1170000 | 1170000 | 211000 | 204000 | 205000 | 240000 | 230000 | 239000 | 4670000 | 4440000 | 4160000 | 1390000 | 1430000 | 1380000
| T1D 3 | 247000 | 245000 | 248000 | 141000 | 139000 | 147000 | 3730000 | 3630000 | 3690000 | 1220000 | 1130000 | 1140000 | 211000 | 203000 | 203000 | 241000 | 236000 | 228000 | 4500000 | 4190000 | 4190000 | 1420000 | 1390000 | 1330000
| T1D 4 | 249000 | 248000 | 248000 | 137000 | 140000 | 134000 | 3810000 | 3590000 | 3610000 | 1210000 | 1140000 | 1140000 | 209000 | 214000 | 210000 | 246000 | 237000 | 232000 | 4410000 | 4280000 | 4120000 | 1380000 | 1340000 | 1340000
| T1D 5 | 249000 | 249000 | 248000 | 147000 | 139000 | 138000 | 3640000 | 3500000 | 3570000 | 1180000 | 1100000 | 1130000 | 212000 | 210000 | 210000 | 240000 | 225000 | 237000 | 4420000 | 4090000 | 4140000 | 1410000 | 1360000 | 1350000
| T1D 6 | 261000 | 254000 | 251000 | 149000 | 146000 | 141000 | 3820000 | 3750000 | 3830000 | 1230000 | 1190000 | 1170000 | 206000 | 209000 | 210000 | 248000 | 240000 | 251000 | 4380000 | 4250000 | 4420000 | 1410000 | 1390000 | 1440000
| NIST Young AA | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| yA-A 1 | 185000 | 188000 | 191000 | 78600 | 75800 | 75100 | 2690000 | 2650000 | 2540000 | 782000 | 749000 | 770000 | 229000 | 233000 | 224000 | 249000 | 238000 | 243000 | 4670000 | 4420000 | 4110000 | 1410000 | 1410000 | 1450000
| yA-A 2 | 191000 | 186000 | 187000 | 73300 | 71000 | 75100 | 2560000 | 2560000 | 2580000 | 811000 | 778000 | 732000 | 225000 | 225000 | 224000 | 236000 | 223000 | 230000 | 4550000 | 4180000 | 4200000 | 1450000 | 1500000 | 1360000
| yA-A 3 | 174000 | 183000 | 181000 | 71900 | 66600 | 69900 | 2580000 | 2450000 | 2450000 | 767000 | 732000 | 744000 | 214000 | 212000 | 211000 | 231000 | 215000 | 231000 | 4280000 | 3810000 | 3920000 | 1360000 | 1350000 | 1360000
| yA-A 4 | 184000 | 175000 | 188000 | 73200 | 64900 | 68100 | 2540000 | 2420000 | 2350000 | 739000 | 745000 | 713000 | 219000 | 207000 | 207000 | 224000 | 223000 | 210000 | 4210000 | 3730000 | 3870000 | 1360000 | 1300000 | 1310000
| yA-A 5 | 182000 | 182000 | 185000 | 72600 | 75500 | 69700 | 2650000 | 2550000 | 2650000 | 761000 | 758000 | 795000 | 223000 | 224000 | 224000 | 231000 | 239000 | 233000 | 4470000 | 4210000 | 4390000 | 1430000 | 1370000 | 1420000
| yA-A 6 | 187000 | 0 | 191000 | 72800 | 0 | 71100 | 2570000 | 0 | 2470000 | 773000 | 0 | 764000 | 226000 | 0 | 222000 | 229000 | 0 | 224000 | 4460000 | 0 | 4100000 | 1400000 | 0 | 1450000
Table 15 LC-MRM Intra Assay QC Data. Please report the values of each of the three replicates separately for each compound and standard. QCs have to be measured in one occasion only. The independent 6 replicates injected 3 times each
| Lowest Level QC | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| LLQC 1 | 101000 | 98600 | 96500 | 36100 | 41300 | 38700 | 1250000 | 1290000 | 1230000 | 379000 | 370000 | 375000 | 228000 | 223000 | 239000 | 210000 | 208000 | 214000 | 4090000 | 4150000 | 3870000 | 1340000 | 1300000 | 1310000
| LLQC 2 | 99000 | 101000 | 103000 | 37400 | 39800 | 37800 | 1310000 | 1240000 | 1270000 | 378000 | 392000 | 376000 | 234000 | 239000 | 240000 | 214000 | 222000 | 207000 | 4190000 | 3990000 | 4140000 | 1300000 | 1310000 | 1360000
| LLQC 3 | 102000 | 98900 | 99000 | 38400 | 36000 | 36100 | 1240000 | 1280000 | 1240000 | 381000 | 379000 | 377000 | 230000 | 226000 | 230000 | 211000 | 196000 | 210000 | 4250000 | 3830000 | 4000000 | 1330000 | 1260000 | 1280000
| LLQC 4 | 101000 | 99700 | 100000 | 40600 | 36800 | 39900 | 1320000 | 1240000 | 1280000 | 384000 | 399000 | 388000 | 242000 | 222000 | 226000 | 215000 | 212000 | 215000 | 4320000 | 4150000 | 4090000 | 1380000 | 1340000 | 1320000
| LLQC 5 | 100000 | 98600 | 92100 | 39000 | 37600 | 37200 | 1260000 | 1220000 | 1200000 | 392000 | 366000 | 369000 | 241000 | 233000 | 234000 | 212000 | 208000 | 217000 | 4140000 | 4160000 | 4020000 | 1370000 | 1320000 | 1260000
| LLQC 6 | 95500 | 103000 | 101000 | 40800 | 36600 | 38500 | 1280000 | 1230000 | 1190000 | 395000 | 369000 | 370000 | 242000 | 235000 | 229000 | 209000 | 223000 | 217000 | 4400000 | 4190000 | 4100000 | 1390000 | 1310000 | 1330000
| Low QC | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| LQC 1 | 148000 | 140000 | 152000 | 61900 | 53700 | 58300 | 1940000 | 1910000 | 1820000 | 587000 | 583000 | 583000 | 235000 | 218000 | 232000 | 217000 | 217000 | 219000 | 4410000 | 4160000 | 4090000 | 1390000 | 1300000 | 1390000
| LQC 2 | 145000 | 145000 | 134000 | 54300 | 61400 | 55500 | 1920000 | 1850000 | 1860000 | 597000 | 581000 | 558000 | 231000 | 235000 | 220000 | 221000 | 213000 | 216000 | 4360000 | 3910000 | 4150000 | 1330000 | 1350000 | 1350000
| LQC 3 | 143000 | 142000 | 141000 | 56500 | 56100 | 57200 | 1870000 | 1860000 | 1890000 | 578000 | 592000 | 571000 | 233000 | 224000 | 223000 | 211000 | 208000 | 213000 | 4240000 | 4270000 | 4120000 | 1310000 | 1360000 | 1390000
| LQC 4 | 150000 | 147000 | 144000 | 59600 | 55200 | 57300 | 1970000 | 1960000 | 1870000 | 602000 | 567000 | 580000 | 236000 | 228000 | 232000 | 221000 | 211000 | 223000 | 4340000 | 4130000 | 4010000 | 1390000 | 1340000 | 1340000
| LQC 5 | 147000 | 142000 | 142000 | 58500 | 58200 | 57500 | 1990000 | 1860000 | 1930000 | 600000 | 563000 | 587000 | 229000 | 218000 | 237000 | 216000 | 223000 | 217000 | 4380000 | 4040000 | 4410000 | 1350000 | 1390000 | 1330000
| LQC 6 | 147000 | 144000 | 140000 | 63400 | 61100 | 59200 | 1990000 | 1870000 | 1860000 | 606000 | 553000 | 561000 | 238000 | 227000 | 232000 | 230000 | 229000 | 220000 | 4440000 | 4170000 | 4050000 | 1370000 | 1350000 | 1370000
| Middle QC | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| MQC 1 | 263000 | 252000 | 252000 | 120000 | 112000 | 120000 | 3800000 | 3690000 | 3550000 | 1200000 | 1090000 | 1130000 | 216000 | 204000 | 200000 | 227000 | 216000 | 225000 | 4250000 | 4040000 | 3940000 | 1420000 | 1270000 | 1300000
| MQC 2 | 264000 | 253000 | 256000 | 130000 | 116000 | 127000 | 3910000 | 3810000 | 3960000 | 1210000 | 1170000 | 1180000 | 205000 | 206000 | 209000 | 238000 | 226000 | 229000 | 4520000 | 4160000 | 4270000 | 1450000 | 1370000 | 1390000
| MQC 3 | 263000 | 246000 | 257000 | 127000 | 115000 | 118000 | 4060000 | 3980000 | 3810000 | 1220000 | 1210000 | 1180000 | 206000 | 199000 | 198000 | 214000 | 224000 | 225000 | 4340000 | 4130000 | 3960000 | 1350000 | 1330000 | 1310000
| MQC 4 | 251000 | 245000 | 247000 | 123000 | 118000 | 119000 | 4000000 | 3670000 | 3900000 | 1220000 | 1180000 | 1200000 | 204000 | 203000 | 201000 | 220000 | 211000 | 225000 | 4280000 | 3910000 | 4130000 | 1350000 | 1350000 | 1310000
| MQC 5 | 256000 | 255000 | 257000 | 118000 | 122000 | 126000 | 3940000 | 3830000 | 4020000 | 1200000 | 1150000 | 1190000 | 199000 | 204000 | 200000 | 222000 | 217000 | 223000 | 4460000 | 4110000 | 4070000 | 1390000 | 1300000 | 1350000
| MQC 6 | 254000 | 237000 | 251000 | 125000 | 125000 | 130000 | 4020000 | 3970000 | 3800000 | 1210000 | 1150000 | 1110000 | 203000 | 195000 | 194000 | 221000 | 221000 | 221000 | 4270000 | 4110000 | 3930000 | 1350000 | 1300000 | 1310000
| High QC | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| HQC 1 | 355000 | 353000 | 355000 | 263000 | 254000 | 240000 | 5300000 | 5240000 | 5210000 | 2700000 | 2450000 | 2480000 | 213000 | 202000 | 202000 | 221000 | 215000 | 220000 | 4460000 | 3980000 | 4040000 | 1370000 | 1350000 | 1360000
| HQC 2 | 354000 | 347000 | 356000 | 263000 | 241000 | 246000 | 5490000 | 5140000 | 4990000 | 2560000 | 2500000 | 2350000 | 200000 | 202000 | 201000 | 229000 | 217000 | 216000 | 4310000 | 3930000 | 3860000 | 1340000 | 1280000 | 1360000
| HQC 3 | 350000 | 347000 | 349000 | 259000 | 251000 | 248000 | 5260000 | 5130000 | 5230000 | 2480000 | 2390000 | 2420000 | 205000 | 195000 | 204000 | 232000 | 216000 | 215000 | 4320000 | 4090000 | 4150000 | 1340000 | 1390000 | 1300000
| HQC 4 | 365000 | 366000 | 352000 | 253000 | 258000 | 254000 | 5630000 | 5340000 | 5280000 | 2600000 | 2470000 | 2490000 | 216000 | 225000 | 205000 | 225000 | 229000 | 225000 | 4550000 | 4060000 | 4130000 | 1390000 | 1350000 | 1290000
| HQC 5 | 369000 | 349000 | 361000 | 265000 | 252000 | 248000 | 5390000 | 5210000 | 4920000 | 2600000 | 2450000 | 2470000 | 205000 | 200000 | 199000 | 222000 | 213000 | 223000 | 4360000 | 3970000 | 3780000 | 1340000 | 1260000 | 1300000
| HQC 6 | 355000 | 355000 | 348000 | 247000 | 240000 | 246000 | 5380000 | 5030000 | 4960000 | 2510000 | 2380000 | 2490000 | 207000 | 201000 | 199000 | 217000 | 222000 | 207000 | 4180000 | 3890000 | 3850000 | 1370000 | 1290000 | 1310000
| Highest Level QC | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| HLQC 1 | 462000 | 462000 | 463000 | 462000 | 462000 | 463000 | 6900000 | 6690000 | 6700000 | 4020000 | 3790000 | 3810000 | 207000 | 204000 | 204000 | 207000 | 204000 | 204000 | 4400000 | 3940000 | 4020000 | 1390000 | 1320000 | 1280000
| HLQC 2 | 468000 | 465000 | 461000 | 468000 | 465000 | 461000 | 6950000 | 6430000 | 6810000 | 3830000 | 3820000 | 3750000 | 211000 | 202000 | 205000 | 211000 | 202000 | 205000 | 4500000 | 4030000 | 4040000 | 1370000 | 1290000 | 1320000
| HLQC 3 | 483000 | 469000 | 460000 | 483000 | 469000 | 460000 | 7080000 | 6770000 | 6860000 | 4080000 | 3980000 | 3890000 | 215000 | 209000 | 208000 | 215000 | 209000 | 208000 | 4340000 | 4160000 | 3970000 | 1440000 | 1370000 | 1330000
| HLQC 4 | 481000 | 478000 | 481000 | 481000 | 478000 | 481000 | 7140000 | 7030000 | 6930000 | 4000000 | 3910000 | 3760000 | 205000 | 195000 | 196000 | 205000 | 195000 | 196000 | 4250000 | 4160000 | 3970000 | 1360000 | 1320000 | 1320000
| HLQC 5 | 538000 | 572000 | 561000 | 538000 | 572000 | 561000 | 7020000 | 6910000 | 6650000 | 4030000 | 3690000 | 3710000 | 207000 | 198000 | 209000 | 207000 | 198000 | 209000 | 4410000 | 3890000 | 3950000 | 1340000 | 1260000 | 1280000
| HLQC 6 | 467000 | 448000 | 462000 | 467000 | 448000 | 462000 | 6910000 | 6740000 | 6740000 | 4150000 | 3650000 | 3930000 | 201000 | 206000 | 208000 | 201000 | 206000 | 208000 | 4360000 | 3890000 | 4220000 | 1400000 | 1280000 | 1380000
| Blank QC | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| Matrix Blank | 243000 | 262000 | 0 | 120000 | 127000 | 0 | 3980000 | 4780000 | 0 | 1210000 | 1440000 | 0 | 254 | 419 | 0 | 446 | 357 | 0 | 657 | 1190 | 0 | 182 | 536 | 0
| Total Blank | 17200 | 4150 | 3830 | 6170 | 1220 | 2180 | 28900 | 21100 | 19000 | 4570 | 3040 | 2520 | 478 | 402 | 267 | 553 | 300 | 354 | 683 | 1010 | 472 | 197 | 374 | 528

## Raw Peak Areas
Sample_Name | Cer d18:1/16:0 [M>264] | Cer d18:1-d7/16:0 [M>271] | Cer d18:1/18:0 [M>264] | Cer d18:1-d7/18:0 [M>271] | Cer d18:1/24:0 [M>264] | Cer d18:1-d7/24:0 [M>271] | Cer d18:1/24:1 [M>264] | Cer d18:1-d7/24:1 [M>271]
Total Blank 1 | 17200 | 478 | 6170 | 553 | 28900 | 683 | 4570 | 197
Matrix Blank | 243000 | 254 | 120000 | 446 | 3980000 | 657 | 1210000 | 182
STD6-1 | 18700 | 188000 | 21900 | 176000 | 235000 | 3790000 | 195000 | 1200000
STD6-2 | 19200 | 200000 | 21600 | 186000 | 229000 | 3860000 | 206000 | 1230000
STD6-3 | 19200 | 196000 | 20700 | 192000 | 229000 | 3770000 | 192000 | 1250000
STD5-1 | 23500 | 193000 | 24500 | 195000 | 297000 | 4010000 | 241000 | 1240000
STD5-2 | 24100 | 190000 | 25300 | 190000 | 291000 | 3960000 | 244000 | 1220000
STD5-3 | 20500 | 197000 | 24000 | 186000 | 289000 | 3880000 | 232000 | 1240000
STD4-1 | 49500 | 199000 | 52800 | 177000 | 553000 | 3760000 | 496000 | 1240000
STD4-2 | 40400 | 203000 | 47700 | 191000 | 519000 | 3890000 | 452000 | 1280000
STD4-3 | 44200 | 198000 | 50200 | 202000 | 569000 | 3920000 | 501000 | 1230000
STD3-1 | 223000 | 201000 | 246000 | 195000 | 2770000 | 3860000 | 2480000 | 1190000
STD3-2 | 218000 | 182000 | 236000 | 191000 | 2860000 | 3930000 | 2580000 | 1260000
STD3-3 | 221000 | 198000 | 241000 | 189000 | 2770000 | 3930000 | 2510000 | 1240000
STD2-1 | 2130000 | 193000 | 2340000 | 187000 | 30400000 | 4340000 | 25100000 | 1270000
STD2-2 | 2180000 | 195000 | 2380000 | 186000 | 30100000 | 4330000 | 25800000 | 1250000
STD2-3 | 2160000 | 194000 | 2370000 | 186000 | 31200000 | 4380000 | 25300000 | 1280000
STD1-1 | 4320000 | 194000 | 4590000 | 194000 | 68100000 | 4910000 | 54600000 | 1400000
STD1-2 | 4330000 | 197000 | 4750000 | 191000 | 66500000 | 4760000 | 55800000 | 1420000
STD1-3 | 4210000 | 200000 | 4790000 | 195000 | 67800000 | 5060000 | 56600000 | 1450000
Pooled QC 1 | 254000 | 212000 | 134000 | 235000 | 3990000 | 4460000 | 1280000 | 1410000
Pooled QC 1 | 266000 | 201000 | 132000 | 232000 | 3950000 | 4070000 | 1180000 | 1350000
SRM 1950-1_1 | 279000 | 195000 | 125000 | 232000 | 3720000 | 4240000 | 1170000 | 1350000
SRM 1950-1_2 | 266000 | 202000 | 129000 | 235000 | 3470000 | 4170000 | 1140000 | 1320000
SRM 1950-1_3 | 270000 | 196000 | 123000 | 227000 | 3620000 | 4030000 | 1130000 | 1350000
SRM 1950-2_1 | 277000 | 207000 | 125000 | 231000 | 3690000 | 4240000 | 1180000 | 1360000
SRM 1950-2_2 | 265000 | 197000 | 124000 | 233000 | 3590000 | 4000000 | 1110000 | 1320000
SRM 1950-2_3 | 270000 | 204000 | 122000 | 229000 | 3670000 | 4090000 | 1110000 | 1290000
SRM 1950-3_1 | 273000 | 199000 | 126000 | 239000 | 3780000 | 4360000 | 1220000 | 1370000
SRM 1950-3_2 | 266000 | 205000 | 122000 | 236000 | 3670000 | 4110000 | 1170000 | 1360000
SRM 1950-3_3 | 263000 | 203000 | 126000 | 231000 | 3590000 | 4080000 | 1140000 | 1360000
SRM 1950-4_1 | 267000 | 202000 | 129000 | 235000 | 3550000 | 4120000 | 1170000 | 1360000
SRM 1950-4_2 | 269000 | 201000 | 123000 | 226000 | 3630000 | 4140000 | 1080000 | 1350000
SRM 1950-4_3 | 278000 | 204000 | 127000 | 227000 | 3490000 | 4200000 | 1110000 | 1370000
SRM 1950-5_1 | 274000 | 211000 | 121000 | 238000 | 3780000 | 4420000 | 1210000 | 1380000
SRM 1950-5_2 | 275000 | 206000 | 127000 | 233000 | 3710000 | 4070000 | 1130000 | 1370000
SRM 1950-5_3 | 270000 | 201000 | 109000 | 241000 | 3560000 | 4020000 | 1110000 | 1320000
SRM 1950-6_1 | 271000 | 212000 | 128000 | 224000 | 3750000 | 4350000 | 1160000 | 1350000
SRM 1950-6_2 | 275000 | 211000 | 126000 | 225000 | 3750000 | 4400000 | 1190000 | 1410000
SRM 1950-6_3 | 268000 | 206000 | 119000 | 225000 | 3600000 | 4030000 | 1090000 | 1380000
Pooled QC 2 | 264000 | 204000 | 122000 | 228000 | 4100000 | 4420000 | 1230000 | 1380000
Pooled QC 2 | 255000 | 199000 | 122000 | 230000 | 3830000 | 4080000 | 1180000 | 1280000
T1D-1_1 | 253000 | 214000 | 143000 | 226000 | 3760000 | 4500000 | 1180000 | 1280000
T1D-1_2 | 245000 | 208000 | 148000 | 232000 | 3590000 | 4290000 | 1120000 | 1370000
T1D-1_3 | 244000 | 197000 | 146000 | 241000 | 3570000 | 4270000 | 1170000 | 1330000
T1D-2_1 | 256000 | 211000 | 140000 | 240000 | 3900000 | 4670000 | 1190000 | 1390000
T1D-2_2 | 243000 | 204000 | 142000 | 230000 | 3860000 | 4440000 | 1170000 | 1430000
T1D-2_3 | 246000 | 205000 | 148000 | 239000 | 3720000 | 4160000 | 1170000 | 1380000
T1D-3_1 | 247000 | 211000 | 141000 | 241000 | 3730000 | 4500000 | 1220000 | 1420000
T1D-3_2 | 245000 | 203000 | 139000 | 236000 | 3630000 | 4190000 | 1130000 | 1390000
T1D-3_3 | 248000 | 203000 | 147000 | 228000 | 3690000 | 4190000 | 1140000 | 1330000
T1D-4_1 | 249000 | 209000 | 137000 | 246000 | 3810000 | 4410000 | 1210000 | 1380000
T1D-4_2 | 248000 | 214000 | 140000 | 237000 | 3590000 | 4280000 | 1140000 | 1340000
T1D-4_3 | 248000 | 210000 | 134000 | 232000 | 3610000 | 4120000 | 1140000 | 1340000
T1D-5_1 | 249000 | 212000 | 147000 | 240000 | 3640000 | 4420000 | 1180000 | 1410000
T1D-5_2 | 249000 | 210000 | 139000 | 225000 | 3500000 | 4090000 | 1100000 | 1360000
T1D-5_3 | 248000 | 210000 | 138000 | 237000 | 3570000 | 4140000 | 1130000 | 1350000
T1D-6_1 | 261000 | 206000 | 149000 | 248000 | 3820000 | 4380000 | 1230000 | 1410000
T1D-6_2 | 254000 | 209000 | 146000 | 240000 | 3750000 | 4250000 | 1190000 | 1390000
T1D-6_3 | 251000 | 210000 | 141000 | 251000 | 3830000 | 4420000 | 1170000 | 1440000
Total Blank 2 | 4150 | 402 | 1220 | 300 | 21100 | 1010 | 3040 | 374
Pooled QC 3 | 261000 | 203000 | 124000 | 231000 | 4240000 | 4370000 | 1270000 | 1410000
Pooled QC 3 | 249000 | 194000 | 127000 | 226000 | 3820000 | 4130000 | 1140000 | 1350000
Young AA-1_1 | 185000 | 229000 | 78600 | 249000 | 2690000 | 4670000 | 782000 | 1410000
Young AA-1_2 | 188000 | 233000 | 75800 | 238000 | 2650000 | 4420000 | 749000 | 1410000
Young AA-1_3 | 191000 | 224000 | 75100 | 243000 | 2540000 | 4110000 | 770000 | 1450000
Young AA-2_1 | 191000 | 225000 | 73300 | 236000 | 2560000 | 4550000 | 811000 | 1450000
Young AA-2_2 | 186000 | 225000 | 71000 | 223000 | 2560000 | 4180000 | 778000 | 1500000
Young AA-2_3 | 187000 | 224000 | 75100 | 230000 | 2580000 | 4200000 | 732000 | 1360000
Young AA-3_1 | 174000 | 214000 | 71900 | 231000 | 2580000 | 4280000 | 767000 | 1360000
Young AA-3_2 | 183000 | 212000 | 66600 | 215000 | 2450000 | 3810000 | 732000 | 1350000
Young AA-3_3 | 181000 | 211000 | 69900 | 231000 | 2450000 | 3920000 | 744000 | 1360000
Young AA-4_1 | 184000 | 219000 | 73200 | 224000 | 2540000 | 4210000 | 739000 | 1360000
Young AA-4_2 | 175000 | 207000 | 64900 | 223000 | 2420000 | 3730000 | 745000 | 1300000
Young AA-4_3 | 188000 | 207000 | 68100 | 210000 | 2350000 | 3870000 | 713000 | 1310000
Young AA-5_1 | 182000 | 223000 | 72600 | 231000 | 2650000 | 4470000 | 761000 | 1430000
Young AA-5_2 | 182000 | 224000 | 75500 | 239000 | 2550000 | 4210000 | 758000 | 1370000
Young AA-5_3 | 185000 | 224000 | 69700 | 233000 | 2650000 | 4390000 | 795000 | 1420000
Young AA-6_1 | 187000 | 226000 | 72800 | 229000 | 2570000 | 4460000 | 773000 | 1400000
Young AA-6_2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Young AA-6_3 | 191000 | 222000 | 71100 | 224000 | 2470000 | 4100000 | 764000 | 1450000
Pooled QC 4 | 243000 | 195000 | 124000 | 233000 | 3710000 | 4130000 | 1220000 | 1330000
Pooled QC 4 | 243000 | 192000 | 113000 | 220000 | 3780000 | 3750000 | 1160000 | 1270000
high TAG-1_1 | 312000 | 179000 | 154000 | 238000 | 5920000 | 4400000 | 1830000 | 1380000
high TAG-1_2 | 304000 | 182000 | 148000 | 233000 | 5990000 | 4290000 | 1650000 | 1270000
high TAG-1_3 | 295000 | 183000 | 149000 | 210000 | 5900000 | 4160000 | 1680000 | 1350000
high TAG-2_1 | 293000 | 194000 | 154000 | 234000 | 5870000 | 4310000 | 1720000 | 1420000
high TAG-2_2 | 304000 | 193000 | 152000 | 230000 | 6010000 | 4250000 | 1740000 | 1360000
high TAG-2_3 | 314000 | 184000 | 152000 | 241000 | 5870000 | 4190000 | 1730000 | 1350000
high TAG-3_1 | 296000 | 197000 | 160000 | 239000 | 5990000 | 4500000 | 1720000 | 1380000
high TAG-3_2 | 304000 | 187000 | 146000 | 236000 | 5680000 | 4040000 | 1670000 | 1340000
high TAG-3_3 | 296000 | 189000 | 144000 | 224000 | 5900000 | 4150000 | 1620000 | 1360000
high TAG-4_1 | 296000 | 193000 | 152000 | 220000 | 6120000 | 4500000 | 1650000 | 1320000
high TAG-4_2 | 296000 | 189000 | 148000 | 226000 | 5850000 | 4020000 | 1650000 | 1340000
high TAG-4_3 | 294000 | 184000 | 138000 | 224000 | 5710000 | 3960000 | 1630000 | 1330000
high TAG-5_1 | 300000 | 188000 | 150000 | 230000 | 5790000 | 4360000 | 1660000 | 1320000
high TAG-5_2 | 305000 | 188000 | 143000 | 227000 | 5820000 | 4200000 | 1610000 | 1320000
high TAG-5_3 | 311000 | 186000 | 144000 | 231000 | 5860000 | 3950000 | 1620000 | 1300000
high TAG-6_1 | 294000 | 192000 | 157000 | 234000 | 6090000 | 4380000 | 1770000 | 1430000
high TAG-6_2 | 300000 | 195000 | 150000 | 226000 | 5700000 | 4180000 | 1640000 | 1340000
high TAG-6_3 | 290000 | 183000 | 147000 | 231000 | 5600000 | 4100000 | 1670000 | 1340000
Pooled QC 5 | 255000 | 208000 | 124000 | 227000 | 4010000 | 4350000 | 1240000 | 1420000
Pooled QC 5 | 245000 | 200000 | 121000 | 219000 | 3880000 | 4090000 | 1160000 | 1310000
STD1-4 | 4330000 | 204000 | 4580000 | 193000 | 70400000 | 5040000 | 57300000 | 1470000
STD1-5 | 4300000 | 200000 | 4780000 | 198000 | 69600000 | 5030000 | 57300000 | 1520000
STD1-6 | 4310000 | 200000 | 4860000 | 186000 | 67500000 | 5010000 | 59300000 | 1480000
STD2-4 | 2110000 | 199000 | 2180000 | 187000 | 31000000 | 4460000 | 26400000 | 1320000
STD2-5 | 2200000 | 201000 | 2390000 | 198000 | 32000000 | 4520000 | 26800000 | 1310000
STD2-6 | 2110000 | 203000 | 2360000 | 200000 | 32200000 | 4580000 | 27100000 | 1310000
STD3-4 | 231000 | 205000 | 255000 | 196000 | 2950000 | 4260000 | 2680000 | 1300000
STD3-5 | 224000 | 209000 | 249000 | 202000 | 2840000 | 4210000 | 2610000 | 1320000
STD3-6 | 219000 | 198000 | 252000 | 197000 | 3030000 | 4160000 | 2670000 | 1330000
STD4-4 | 50600 | 212000 | 46400 | 199000 | 616000 | 4140000 | 531000 | 1310000
STD4-5 | 46100 | 195000 | 50900 | 204000 | 566000 | 4070000 | 521000 | 1330000
STD4-6 | 84600 | 210000 | 74700 | 181000 | 710000 | 4110000 | 512000 | 1270000
STD5-4 | 24200 | 207000 | 23700 | 195000 | 299000 | 4010000 | 254000 | 1310000
STD5-5 | 4940 | 41700 | 7780 | 36600 | 73500 | 880000 | 55900 | 280000
STD5-6 | 27100 | 199000 | 27400 | 190000 | 297000 | 4080000 | 257000 | 1270000
STD6-4 | 20400 | 211000 | 20600 | 187000 | 250000 | 3940000 | 209000 | 1320000
STD6-5 | 18600 | 204000 | 22100 | 188000 | 248000 | 4230000 | 213000 | 1280000
STD6-6 | 19900 | 211000 | 18900 | 201000 | 250000 | 4010000 | 213000 | 1310000
LLQC-1_1 | 101000 | 228000 | 36100 | 210000 | 1250000 | 4090000 | 379000 | 1340000
LLQC-1_2 | 98600 | 223000 | 41300 | 208000 | 1290000 | 4150000 | 370000 | 1300000
LLQC-1_3 | 96500 | 239000 | 38700 | 214000 | 1230000 | 3870000 | 375000 | 1310000
LLQC-2_1 | 99000 | 234000 | 37400 | 214000 | 1310000 | 4190000 | 378000 | 1300000
LLQC-2_2 | 101000 | 239000 | 39800 | 222000 | 1240000 | 3990000 | 392000 | 1310000
LLQC-2_3 | 103000 | 240000 | 37800 | 207000 | 1270000 | 4140000 | 376000 | 1360000
LLQC-3_1 | 102000 | 230000 | 38400 | 211000 | 1240000 | 4250000 | 381000 | 1330000
LLQC-3_2 | 98900 | 226000 | 36000 | 196000 | 1280000 | 3830000 | 379000 | 1260000
LLQC-3_3 | 99000 | 230000 | 36100 | 210000 | 1240000 | 4000000 | 377000 | 1280000
LLQC-4_1 | 101000 | 242000 | 40600 | 215000 | 1320000 | 4320000 | 384000 | 1380000
LLQC-4_2 | 99700 | 222000 | 36800 | 212000 | 1240000 | 4150000 | 399000 | 1340000
LLQC-4_3 | 100000 | 226000 | 39900 | 215000 | 1280000 | 4090000 | 388000 | 1320000
LLQC-5_1 | 100000 | 241000 | 39000 | 212000 | 1260000 | 4140000 | 392000 | 1370000
LLQC-5_2 | 98600 | 233000 | 37600 | 208000 | 1220000 | 4160000 | 366000 | 1320000
LLQC-5_3 | 92100 | 234000 | 37200 | 217000 | 1200000 | 4020000 | 369000 | 1260000
LLQC-6_1 | 95500 | 242000 | 40800 | 209000 | 1280000 | 4400000 | 395000 | 1390000
LLQC-6_2 | 103000 | 235000 | 36600 | 223000 | 1230000 | 4190000 | 369000 | 1310000
LLQC-6_3 | 101000 | 229000 | 38500 | 217000 | 1190000 | 4100000 | 370000 | 1330000
LQC-1_1 | 148000 | 235000 | 61900 | 217000 | 1940000 | 4410000 | 587000 | 1390000
LQC-1_2 | 140000 | 218000 | 53700 | 217000 | 1910000 | 4160000 | 583000 | 1300000
LQC-1_3 | 152000 | 232000 | 58300 | 219000 | 1820000 | 4090000 | 583000 | 1390000
LQC-2_1 | 145000 | 231000 | 54300 | 221000 | 1920000 | 4360000 | 597000 | 1330000
LQC-2_2 | 145000 | 235000 | 61400 | 213000 | 1850000 | 3910000 | 581000 | 1350000
LQC-2_3 | 134000 | 220000 | 55500 | 216000 | 1860000 | 4150000 | 558000 | 1350000
LQC-3_1 | 143000 | 233000 | 56500 | 211000 | 1870000 | 4240000 | 578000 | 1310000
LQC-3_2 | 142000 | 224000 | 56100 | 208000 | 1860000 | 4270000 | 592000 | 1360000
LQC-3_3 | 141000 | 223000 | 57200 | 213000 | 1890000 | 4120000 | 571000 | 1390000
LQC-4_1 | 150000 | 236000 | 59600 | 221000 | 1970000 | 4340000 | 602000 | 1390000
LQC-4_2 | 147000 | 228000 | 55200 | 211000 | 1960000 | 4130000 | 567000 | 1340000
LQC-4_3 | 144000 | 232000 | 57300 | 223000 | 1870000 | 4010000 | 580000 | 1340000
LQC-5_1 | 147000 | 229000 | 58500 | 216000 | 1990000 | 4380000 | 600000 | 1350000
LQC-5_2 | 142000 | 218000 | 58200 | 223000 | 1860000 | 4040000 | 563000 | 1390000
LQC-5_3 | 142000 | 237000 | 57500 | 217000 | 1930000 | 4410000 | 587000 | 1330000
LQC-6_1 | 147000 | 238000 | 63400 | 230000 | 1990000 | 4440000 | 606000 | 1370000
LQC-6_2 | 144000 | 227000 | 61100 | 229000 | 1870000 | 4170000 | 553000 | 1350000
LQC-6_3 | 140000 | 232000 | 59200 | 220000 | 1860000 | 4050000 | 561000 | 1370000
MQC-1_1 | 263000 | 216000 | 120000 | 227000 | 3800000 | 4250000 | 1200000 | 1420000
MQC-1_2 | 252000 | 204000 | 112000 | 216000 | 3690000 | 4040000 | 1090000 | 1270000
MQC-1_3 | 252000 | 200000 | 120000 | 225000 | 3550000 | 3940000 | 1130000 | 1300000
MQC-2_1 | 264000 | 205000 | 130000 | 238000 | 3910000 | 4520000 | 1210000 | 1450000
MQC-2_2 | 253000 | 206000 | 116000 | 226000 | 3810000 | 4160000 | 1170000 | 1370000
MQC-2_3 | 256000 | 209000 | 127000 | 229000 | 3960000 | 4270000 | 1180000 | 1390000
MQC-3_1 | 263000 | 206000 | 127000 | 214000 | 4060000 | 4340000 | 1220000 | 1350000
MQC-3_2 | 246000 | 199000 | 115000 | 224000 | 3980000 | 4130000 | 1210000 | 1330000
MQC-3_3 | 257000 | 198000 | 118000 | 225000 | 3810000 | 3960000 | 1180000 | 1310000
MQC-4_1 | 251000 | 204000 | 123000 | 220000 | 4000000 | 4280000 | 1220000 | 1350000
MQC-4_2 | 245000 | 203000 | 118000 | 211000 | 3670000 | 3910000 | 1180000 | 1350000
MQC-4_3 | 247000 | 201000 | 119000 | 225000 | 3900000 | 4130000 | 1200000 | 1310000
MQC-5_1 | 256000 | 199000 | 118000 | 222000 | 3940000 | 4460000 | 1200000 | 1390000
MQC-5_2 | 255000 | 204000 | 122000 | 217000 | 3830000 | 4110000 | 1150000 | 1300000
MQC-5_3 | 257000 | 200000 | 126000 | 223000 | 4020000 | 4070000 | 1190000 | 1350000
MQC-6_1 | 254000 | 203000 | 125000 | 221000 | 4020000 | 4270000 | 1210000 | 1350000
MQC-6_2 | 237000 | 195000 | 125000 | 221000 | 3970000 | 4110000 | 1150000 | 1300000
MQC-6_3 | 251000 | 194000 | 130000 | 221000 | 3800000 | 3930000 | 1110000 | 1310000
HQC-1_1 | 355000 | 213000 | 263000 | 221000 | 5300000 | 4460000 | 2700000 | 1370000
HQC-1_2 | 353000 | 202000 | 254000 | 215000 | 5240000 | 3980000 | 2450000 | 1350000
HQC-1_3 | 355000 | 202000 | 240000 | 220000 | 5210000 | 4040000 | 2480000 | 1360000
HQC-2_1 | 354000 | 200000 | 263000 | 229000 | 5490000 | 4310000 | 2560000 | 1340000
HQC-2_2 | 347000 | 202000 | 241000 | 217000 | 5140000 | 3930000 | 2500000 | 1280000
HQC-2_3 | 356000 | 201000 | 246000 | 216000 | 4990000 | 3860000 | 2350000 | 1360000
HQC-3_1 | 350000 | 205000 | 259000 | 232000 | 5260000 | 4320000 | 2480000 | 1340000
HQC-3_2 | 347000 | 195000 | 251000 | 216000 | 5130000 | 4090000 | 2390000 | 1390000
HQC-3_3 | 349000 | 204000 | 248000 | 215000 | 5230000 | 4150000 | 2420000 | 1300000
HQC-4_1 | 365000 | 216000 | 253000 | 225000 | 5630000 | 4550000 | 2600000 | 1390000
HQC-4_2 | 366000 | 225000 | 258000 | 229000 | 5340000 | 4060000 | 2470000 | 1350000
HQC-4_3 | 352000 | 205000 | 254000 | 225000 | 5280000 | 4130000 | 2490000 | 1290000
HQC-5_1 | 369000 | 205000 | 265000 | 222000 | 5390000 | 4360000 | 2600000 | 1340000
HQC-5_2 | 349000 | 200000 | 252000 | 213000 | 5210000 | 3970000 | 2450000 | 1260000
HQC-5_3 | 361000 | 199000 | 248000 | 223000 | 4920000 | 3780000 | 2470000 | 1300000
HQC-6_1 | 355000 | 207000 | 247000 | 217000 | 5380000 | 4180000 | 2510000 | 1370000
HQC-6_2 | 355000 | 201000 | 240000 | 222000 | 5030000 | 3890000 | 2380000 | 1290000
HQC-6_3 | 348000 | 199000 | 246000 | 207000 | 4960000 | 3850000 | 2490000 | 1310000
HLQC-1_1 | 462000 | 207000 | 401000 | 224000 | 6900000 | 4400000 | 4020000 | 1390000
HLQC-1_2 | 462000 | 204000 | 397000 | 228000 | 6690000 | 3940000 | 3790000 | 1320000
HLQC-1_3 | 463000 | 204000 | 380000 | 231000 | 6700000 | 4020000 | 3810000 | 1280000
HLQC-2_1 | 468000 | 211000 | 393000 | 216000 | 6950000 | 4500000 | 3830000 | 1370000
HLQC-2_2 | 465000 | 202000 | 388000 | 221000 | 6430000 | 4030000 | 3820000 | 1290000
HLQC-2_3 | 461000 | 205000 | 379000 | 220000 | 6810000 | 4040000 | 3750000 | 1320000
HLQC-3_1 | 483000 | 215000 | 404000 | 229000 | 7080000 | 4340000 | 4080000 | 1440000
HLQC-3_2 | 469000 | 209000 | 407000 | 221000 | 6770000 | 4160000 | 3980000 | 1370000
HLQC-3_3 | 460000 | 208000 | 403000 | 220000 | 6860000 | 3970000 | 3890000 | 1330000
HLQC-4_1 | 481000 | 205000 | 392000 | 229000 | 7140000 | 4250000 | 4000000 | 1360000
HLQC-4_2 | 478000 | 195000 | 388000 | 221000 | 7030000 | 4160000 | 3910000 | 1320000
HLQC-4_3 | 481000 | 196000 | 405000 | 228000 | 6930000 | 3970000 | 3760000 | 1320000
HLQC-5_1 | 538000 | 207000 | 445000 | 225000 | 7020000 | 4410000 | 4030000 | 1340000
HLQC-5_2 | 572000 | 198000 | 473000 | 213000 | 6910000 | 3890000 | 3690000 | 1260000
HLQC-5_3 | 561000 | 209000 | 455000 | 212000 | 6650000 | 3950000 | 3710000 | 1280000
HLQC-6_1 | 467000 | 201000 | 410000 | 231000 | 6910000 | 4360000 | 4150000 | 1400000
HLQC-6_2 | 448000 | 206000 | 355000 | 223000 | 6740000 | 3890000 | 3650000 | 1280000
HLQC-6_3 | 462000 | 208000 | 395000 | 229000 | 6740000 | 4220000 | 3930000 | 1380000
Matrix Blank | 262000 | 419 | 127000 | 357 | 4780000 | 1190 | 1440000 | 536
Total Blank 3 | 3830 | 267 | 2180 | 354 | 19000 | 472 | 2520 | 528