## Participant Information
Participant Information
Participant Information |  | Group Name
|  | Lab code | 38
| Institution | Institution name
Street
City
Postal code
Country
| Contact person | Name
Address
Email
Phone

## Summary
Summary
Summary | This protocol describes the extraction and LC-MRM procedure for absolute quantification of Cer d18:1/16:0, Cer d18:1/18:0, Cer d18:1/24:0 and Cer d18:1/24:1 in human plasma according to Kauhanen et al, Anal Bioanal Chem (2016) 408:3475–3483.
A calibration curve is built by injecting several dilutions of the non-labelled standards mixed with a fixed concentration of labelled standards. Each calibration point is prepared and analysed in triplicates (n=3).
To establish the endogenous concentration of the four selected ceramides in the study samples (NIST SRM1950, NIST high TAG, NIST T1D, and NIST young A-A) each sample is extracted and analyzed in six replicates (n =6).
Prepare the pooled QC by mixing 100 mL each of NIST SRM1950, NIST high TAG, NIST T1D, and NIST young AA (total 400 mL).
We suggest proceeding first to measure the ceramides concentration in the four study samples and then continue with the intra-assay validation to avoid wasting material.

## Extraction - Overview
Extraction Overview
Extraction Overview | Table 1 - Extraction - Materials
Chemicals, solvents and materials
Ethyl acetate (EtAc)
2-propanol (IPA)
5% Bovine Serum Albumin (BSA)
96-well plates/tubes
Sealing foil
2 mL DeepWell plates/tubes/LC vials
Table 2 - Extraction - Sample Types
| Sample type | Components
| Pooled Quality Control samples (QC) | 10 µL pooled QC |  |  | 20 µL Labelled IS |  | 570 µL EtAc:IPA 2:8 (v/v)
| Matrix Blank (MB) | 10 µL Matrix (pooled QC) only |  |  | - |  | 590 µL EtAc:IPA 2:8
| Study Samples (S) | 10 µL Human plasma (NIST SRM1950, high TAG, T1D, young AA) |  |  | 20 µL Labelled IS |  | 570 µL EtAc:IPA 2:8
| Total Blank (TB) | 10 µL 5% BSA |  |  | - |  | 590 µL EtAc:IPA 2:8
| Calibration line standards (STD) | 10 µL 5% BSA + 20 µL non labelled STD mixture (Table 5) |  |  | 20 µL Labelled IS |  | 550 µL EtAc:IPA 2:8
Table 3 - Extraction - Labelled Internal Standard Mixture Solution (provided)
| Labelled IS mixture in EtAc:IPA 2:8 (v/v) | Concentration |  |  |  |  | Unit | Vendor
| D7-Cer d18:1/16:0 | 0.125 |  |  |  |  | pmol/µL [µM] | Avanti Polar Lipids
| D7-Cer d18:1/18:0 | 0.050
| D7-Cer d18:1/24:0 | 1500
| D7-Cer d18:1/24:1 | 0.500
Table 4 - Extraction - Non-Labelled Standards - for STD1 Mixture Solution (provided)
| Non-labelled standards in EtAc:IPA 2:8 (v/v) | Concentration |  |  |  |  | Unit | Vendor
| Cer d18:1/16:0 | 2.0 |  |  |  |  | pmol/µL [µM] | Avanti Polar Lipids
| Cer d18:1/18:0 | 2.0
| Cer d18:1/24:0 | 20.0
| Cer d18:1/24:1 | 20.0
Table 5 - Serial dilutions of Non-Labelled Standards from STD1 Mixture Solution for calibration line
| Add 20 µL of | STD1 Conc. | STD2 Conc. | STD3 Conc. | STD4 Conc. | STD5 Conc. | STD6 Conc. | Unit
| Cer d18:1/16:0 | 2.0 | 1.0 | 0.10 | 0.02 | 0.01 | 0.008 | pmol/µL [µM]
| Cer d18:1/18:0 | 2.0 | 1.0 | 0.10 | 0.02 | 0.01 | 0.008
| Cer d18:1/24:0 | 20.0 | 10.0 | 1.0 | 0.2 | 0.1 | 0.08
| Cer d18:1/24:1 | 20.0 | 10.0 | 1.0 | 0.2 | 0.1 | 0.08

## Extraction - Procedure
Extraction Procedure
Extraction Procedure | Extraction Procedure
| Step | Description | Reference Table
| 1 | Prepare different dilutions of STD1 by diluting the provided mixture (Extraction - Overview, Table 4) with EtAc:IPA 2:8 as suggested in table | Extraction - Overview'!B34
| 2 | Thaw plasma on ice (+4°C) and bring to room temperature prior to extraction. Mix by pipetting before use.
| 3 | Prepare samples in plates/tubes manually or with a robot, according to specifications in column C of table | Extraction - Overview'!C14
| 4 | Add 20 µL of labelled internal standard mixture to samples according to column F of table | Extraction - Overview'!F14
| 5 | For the calibration line, add 20 µL of different dilutions of the non-labelled mixture (STD1-6).
| 6 | Add 590 µL of ethyl acetate:isopropanol (2:8, vol/vol) to Total Blank (TB) and Matrix Blank (MB); 550 µL to calibration line (STD); 570 µL to study samples and QCs according to column H of table | Extraction - Overview'!H14
| 7 | Mix samples 10 min by pipetting or vortexing
| 8 | Centrifuge samples for 10 min at 3000 x gat room temperature
| 9 | Transfer 50 µL of the clear supernatant to MS-vials or MS plates, close/seal, store samples at -20 °C prior to LC-MRM analysis.

## LC-MRM - Overview
LC-MRM Overview
LC-MRM Overview | Table 6 Chemicals and solvents
Ammonium formate
Formic acid
Ultra-pure water
Acetonitrile
Table 7 LC-MRM LC-MS Details
| Property | Specifications
| Mass spectrometry | QQQ Waters Xevo TQ-XS
| LC | UPLC
| Injection volume (indicative) | 5 µL
| Purge | 90:10 water:ACN
| Strong wash | Acetonitrile
| Pre-column | Waters Acquity BEH C18, 1.7 µm VanGuard Pre-Column
| Column | Waters Acquity BEH C18, 2.1 × 50 mm id. 1.7 µm
| Column temperature | 60 °C
| Mobile phase A | 10 mM ammonium acetate in water with 0.1% formic acid
| Mobile phase B | 10 mM ammonium acetate in acetonitrile:2-propanol (4:3, vol/vol) with 0.1% formic acid
Table 8 LC-MRM LC gradient
| Time (minutes) | A% | B% |  | Flow (µL/min)
| 0.00 | 15 | 85 |  | 500
| 0.50 | 15 | 85 |  | 500
| 1.50 | 0 | 100 |  | 500
| 4.00 | 0 | 100 |  | 500
| 4.10 | 15 | 85 |  | 500
| 5.00 | 15 | 85 |  | 500
Table 9 LC-MRM MS conditions (EXAMPLE: this only applies to Sciex QTRAP, please adapt)
| Polarity mode | Positive
| Capillary | 3500 V
| Cone | 35 V
| Source temperature | 120°C
| Desolvation gas (l/hr) | 1200
| Desolvation temp (°C) | 600
| Cone gas (l/h) | 150
| Nebuliser (Bar) | 6.8
| Collision cell exit potential | n/a
| Collision energy (CE) | 30 eV
Table 10 LC-MRM Mass transitions used for absolute quantification. On different systems, RT and CE might be different!
| Analyte | Mass transitions | IS | Mass transitions | RT (min) | CE
| Cer d18:1/16:0 | 538.5 – 264.3 | D7- Cer d18:1/16:0 | 545.6 - 271.3 | 1.9-2.0 | 30
C34H68NO3 – C18H34N
| Cer d18:1/18:0 | 566.6 – 264.3 | D7- Cer d18:1/18:0 | 573.6 – 271.3 | 2.1-2.2 | 30
C36H72NO3 – C18H34N
| Cer d18:1/24:0 | 650.6 – 264.3 | D7- Cer d18:1/24:0 | 657.7 – 271.3 | 2.4-2.5 | 30
C42H84NO3 – C18H34N
| Cer d18:1/24:1 | 648.6 – 264.3 | D7- Cer d18:1/24:1 | 655.7 – 271.3 | 2.3-2.4 | 30
C42H82NO3 – C18H34N
Table 11 Data processing
| Processing software | (Add here name and version of the software used)

## LC-MRM - Sample Sequence
Sample Sequence (Example)
Sample Sequence (Example) | Sample Sequence 1
| Number | Sample Name | Remarks
| 1 | Total Blank 1
| 2 | Matrix blank
| 3 | Calibration line 1 from STD6 to 1 (n=3)
| 4 | […]
| 5 | Pooled QC
| 6 | Pooled QC
| 7 | SRM 1950-1
| 8 | SRM 1950-2
| 9 | SRM 1950-3
| 10 | SRM 1950-4
| 11 | SRM 1950-5
| 12 | SRM 1950-6
| 13 | Pooled QC
| 14 | Pooled QC
| 15 | T1D-1
| 16 | T1D-2
| 17 | T1D-3
| 18 | T1D-4
| 19 | T1D-5
| 20 | T1D-6
| 21 | Total Blank 2
| 22 | Pooled QC
| 23 | Pooled QC
| 24 | Young AA-1
| 25 | Young AA-2
| 26 | Young AA-3
| 27 | Young AA-4
| 28 | Young AA-5
| 29 | Young AA-6
| 30 | Pooled QC
| 31 | Pooled QC
| 32 | hTAG-1
| 33 | hTAG-2
| 34 | hTAG-3
| 35 | hTAG-4
| 36 | hTAG-5
| 37 | hTAG-6
| 38 | Pooled QC
| 39 | Pooled QC
| 40 | Calibration line 2 from STD1 to 6 (n=3)
| 41 | Intra-assay QCs: LLQC, LQC, MQC, HQC
| 42 | Matrix blank
| 43 | Total Blank 3

## LC-MRM - Intra Assay QC Proc
Intra Assay Procedure
Intra Assay Procedure | Quality Controls for Intra-Assay Variation
| Step | Description | Reference Table
| 1 | Prepare the Pooled NIST QC sample containing equal volumes of the four NIST reference plasma you received.
| 2 | The pooled QC representing middle quality control (MQC), low QC (LQC), and lower limit of quantification QC (LLOQ) is prepared by diluting MQC (e.g. two- and three-fold) with water.
| 3 | High QC (HQC) and upper limit of quantification (ULOQ) QC are prepared by spiking in 10 µL and 20 µL respectively of STD3 endogenous standards from table | Extraction - Overview'!B34
| 4 | Internal standards are added to all QC samples.
Intra-Assay Variation
| 6 | The precision and accuracy of the assay is determined for each of the four ceramides measured in LLOQ, LQC, MQC, HQC, and HLOQ.
| 7 | Intra-assay variation, precision, and accuracy is calculated for each of the ceramide in replicates of six (n = 6) at each QC concentration.
| 8 | The intra-assay precision (percentage coefficient variance, %CV) and accuracy (percentage accuracy, %Accuracy) are calculated from the nominal concentrations according to the formulas reported in corresponding sheets and tables of this document.
| 9 | Recovery estimation is not necessary for the published method but only for different ones. Stability estimation is not necessary for the published method but only for different ones. Calibration line is prepared in 5% BSA.

## LC-MRM - Intra Assay QC Res
INTRA ASSAY - QC RESULTS
INTRA ASSAY - QC RESULTS | Table 12 LC-MRM Calibration Line 1 Data (Beginning). Please report the values of each of the three replicates separately for each compound and standard. Each STD prepared in triplicate, injected once from each vial
| Calibration Line 1 | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| STD 1 | 161670.266 | 188843.281 | 138922.375 | 431775.375 | 492482.344 | 366987.5 | 8043190 | 8987714 | 6841481.5 | 6912193.5 | 7731415 | 5683858 | 13193.428 | 12360.023 | 9367.42 | 29792.49 | 27829.625 | 22106 | 945122.75 | 855688.688 | 686402.188 | 288977 | 275952.594 | 217880.781
| STD 2 | 162920.375 | 99568.086 | 109124.891 | 435999.625 | 265772.406 | 290046.219 | 8146007 | 4970327 | 5663704 | 7098589 | 4277954.5 | 4780396.5 | 19071.592 | 12869.571 | 15976.856 | 43532.461 | 29771.764 | 38288.773 | 1374526.75 | 953609.563 | 1244105.875 | 443021.406 | 303883.344 | 396407.094
| STD 3 | 13680.958 | 8577.718 | 6304.505 | 34671.465 | 24212.867 | 18870.303 | 629642 | 446231.75 | 358521.594 | 558913.938 | 394434.656 | 310575.531 | 17600.848 | 13390.945 | 10123.553 | 38770.746 | 29789.727 | 22531.334 | 1185097.5 | 923869.75 | 709416.938 | 388887.688 | 283552 | 211377.344
| STD 4 | 1973.699 | 1449.058 | 633.647 | 5899.858 | 4888.85 | 2934.986 | 119461.469 | 90872.859 | 67757.625 | 105216.063 | 80641.789 | 58071.98 | 16113.021 | 9719.227 | 9837.344 | 33369.02 | 23841.447 | 23040.943 | 1026804.875 | 719379.625 | 728291.313 | 326629.094 | 230239.844 | 229324.578
| STD 5 | 940.139 | 467.968 | 296.453 | 2485.385 | 2164.195 | 1493.584 | 47810.879 | 40317.414 | 36760.246 | 42945.367 | 37083.055 | 35143.426 | 14261.565 | 7564.1 | 8597.471 | 30040.352 | 17709.107 | 19285.578 | 887241.813 | 540758.75 | 583021.063 | 288003.688 | 175812.063 | 189350.922
| STD 6 | 2639.93 | 1538.747 | 819.555 | 7274.613 | 5125.318 | 2907.305 | 133056.453 | 98550.797 | 60979.008 | 115961.359 | 90090.375 | 55806.691 | 36720.559 | 27537.49 | 15694.344 | 78320.758 | 60840.969 | 36304.375 | 2321766.25 | 1749863.75 | 1065339.75 | 776035.75 | 595177.188 | 358447.313
Table 13 LC-MRM Calibration Line 2 Data (End). Please report the values of each of the three replicates separately for each compound and standard. Each STD prepared in triplicate, injected once from each vial
| Calibration Line 2 | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| STD 1 | 277289.563 | 255814.391 | 203670.125 | 700361.875 | 669241 | 568867.25 | 14575619 | 13329707 | 11794922 | 11897055 | 11154217 | 9360104 | 16355.622 | 15111.695 | 13766.348 | 35305.082 | 35629.559 | 35497.438 | 1267685.375 | 1234273.875 | 1232357.625 | 381619.906 | 372679.438 | 360469.656
| STD 2 | 347266.25 | 315365.719 | 348983.531 | 980989.75 | 901498.625 | 1011734.063 | 21319064 | 18968978 | 21518826 | 16963404 | 15407429 | 17533704 | 36673.98 | 31191.121 | 34747.746 | 91389.188 | 77063.336 | 89799.391 | 3265680.25 | 2830980.25 | 3223640.75 | 999383.063 | 844960.625 | 967858.313
| STD 3 | 14533.249 | 10892.807 | 27592.035 | 41671.93 | 31974.789 | 70192.281 | 894912 | 674329.063 | 1527557.625 | 721902.875 | 544451.5 | 1157965.25 | 14343.398 | 12222.679 | 20205.186 | 34841.352 | 30053.99 | 50203.668 | 1241256 | 1062834.5 | 1665800.75 | 369569.781 | 316439.438 | 516793.719
| STD 4 | 886.561 | 1214.582 | 1273.573 | 3637.833 | 4593.059 | 5198.683 | 82246.516 | 96773.281 | 111277.211 | 70413.203 | 79030.664 | 90581.273 | 11251.16 | 10266.151 | 10747.004 | 26524.496 | 24677.238 | 27329.666 | 922217 | 866045.063 | 955618.625 | 271849.75 | 256217.266 | 284815.375
| STD 5 | 584.332 | 784.275 | 937.273 | 1909.491 | 2837.258 | 2945.998 | 52792.43 | 67830.656 | 68206.016 | 45772.688 | 55226.453 | 52085.949 | 13211.278 | 13242.304 | 11357.174 | 30803.756 | 30896.357 | 28448.365 | 1071927.25 | 1050741.125 | 994547.313 | 316065.594 | 319296.375 | 295500.438
| STD 6 | 2019.086 | 1946.482 | 2514.107 | 7740.59 | 6290.741 | 7113.292 | 164199.766 | 143158.172 | 148346.438 | 135193.359 | 120758.313 | 115555.281 | 30963.518 | 30014.457 | 34929.848 | 84421.281 | 71379.758 | 84000.766 | 2913704.25 | 2568921.75 | 2921018.75 | 851272.438 | 768740.75 | 866205.25
Table 14 LC-MRM Sample Data. Please report the values of each of the three replicates separately for each compound and standard. Each vial injected 3 times
| NIST SRM | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| SRM 1 | 15436.854 | 16314.774 | 14986.483 | 14500.686 | 14473.391 | 15033.943 | 681805.5 | 677462.75 | 684014.938 | 224244.469 | 223780.453 | 227862.234 | 65022.191 | 64518.461 | 64145.637 | 151160.313 | 151860.813 | 153184.578 | 4341372.5 | 4312498 | 4414692 | 1392100.5 | 1408901.75 | 1430026.625
| SRM 2 | 13104.801 | 12761.136 | 13306.455 | 12462.881 | 11995.347 | 11964.05 | 596229.688 | 600847.438 | 600883.375 | 198338.438 | 196484.156 | 200121.906 | 45447.992 | 45167.207 | 45271.977 | 106584.813 | 107404.82 | 106785.172 | 3138495.25 | 3231560.25 | 3210960.75 | 1025875.438 | 1052141.125 | 1052931.125
| SRM 3 | 11630.096 | 12178.446 | 12149.92 | 11231.265 | 11200.254 | 11399.455 | 536108 | 552527.375 | 552447.875 | 181562.531 | 180861.5 | 183923.594 | 47679.172 | 48279.688 | 47575.801 | 114303.508 | 111957.086 | 113698.469 | 3396372.5 | 3389653.25 | 3396195.75 | 1094107.625 | 1100481.25 | 1117860.75
| SRM 4 | 12687.581 | 12242.174 | 12127.851 | 10719.236 | 11039.733 | 10896.063 | 519237.625 | 521329.594 | 526884 | 171941.953 | 172139.078 | 171500.188 | 66290.953 | 66059.961 | 64484.008 | 147548.516 | 145785.234 | 145517.406 | 4454893.5 | 4386499.5 | 4443964.5 | 1455284.625 | 1438925.125 | 1424130.875
| SRM 5 | 12537.665 | 13009.723 | 11847.886 | 11580.809 | 11507.396 | 11360.44 | 539459.5 | 544855.938 | 555169.063 | 177839.063 | 180934.813 | 182172.469 | 61886.465 | 61876.984 | 63271.77 | 141601.109 | 144288.984 | 144576.672 | 4169094.25 | 4344303 | 4405437.5 | 1379458.75 | 1426364.375 | 1395926.375
| SRM 6 | 14234.146 | 14364.467 | 14327.737 | 13601.656 | 13476.734 | 13838.013 | 680419.188 | 692953.625 | 701439.063 | 216232.906 | 222140.609 | 221483.75 | 54015.633 | 52911.805 | 54039.949 | 128640.758 | 129781.938 | 131661.922 | 3956682.75 | 4068984.25 | 4052581.75 | 1241749.125 | 1276584.25 | 1296079.625
| NIST hTAG | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| hTAG 1 | 44652.141 | 45517.219 | 45129.336 | 47025.496 | 47240.508 | 46947.801 | 2610320 | 2595654.5 | 2605426.5 | 786314.688 | 777633.25 | 791885.875 | 83913.664 | 85090.391 | 84928.813 | 218882.219 | 220236.469 | 220604.75 | 6246800 | 6205797 | 6241036 | 1986632.125 | 1991988.875 | 1998425.75
| hTAG 2 | 29460.719 | 30770.299 | 29673.1 | 31535.877 | 32624.186 | 32892.641 | 1836938.5 | 1899222.5 | 1912866.125 | 553689.5 | 566766.875 | 564470.625 | 70247.281 | 69570.063 | 70397.828 | 202326.031 | 202238.609 | 204421.172 | 5820052.5 | 5885250.5 | 5957444 | 1852259.5 | 1899622.125 | 1872637
| hTAG 3 | 29454.91 | 29349.404 | 28824.598 | 31397.459 | 30707.348 | 30842.252 | 1803000.625 | 1807376.625 | 1803893.625 | 523551.594 | 528989.5 | 534711.938 | 75257.297 | 74506.633 | 74876.609 | 211177.031 | 209880.984 | 212289.781 | 6332020.5 | 6390343.5 | 6292947 | 1972418.375 | 1986380.75 | 2028472.25
| hTAG 4 | 36212.043 | 36818.883 | 36837.238 | 37206.805 | 37126.254 | 36960.918 | 2124340.75 | 2128993.25 | 2132806.75 | 646113.625 | 635747.875 | 645196.875 | 80454.641 | 82028.547 | 83778.914 | 217845.281 | 220091.75 | 220511.797 | 6484776.5 | 6444215.5 | 6358559 | 2036319.125 | 2015724.75 | 2023588
| hTAG 5 | 30079.861 | 29434.121 | 30090.055 | 30211.896 | 29707.402 | 29400.451 | 1678585 | 1702684.375 | 1714143 | 500639.5 | 507014.625 | 504837.75 | 81222.531 | 80021.359 | 82054.398 | 213995.438 | 214363.922 | 216528.547 | 6467977 | 6426290 | 6436785 | 2056280.75 | 2038441.375 | 2057209.25
| hTAG 6 | 23444.48 | 23542.252 | 23469.393 | 24017.633 | 24286.73 | 24644.158 | 1356627.75 | 1380889.5 | 1386984.25 | 409790.906 | 412467.313 | 416608.219 | 70214.563 | 69976.602 | 71164.398 | 199561.734 | 197658.453 | 201788.438 | 6077615.5 | 6147816.5 | 6174492 | 1904866.875 | 1940109.25 | 1922105.375
| NIST T1D | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| T1D 1 | 14161.296 | 14397.564 | 14406.065 | 17747.752 | 17411.156 | 18351.234 | 742592.375 | 762276.625 | 758625.563 | 247402.656 | 242761.141 | 241558.594 | 56261.785 | 57254.949 | 57748.434 | 139005.906 | 136908.625 | 138661.297 | 4055878.75 | 3987511 | 4017309.25 | 1317272.75 | 1310546.625 | 1315208.625
| T1D 2 | 9518.183 | 9495.016 | 9483.618 | 11294.468 | 11380.858 | 11498.029 | 501083.813 | 509048.188 | 510878.75 | 163424.703 | 163596.359 | 164701.375 | 45833.945 | 46876.441 | 47228.215 | 109346.484 | 110182.289 | 110911.039 | 3332506.25 | 3407169.75 | 3456803.75 | 1088271.5 | 1091701.75 | 1110566.625
| T1D 3 | 13140.48 | 12792.008 | 12754.989 | 15325.131 | 15414.313 | 15853.961 | 653085.75 | 654023 | 659736.188 | 217421.438 | 213144.719 | 215131.016 | 50490.781 | 50856.609 | 50877.453 | 120801.672 | 120864.211 | 122842.555 | 3654831.75 | 3642395 | 3674864.5 | 1180001.875 | 1160410.5 | 1177917.5
| T1D 4 | 11986.611 | 11862.257 | 11765.321 | 13646.484 | 12906.147 | 13779.42 | 583297.688 | 572196.75 | 569588.438 | 189285.828 | 189656.828 | 186851.734 | 45057.422 | 44912.051 | 44584.047 | 100924.516 | 100722.695 | 100224.813 | 2962037.75 | 2940367.75 | 2998590.25 | 969644.813 | 959750.5 | 968603.688
| T1D 5 | 10442.962 | 10596.662 | 10833.356 | 12636.963 | 12340.86 | 12948.35 | 529280.813 | 545595.938 | 550630.813 | 174805.438 | 176898.266 | 178246.688 | 48193.25 | 47130.766 | 46869.039 | 112195.266 | 111242.828 | 112755.992 | 3328376.75 | 3440922.75 | 3438498.75 | 1094013 | 1097847.75 | 1100645.5
| T1D 6 | 7923.998 | 8416.837 | 8620.573 | 10360.113 | 10212.473 | 10213.71 | 438901.906 | 431897.188 | 438276.844 | 142017.266 | 144094.406 | 144796.875 | 59848.52 | 61268.84 | 61921.23 | 144518.75 | 145031.813 | 146930.469 | 4463162.5 | 4581243.5 | 4509406 | 1422120.375 | 1437266.875 | 1441855.75
| NIST Young AA | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| yA-A 1 | 18170.344 | 18974.5 | 19668.578 | 16782.332 | 17065.184 | 17147.867 | 914714.313 | 912233.25 | 897802.313 | 273145.75 | 270449.094 | 270825.688 | 74743.992 | 75085.891 | 75517.203 | 183911.484 | 184805.188 | 185306.531 | 5195973.5 | 5183351.5 | 5206968.5 | 1694428.375 | 1688717.75 | 1704857.875
| yA-A 2 | 6231.271 | 6230.36 | 6095.92 | 5671.094 | 6027.101 | 5984.772 | 333246.156 | 340202.938 | 345186.531 | 100858.836 | 102288.617 | 103882.672 | 63004.539 | 62231.629 | 63180.504 | 151965.078 | 151569.781 | 150679.938 | 4606121 | 4696746 | 4729148 | 1503073.625 | 1530413.875 | 1530609.375
| yA-A 3 | 9915.622 | 10408.454 | 10346.612 | 9563.206 | 9654.504 | 9563.626 | 521022.313 | 532983 | 530973.063 | 152300.75 | 152556.094 | 155066.859 | 58546.117 | 60327.984 | 60132.367 | 151101.078 | 152567.688 | 152877.156 | 4832932 | 4906300 | 4951557.5 | 1497946.625 | 1500713.5 | 1533961
| yA-A 4 | 5739.433 | 5974.818 | 5467.194 | 5382.24 | 4979.066 | 5006.184 | 277700.656 | 272699.156 | 274479.344 | 81820.461 | 80373.602 | 81159.359 | 57507.973 | 58134.668 | 58226.402 | 137858.328 | 137522.5 | 137021.453 | 4422203 | 4399080.5 | 4385594 | 1387558.75 | 1383430.25 | 1373742.875
| yA-A 5 | 9656.237 | 9716.562 | 10370.434 | 8904.438 | 8791.323 | 8728.627 | 513083.906 | 515221.094 | 522171.75 | 151723.438 | 153093.125 | 153252.938 | 51172.836 | 51625.113 | 51747.164 | 124339.023 | 124641.727 | 127274.102 | 3852741.75 | 3861304 | 3886983 | 1196626.375 | 1218202 | 1232262.875
| yA-A 6 | 11470.36 | 11802.071 | 11935.199 | 11271.392 | 11607.928 | 11170.978 | 645862 | 664375.813 | 678813.813 | 189095.844 | 193314.625 | 190869.188 | 54369.215 | 55299.727 | 55115.09 | 146578.453 | 146616.5 | 149231.344 | 4564896.5 | 4651256 | 4718967 | 1437475.75 | 1443033.625 | 1472440.25
Table 15 LC-MRM Intra Assay QC Data. Please report the values of each of the three replicates separately for each compound and standard. QCs have to be measured in one occasion only. The independent 6 replicates injected 3 times each
| Lowest Level QC | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| LLQC 1 | 2215 | 2189 | 2249.1 | 3048.4 | 3245.2 | 3039.7 | 157440.4 | 176017.9 | 174349.2 | 48261.6 | 52144.6 | 51283.2 | 27349.7 | 27634.8 | 27986.1 | 81091.6 | 80249.7 | 81524.7 | 2954668 | 3149770.8 | 3228822.5 | 865003.6 | 923655.8 | 929043.9
| LLQC 2 | 3157.8 | 3624.3 | 3608.3 | 4057.2 | 4349.9 | 4303.6 | 281479.4 | 270833.8 | 267203 | 85068.2 | 82196.5 | 81364.5 | 42305.8 | 41605.3 | 41998 | 108120 | 107335.4 | 109488.2 | 4188713.5 | 4005877.5 | 3948630 | 1219580.6 | 1166085.5 | 1189942.8
| LLQC 3 | 2142.2 | 2032.6 | 1817.2 | 2418.6 | 2685.9 | 2380.1 | 159266.9 | 155744 | 158321.7 | 49953.8 | 49116.7 | 49334 | 30042.2 | 30026.8 | 29883.1 | 77431.9 | 78183.6 | 77338.9 | 2952574.5 | 2937283.8 | 2909791.5 | 861808.1 | 851500.9 | 859901
| LLQC 4 | 3555.2 | 3493.2 | 3522.6 | 3595 | 3640.2 | 3886.7 | 228410.2 | 226494.9 | 225514.5 | 69440.1 | 70130.9 | 69894.6 | 47330.3 | 47036.5 | 48023.5 | 112404 | 111031.4 | 113062.7 | 4130891.8 | 4139636.3 | 4079943.3 | 1232154.1 | 1215391.1 | 1210097.1
| LLQC 5 | 3407.3 | 3623.4 | 3724.2 | 3820.9 | 3921.8 | 3990.4 | 225764.8 | 227884.5 | 227222.1 | 70188.6 | 71556.3 | 70531 | 38606.4 | 39589 | 40184.1 | 94012.4 | 93648.3 | 94356.1 | 3387620.5 | 3433738 | 3510024.5 | 1012335 | 1017171.8 | 1023287.6
| LLQC 6 | 2631.1 | 2655 | 2796.9 | 3264.3 | 3291.9 | 3191.8 | 185509.6 | 189813 | 194429 | 58258.7 | 58775.2 | 58095.9 | 40885.8 | 41256.9 | 40957.8 | 104310.2 | 104879.9 | 106040.6 | 3832305.8 | 3867651.8 | 3919123.8 | 1143130.3 | 1153637.5 | 1137494.9
| Low QC | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| LQC 1 | 2282.7 | 2161.6 | 2158.7 | 2829.7 | 2768.6 | 2606.2 | 126973.7 | 124724.8 | 124605.3 | 36833.8 | 37397.5 | 36703.4 | 17262 | 17723.8 | 17818.5 | 47128.9 | 47195.7 | 46909.9 | 1751138.9 | 1732633.5 | 1746793.3 | 513529.2 | 514645.4 | 514568.3
| LQC 2 | 1285.9 | 1550.5 | 1380.5 | 2008.1 | 1826.4 | 1853.6 | 114605.8 | 111471 | 112375 | 35995 | 36337.3 | 36419.9 | 16423.4 | 16068.1 | 16320.7 | 42540.3 | 41503.2 | 41572.8 | 1560468.3 | 1557149 | 1555178.9 | 454617.9 | 452970.4 | 455091.3
| LQC 3 | 1293.7 | 1327.6 | 1105.4 | 1755.7 | 1819.8 | 1911.7 | 119564.3 | 119973.5 | 119414.4 | 36502.1 | 37414.7 | 37456.9 | 13658.9 | 13614.7 | 13634.6 | 36825.4 | 37030.7 | 36888.8 | 1386879 | 1407185.3 | 1398082.4 | 408236.3 | 404321.7 | 406196.9
| LQC 4 | 1738.2 | 1958.2 | 2099.3 | 2603.9 | 2267.5 | 2384.9 | 138628.1 | 143235.4 | 142828.9 | 42689.1 | 43039.6 | 42849.8 | 20363.1 | 20270 | 20075.5 | 51974.6 | 52737.5 | 51645.2 | 1984280.9 | 1988680.3 | 2011022 | 581385.3 | 579378.2 | 585329.9
| LQC 5 | 2049.1 | 2041.4 | 2228.9 | 3025 | 3034.1 | 3531.8 | 166018.1 | 180115.2 | 183743.8 | 48094.7 | 52511.1 | 53006 | 16663.1 | 16849.1 | 17125.1 | 49650.6 | 50119.4 | 50371.9 | 1950936.1 | 2109272.8 | 2133546.3 | 556536.7 | 588172.8 | 594159.9
| LQC 6 | 2026.5 | 2043.9 | 1834.5 | 3191.5 | 2835.4 | 2682.4 | 168603.7 | 164554.3 | 135256.4 | 50821.1 | 48586.5 | 40859.3 | 22063.8 | 21688.3 | 18508.2 | 64526 | 64369.2 | 53028.2 | 2675461.3 | 2645656 | 2134449.3 | 754791.4 | 743783.3 | 615701.5
| Middle QC | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| MQC 1 | 36513.3 | 36350.9 | 35932.2 | 84409.1 | 85377.1 | 83565 | 2340367.3 | 2208925.5 | 2185714.5 | 1433455.1 | 1373040.9 | 1354160.5 | 29825 | 29258.5 | 28385.2 | 76301 | 75771.8 | 75081.7 | 2759841.5 | 2574213 | 2580786.8 | 786586.2 | 759448.7 | 753951.9
| MQC 2 | 25289.7 | 26402.9 | 26592.6 | 67902.7 | 66479.1 | 67465.2 | 1611430.9 | 1634735.5 | 1659031.1 | 1113457.5 | 1142341 | 1123257.1 | 20538.3 | 20916 | 20963.5 | 54399.3 | 54076.3 | 53686.5 | 1918545.1 | 1942087.1 | 1964101.9 | 571559.3 | 571835.6 | 571250.9
| MQC 3 | 29573.1 | 28924.1 | 28608.2 | 71474.2 | 72290.3 | 72605.2 | 1785846.9 | 1753614.3 | 1732497.8 | 1193764.9 | 1201713.3 | 1183610.1 | 24186.5 | 24535.8 | 23989.9 | 59186.9 | 59610.6 | 59739.4 | 2102429.8 | 2082044.1 | 2058483 | 621209 | 607949 | 621999.4
| MQC 4 | 30872.6 | 31341.5 | 29179.7 | 77308.6 | 77207.8 | 73126.5 | 1835577.9 | 1838939.6 | 1756278.1 | 1263447.8 | 1287794.4 | 1220977.4 | 22938.3 | 23180 | 21688.4 | 58208.4 | 57641.4 | 55564.1 | 2029118.1 | 2046651.4 | 1982115.9 | 590606.5 | 604362.1 | 574434
| MQC 5 | 30733.8 | 29077.7 | failed inj | 78400.2 | 71896.5 | failed inj | 1814824.8 | 1687426.5 | failed inj | 1355283.4 | 1260689 | failed inj | 23019.8 | 20315.4 | failed inj | 56462 | 51593.6 | failed inj | 1986076.4 | 1850769.9 | failed inj | 596193.9 | 542905.4 | failed inj
| MQC 6 | 16026.7 | 16648.8 | 14896.3 | 42757.5 | 42801.1 | 36702.5 | 945217.7 | 1064119.6 | 930990.4 | 705361.4 | 777467.9 | 634866.9 | 15599 | 15481.3 | 12937.4 | 38573.5 | 39550 | 33471.5 | 1253412.8 | 1410807.1 | 1218711.8 | 384181.8 | 413300.7 | 357559
| High QC | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| HQC 1 | 9798.9 | 10504.5 | 9815.2 | 25841.4 | 26601 | 24717.5 | 713582.9 | 779505.6 | 751134.1 | 439037.3 | 491436.5 | 467768.1 | 13856.2 | 14062.2 | 13301.9 | 39348 | 39201.1 | 37530.9 | 1415767.6 | 1561599.9 | 1504148.1 | 419849.2 | 451913.3 | 429206.8
| HQC 2 | 12419.9 | 13430.5 | 12003.9 | 30578.1 | 30205.5 | 26367.1 | 830325.4 | 799995.4 | 701128.4 | 546305.6 | 537443.2 | 466493.8 | 18103.9 | 18763.1 | 17034.4 | 47615.3 | 47445.1 | 40847 | 1814260.8 | 1739441.4 | 1516333.5 | 541037.9 | 516917.2 | 457435.8
| HQC 3 | 17942.5 | 18560.3 | 18554.4 | 41125.5 | 41084.7 | 40571.7 | 1028188.6 | 1042308.1 | 1024119.3 | 695803.9 | 695467.6 | 702119.5 | 23768.8 | 23555.7 | 24024.9 | 55494 | 56145.4 | 55749.2 | 1999256.9 | 1974953.8 | 1986003.3 | 598283.4 | 588253.4 | 594084.3
| HQC 4 | 21474.6 | 21853.4 | 21812.4 | 48272.8 | 48540 | 48251.2 | 1200430.9 | 1231054.6 | 1215460.1 | 811255 | 839889.6 | 838309.8 | 25205.9 | 25243 | 25118.1 | 62940.2 | 62500.3 | 63257.1 | 2230454.8 | 2287331 | 2302493.8 | 665109.3 | 672355.3 | 681843.7
| HQC 5 | 24103.1 | 24103.1 | failed inj | 53199 | 52351.9 | failed inj | 1379549.3 | 1375388.4 | failed inj | 875940.5 | 879792.7 | failed inj | 26876.1 | 27104.5 | failed inj | 66379.8 | 66553.7 | failed inj | 2452869.3 | 2476920.5 | failed inj | 726077.6 | 718171 | failed inj
| HQC 6 | 24408.3 | 24408.3 | failed inj | 51207.2 | 52791.3 | failed inj | 1263699.4 | 1422036 | failed inj | 822254.8 | 879417.4 | failed inj | 33216.9 | 33870.6 | failed inj | 81521.4 | 82090.8 | failed inj | 2607157.3 | 2896063.3 | failed inj | 794317.9 | 856224.3 | failed inj
| Highest Level QC | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| HLQC 1 | 9534.7 | 9393.3 | 7064.7 | 9799.4 | 10266.8 | 6164.4 | 503618.2 | 570447.7 | 350048.8 | 158729.9 | 172142.7 | 111234.1 | 50121.7 | 50136.3 | 35836.5 | 128790.5 | 131771.3 | 83954.3 | 4003685.8 | 4498481.5 | 2783713.3 | 1249956 | 1326992.1 | 834866
| HLQC 2 | 11924.4 | 11884.8 | 9990.6 | 15334.6 | 15454.3 | 11712.9 | 871911.5 | 988787.8 | 714292.9 | 263201.2 | 287914.6 | 210687.2 | 38737.4 | 39895.3 | 30891.3 | 118442.5 | 122245 | 87648.7 | 4004380.8 | 4576019.5 | 3239274.8 | 1180325.6 | 1301531.1 | 962100
| HLQC 3 | 9810.7 | 10834.6 | failed inj | 11584.8 | 11375.9 | failed inj | 715638.8 | 716498 | failed inj | 208854.4 | 207620.9 | failed inj | 42335.9 | 42203.2 | failed inj | 111030.4 | 114339.2 | failed inj | 4018199.3 | 3988223.3 | failed inj | 1169204.8 | 1174750.6 | failed inj
| HLQC 4 | 11699.4 | 12136.3 | failed inj | 12195.6 | 12250.6 | failed inj | 603888 | 683221.9 | failed inj | 193897.2 | 210226.5 | failed inj | 73304.7 | 73030.3 | failed inj | 194745.2 | 198163.6 | failed inj | 6045185.5 | 6761572.5 | failed inj | 1896909.1 | 2044145.1 | failed inj
| HLQC 5 | 6499.5 | 6861.7 | 5455.5 | 7064.7 | 7257.3 | 5253.6 | 370954.6 | 428804.9 | 302555.7 | 118790.2 | 132795.6 | 94779.6 | 47282.8 | 48722.2 | 36337.9 | 117587.9 | 122565 | 85556 | 3694010.5 | 4210245.5 | 2954733.5 | 1154848.1 | 1277959.9 | 916354.7
| HLQC 6 | 10718.4 | 11134.5 | failed inj | 15045.3 | 15573 | failed inj | 904989.3 | 1042074.3 | failed inj | 263031.6 | 298305.7 | failed inj | 59028.1 | 59158.5 | failed inj | 191564.2 | 197235.8 | failed inj | 6485200.5 | 7531476 | failed inj | 1938201.6 | 2178066.3 | failed inj
| Blank QC | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| Matrix Blank | 6568.5 | 10548.9 | n/a | 5797.7 | 10758.3 | n/a | 294585.8 | 560258.2 | n/a | 97601.5 | 175220.8 | n/a | 114.4 | ND | n/a | 124.1 | ND | n/a | 56.7 | 372.3 | n/a | 172.2 | 194.6 | n/a
| Total Blank | ND | ND | ND | ND | ND | ND | 10136.3 | 14691.5 | 16548.4 | 6954.9 | 6738.8 | 9890 | ND | ND | ND | ND | ND | ND | 188.5 | 417.6 | 265.5 | 113.3 | ND | 197.9
|  | ND | Not detected

## Raw Peak Areas
Sample name | Cer d18:1/16:0 [M>264] | Cer d18:1/16:0-d7 [M>264] | Cer d18:1/18:0 [M>264] | Cer d18:1/18:0-d7 [M>264] | Cer d18:1/24:0 [M>264] | Cer d18:1/24:0-d7 [M>264] | Cer d18:1/24:1 [M>264] | Cer d18:1/24:1-d7 [M>264]
Total blank 1 |  |  |  |  | 10136.3 | 188.5 | 6954.9 | 113.3
Matrix blank | 6568.5 | 144.4 | 5797.7 | 124.1 | 294585.8 | 56.7 | 97601.5 | 172.2
STD-6_1 Line 1 | 2639.9 | 36720.6 | 7274.6 | 78320.8 | 133056.5 | 2321766.3 | 115961.4 | 776035.8
STD-6_2 Line 1 | 1538.7 | 27537.5 | 5125.3 | 60841 | 98550.8 | 1749863.8 | 90090.4 | 595177.2
STD-6_3 Line 1 | 819.6 | 15694.3 | 2907.3 | 36304.4 | 60979 | 1065339.8 | 55806.7 | 358447.3
STD-5_1 Line 1 | 940.1 | 14261.6 | 2485.4 | 30040.4 | 47810.9 | 887241.8 | 42945.4 | 288003.7
STD-5_2 Line 1 | 468 | 7564.1 | 2164.2 | 17709.1 | 40317.4 | 540758.8 | 37083.1 | 175812.1
STD-5_3 Line 1 | 296.5 | 8597.5 | 1493.6 | 19285.6 | 36760.2 | 583021.1 | 35143.4 | 189350.9
STD-4_1 Line 1 | 1973.7 | 16113 | 5899.9 | 33369 | 119461.5 | 1026804.9 | 105216.1 | 326629.1
STD-4_2 Line 1 | 1449.1 | 9719.2 | 4888.9 | 23841.4 | 90872.9 | 719379.6 | 80641.8 | 230239.8
STD-4_3 Line 1 | 633.6 | 9837.3 | 2935 | 23040.9 | 67757.6 | 728291.3 | 58072 | 229324.6
STD-3_1 Line 1 | 13681 | 17600.8 | 34671.5 | 38770.7 | 629642 | 1185097.5 | 558913.9 | 388887.7
STD-3_2 Line 1 | 8577.7 | 13390.9 | 24212.9 | 29789.7 | 446231.8 | 923869.8 | 394434.7 | 283552
STD-3_3 Line 1 | 6304.5 | 10123.6 | 18870.3 | 22531.3 | 358521.6 | 709416.9 | 310575.5 | 211377.3
STD-2_1 Line 1 | 162920.4 | 19071.6 | 435999.6 | 43532.5 | 8146007 | 1374526.8 | 7098589 | 443021.4
STD-2_2 Line 1 | 99568.1 | 12869.6 | 265772.4 | 29771.8 | 4970327 | 953609.6 | 4277954.5 | 303883.3
STD-2_3 Line 1 | 109124.9 | 15976.9 | 290046.2 | 38288.8 | 5663704 | 1244105.9 | 4780396.5 | 396407.1
STD-1_1 Line 1 | 161670.3 | 13193.4 | 431775.4 | 29792.5 | 8043190 | 945122.8 | 6912193.5 | 288977
STD-1_2 Line 1 | 188843.3 | 12360 | 492482.3 | 27829.6 | 8987714 | 855688.7 | 7731415 | 275952.6
STD-1_3 Line 1 | 138922.4 | 9367.4 | 366987.5 | 22106 | 6841481.5 | 686402.2 | 5683858 | 217880.8
Pooled QC | 18348.6 | 15228.7 | 42049.2 | 33521.5 | 882582.4 | 1019178.8 | 651321.4 | 334675.5
Pooled QC | 18789.2 | 16130.9 | 41218.5 | 33594.9 | 875009.2 | 1014786.3 | 645487.2 | 327925.3
SRM 1950_1 | 15436.9 | 65022.2 | 14500.7 | 151160.3 | 681805.5 | 4341372.5 | 224244.5 | 1392100.5
SRM 1950_1 | 16314.8 | 64518.5 | 14473.4 | 151860.8 | 677462.8 | 4312498 | 223780.5 | 1408901.8
SRM 1950_1 | 14986.5 | 64145.6 | 15033.9 | 153184.6 | 684014.9 | 4414692 | 227862.2 | 1430026.6
SRM 1950_2 | 13104.8 | 45448 | 12462.9 | 106584.8 | 596229.7 | 3138495.3 | 198338.4 | 1025875.4
SRM 1950_2 | 12761.1 | 45167.2 | 11995.3 | 107404.8 | 600847.4 | 3231560.3 | 196484.2 | 1052141.1
SRM 1950_2 | 13306.5 | 45272 | 11964 | 106785.2 | 600883.4 | 3210960.8 | 200121.9 | 1052931.1
SRM 1950_3 | 11630.1 | 47679.2 | 11231.3 | 114303.5 | 536108 | 3396372.5 | 181562.5 | 1094107.6
SRM 1950_3 | 12178.4 | 48279.7 | 11200.3 | 111957.1 | 552527.4 | 3389653.3 | 180861.5 | 1100481.3
SRM 1950_3 | 12149.9 | 47575.8 | 11399.5 | 113698.5 | 552447.9 | 3396195.8 | 183923.6 | 1117860.8
SRM 1950_4 | 12687.6 | 66291 | 10719.2 | 147548.5 | 519237.6 | 4454893.5 | 171942 | 1455284.6
SRM 1950_4 | 12242.2 | 66060 | 11039.7 | 145785.2 | 521329.6 | 4386499.5 | 172139.1 | 1438925.1
SRM 1950_4 | 12127.9 | 64484 | 10896.1 | 145517.4 | 526884 | 4443964.5 | 171500.2 | 1424130.9
SRM 1950_5 | 12537.7 | 61886.5 | 11580.8 | 141601.1 | 539459.5 | 4169094.3 | 177839.1 | 1379458.8
SRM 1950_5 | 13009.7 | 61877 | 11507.4 | 144289 | 544855.9 | 4344303 | 180934.8 | 1426364.4
SRM 1950_5 | 11847.9 | 63271.8 | 11360.4 | 144576.7 | 555169.1 | 4405437.5 | 182172.5 | 1395926.4
SRM 1950_6 | 14234.1 | 54015.6 | 13601.7 | 128640.8 | 680419.2 | 3956682.8 | 216232.9 | 1241749.1
SRM 1950_6 | 14364.5 | 52911.8 | 13476.7 | 129781.9 | 692953.6 | 4068984.3 | 222140.6 | 1276584.3
SRM 1950_6 | 14327.7 | 54039.9 | 13838 | 131661.9 | 701439.1 | 4052581.8 | 221483.8 | 1296079.6
Pooled QC | 18823.5 | 15253.1 | 43116.5 | 35696.5 | 944033.9 | 1096454.6 | 682264.8 | 348990.4
Pooled QC | 18687.1 | 16166 | 42615.1 | 34358.1 | 916387.7 | 1059148.1 | 675221.3 | 345017.4
T1D_1 | 14161.3 | 56261.8 | 17747.8 | 139005.9 | 742592.4 | 4055878.8 | 247402.7 | 1317272.8
T1D_1 | 14397.6 | 57254.9 | 17411.2 | 136908.6 | 762276.6 | 3987511 | 242761.1 | 1310546.6
T1D_1 | 14406.1 | 57748.4 | 18351.2 | 138661.3 | 758625.6 | 4017309.3 | 241558.6 | 1315208.6
T1D_2 | 9518.2 | 45833.9 | 11294.5 | 109346.5 | 501083.8 | 3332506.3 | 163424.7 | 1088271.5
T1D_2 | 9495 | 46876.4 | 11380.9 | 110182.3 | 509048.2 | 3407169.8 | 163596.4 | 1091701.8
T1D_2 | 9483.6 | 47228.2 | 11498 | 110911 | 510878.8 | 3456803.8 | 164701.4 | 1110566.6
T1D_3 | 13140.5 | 50490.8 | 15325.1 | 120801.7 | 653085.8 | 3654831.8 | 217421.4 | 1180001.9
T1D_3 | 12792 | 50856.6 | 15414.3 | 120864.2 | 654023 | 3642395 | 213144.7 | 1160410.5
T1D_3 | 12755 | 50877.5 | 15854 | 122842.6 | 659736.2 | 3674864.5 | 215131 | 1177917.5
T1D_4 | 11986.6 | 45057.4 | 13646.5 | 100924.5 | 583297.7 | 2962037.8 | 189285.8 | 969644.8
T1D_4 | 11862.3 | 44912.1 | 12906.1 | 100722.7 | 572196.8 | 2940367.8 | 189656.8 | 959750.5
T1D_4 | 11765.3 | 44584 | 13779.4 | 100224.8 | 569588.4 | 2998590.3 | 186851.7 | 968603.7
T1D_5 | 10443 | 48193.3 | 12637 | 112195.3 | 529280.8 | 3328376.8 | 174805.4 | 1094013
T1D_5 | 10596.7 | 47130.8 | 12340.9 | 111242.8 | 545595.9 | 3440922.8 | 176898.3 | 1097847.8
T1D_5 | 10833.4 | 46869 | 12948.4 | 112756 | 550630.8 | 3438498.8 | 178246.7 | 1100645.5
T1D_6 | 7924 | 59848.5 | 10360.1 | 144518.8 | 438901.9 | 4463162.5 | 142017.3 | 1422120.4
T1D_6 | 8416.8 | 61268.8 | 10212.5 | 145031.8 | 431897.2 | 4581243.5 | 144094.4 | 1437266.9
T1D_6 | 8620.6 | 61921.2 | 10213.7 | 146930.5 | 438276.8 | 4509406 | 144796.9 | 1441855.8
Total blank 2 |  |  |  |  | 14691.5 | 417.6 | 6738.8
Pooled QC | 19406.8 | 16477.9 | 44349.3 | 35616.2 | 978912.4 | 1125219.6 | 692260.3 | 359304.7
Pooled QC | 19394.7 | 16242.6 | 43562.4 | 35362 | 963323.1 | 1111976.4 | 693622.5 | 353690.6
Young AA_1 | 18170.3 | 74744 | 16782.3 | 183911.5 | 914714.3 | 5195973.5 | 273145.8 | 1694428.4
Young AA_1 | 18974.5 | 75085.9 | 17065.2 | 184805.2 | 912233.3 | 5183351.5 | 270449.1 | 1688717.8
Young AA_1 | 19668.6 | 75517.2 | 17147.9 | 185306.5 | 897802.3 | 5206968.5 | 270825.7 | 1704857.9
Young AA_2 | 6231.3 | 63004.5 | 5671.1 | 151965.1 | 333246.2 | 4606121 | 100858.8 | 1503073.6
Young AA_2 | 6230.4 | 62231.6 | 6027.1 | 151569.8 | 340202.9 | 4696746 | 102288.6 | 1530413.9
Young AA_2 | 6095.9 | 63180.5 | 5984.8 | 150679.9 | 345186.5 | 4729148 | 103882.7 | 1530609.4
Young AA_3 | 9915.6 | 58546.1 | 9563.2 | 151101.1 | 521022.3 | 4832932 | 152300.8 | 1497946.6
Young AA_3 | 10408.5 | 60328 | 9654.5 | 152567.7 | 532983 | 4906300 | 152556.1 | 1500713.5
Young AA_3 | 10346.6 | 60132.4 | 9563.6 | 152877.2 | 530973.1 | 4951557.5 | 155066.9 | 1533961
Young AA_4 | 5739.4 | 57508 | 5382.2 | 137858.3 | 277700.7 | 4422203 | 81820.5 | 1387558.8
Young AA_4 | 5974.8 | 58134.7 | 4979.1 | 137522.5 | 272699.2 | 4399080.5 | 80373.6 | 1383430.3
Young AA_4 | 5467.2 | 58226.4 | 5006.2 | 137021.5 | 274479.3 | 4385594 | 81159.4 | 1373742.9
Young AA_5 | 9656.2 | 51172.8 | 8904.4 | 124339 | 513083.9 | 3852741.8 | 151723.4 | 1196626.4
Young AA_5 | 9716.6 | 51625.1 | 8791.3 | 124641.7 | 515221.1 | 3861304 | 153093.1 | 1218202
Young AA_5 | 10370.4 | 51747.2 | 8728.6 | 127274.1 | 522171.8 | 3886983 | 153252.9 | 1232262.9
Young AA_6 | 11470.4 | 54369.2 | 11271.4 | 146578.5 | 645862 | 4564896.5 | 189095.8 | 1437475.8
Young AA_6 | 11802.1 | 55299.7 | 11607.9 | 146616.5 | 664375.8 | 4651256 | 193314.6 | 1443033.6
Young AA_6 | 11935.2 | 55115.1 | 11171 | 149231.3 | 678813.8 | 4718967 | 190869.2 | 1472440.3
Pooled QC | 20497.5 | 16508.6 | 45523.1 | 37271.6 | 1061889.8 | 1214422.6 | 766505.7 | 381317.2
Pooled QC | 19703.4 | 16693.3 | 45335 | 36940.9 | 1011485.1 | 1175409.8 | 721520.6 | 369751.7
hTAG_1 | 44652.1 | 83913.7 | 47025.5 | 218882.2 | 2610320 | 6246800 | 786314.7 | 1986632.1
hTAG_1 | 45517.2 | 85090.4 | 47240.5 | 220236.5 | 2595654.5 | 6205797 | 777633.3 | 1991988.9
hTAG_1 | 45129.3 | 84928.8 | 46947.8 | 220604.8 | 2605426.5 | 6241036 | 791885.9 | 1998425.8
hTAG_2 | 29460.7 | 70247.3 | 31535.9 | 202326 | 1836938.5 | 5820052.5 | 553689.5 | 1852259.5
hTAG_2 | 30770.3 | 69570.1 | 32624.2 | 202238.6 | 1899222.5 | 5885250.5 | 566766.9 | 1899622.1
hTAG_2 | 29673.1 | 70397.8 | 32892.6 | 204421.2 | 1912866.1 | 5957444 | 564470.6 | 1872637
hTAG_3 | 29454.9 | 75257.3 | 31397.5 | 211177 | 1803000.6 | 6332020.5 | 523551.6 | 1972418.4
hTAG_3 | 29349.4 | 74506.6 | 30707.3 | 209881 | 1807376.6 | 6390343.5 | 528989.5 | 1986380.8
hTAG_3 | 28824.6 | 74876.6 | 30842.3 | 212289.8 | 1803893.6 | 6292947 | 534711.9 | 2028472.3
hTAG_4 | 36212 | 80454.6 | 37206.8 | 217845.3 | 2124340.8 | 6484776.5 | 646113.6 | 2036319.1
hTAG_4 | 36818.9 | 82028.5 | 37126.3 | 220091.8 | 2128993.3 | 6444215.5 | 635747.9 | 2015724.8
hTAG_4 | 36837.2 | 83778.9 | 36960.9 | 220511.8 | 2132806.8 | 6358559 | 645196.9 | 2023588
hTAG_5 | 30079.9 | 81222.5 | 30211.9 | 213995.4 | 1678585 | 6467977 | 500639.5 | 2056280.8
hTAG_5 | 29434.1 | 80021.4 | 29707.4 | 214363.9 | 1702684.4 | 6426290 | 507014.6 | 2038441.4
hTAG_5 | 30090.1 | 82054.4 | 29400.5 | 216528.5 | 1714143 | 6436785 | 504837.8 | 2057209.3
hTAG_6 | 23444.5 | 70214.6 | 24017.6 | 199561.7 | 1356627.8 | 6077615.5 | 409790.9 | 1904866.9
hTAG_6 | 23542.3 | 69976.6 | 24286.7 | 197658.5 | 1380889.5 | 6147816.5 | 412467.3 | 1940109.3
hTAG_6 | 23469.4 | 71164.4 | 24644.2 | 201788.4 | 1386984.3 | 6174492 | 416608.2 | 1922105.4
Pooled QC | 20437.3 | 16991.4 | 46380.5 | 38466.5 | 1082368.5 | 1278831.9 | 763143.1 | 387774
Pooled QC | 20800.7 | 17162.5 | 47463.6 | 38422.3 | 1047152.9 | 1251505.1 | 758873.5 | 391446.9
STD-6_1 Line 2 | 2019.1 | 30963.5 | 7740.6 | 84421.3 | 164199.8 | 2913704.3 | 135193.4 | 851272.4
STD-6_2 Line 2 | 1946.5 | 30014.5 | 6290.7 | 71379.8 | 143158.2 | 2568921.8 | 120758.3 | 768740.8
STD-6_3 Line 2 | 2514.1 | 34929.8 | 7113.3 | 84000.8 | 148346.4 | 2921018.8 | 115555.3 | 866205.3
STD-5_1 Line 2 | 584.3 | 13211.3 | 1909.5 | 30803.8 | 52792.4 | 1071927.3 | 45772.7 | 316065.6
STD-5_2 Line 2 | 784.3 | 13242.3 | 2837.3 | 30896.4 | 67830.7 | 1050741.1 | 55226.5 | 319296.4
STD-5_3 Line 2 | 937.3 | 11357.2 | 2946 | 28448.4 | 68206 | 994547.3 | 52085.9 | 295500.4
STD-4_1 Line 2 | 886.6 | 11251.2 | 3637.8 | 26524.5 | 82246.5 | 922217 | 70413.2 | 271849.8
STD-4_2 Line 2 | 1214.6 | 10266.2 | 4593.1 | 24677.2 | 96773.3 | 866045.1 | 79030.7 | 256217.3
STD-4_3 Line 2 | 1273.6 | 10747 | 5198.7 | 27329.7 | 111277.2 | 955618.6 | 90581.3 | 284815.4
STD-3_1 Line 2 | 14533.2 | 14343.4 | 41671.9 | 34841.4 | 894912 | 1241256 | 721902.9 | 369569.8
STD-3_2 Line 2 | 10892.8 | 12222.7 | 31974.8 | 30054 | 674329.1 | 1062834.5 | 544451.5 | 316439.4
STD-3_3 Line 2 | 27592 | 20205.2 | 70192.3 | 50203.7 | 1527557.6 | 1665800.8 | 1157965.3 | 516793.7
STD-2_1 Line 2 | 347266.3 | 36674 | 980989.8 | 91389.2 | 21319064 | 3265680.3 | 16963404 | 999383.1
STD-2_2 Line 2 | 315365.7 | 31191.1 | 901498.6 | 77063.3 | 18968978 | 2830980.3 | 15407429 | 844960.6
STD-2_3 Line 2 | 348983.5 | 34747.7 | 1011734.1 | 89799.4 | 21518826 | 3223640.8 | 17533704 | 967858.3
STD-1_1 Line 2 | 277289.6 | 16355.6 | 700361.9 | 35305.1 | 14575619 | 1267685.4 | 11897055 | 381619.9
STD-1_2 Line 2 | 255814.4 | 15111.7 | 669241 | 35629.6 | 13329707 | 1234273.9 | 11154217 | 372679.4
STD-1_3 Line 2 | 203670.1 | 13766.3 | 568867.3 | 35497.4 | 11794922 | 1232357.6 | 9360104 | 360469.7
Lowest QC_1 | 2215 | 27349.7 | 3048.4 | 81091.6 | 157440.4 | 2954668 | 48261.6 | 865003.6
Lowest QC_1 | 2189 | 27634.8 | 3245.2 | 80249.7 | 176017.9 | 3149770.8 | 52144.6 | 923655.8
Lowest QC_1 | 2249.1 | 27986.1 | 3039.7 | 81524.7 | 174349.2 | 3228822.5 | 51283.2 | 929043.9
Lowest QC_2 | 3157.8 | 42305.8 | 4057.2 | 108120 | 281479.4 | 4188713.5 | 85068.2 | 1219580.6
Lowest QC_2 | 3624.3 | 41605.3 | 4349.9 | 107335.4 | 270833.8 | 4005877.5 | 82196.5 | 1166085.5
Lowest QC_2 | 3608.3 | 41998 | 4303.6 | 109488.2 | 267203 | 3948630 | 81364.5 | 1189942.8
Lowest QC_3 | 2142.2 | 30042.2 | 2418.6 | 77431.9 | 159266.9 | 2952574.5 | 49953.8 | 861808.1
Lowest QC_3 | 2032.6 | 30026.8 | 2685.9 | 78183.6 | 155744 | 2937283.8 | 49116.7 | 851500.9
Lowest QC_3 | 1817.2 | 29883.1 | 2380.1 | 77338.9 | 158321.7 | 2909791.5 | 49334 | 859901
Lowest QC_4 | 3555.2 | 47330.3 | 3595 | 112404 | 228410.2 | 4130891.8 | 69440.1 | 1232154.1
Lowest QC_4 | 3493.2 | 47036.5 | 3640.2 | 111031.4 | 226494.9 | 4139636.3 | 70130.9 | 1215391.1
Lowest QC_4 | 3522.6 | 48023.5 | 3886.7 | 113062.7 | 225514.5 | 4079943.3 | 69894.6 | 1210097.1
Lowest QC_5 | 3407.3 | 38606.4 | 3820.9 | 94012.4 | 225764.8 | 3387620.5 | 70188.6 | 1012335
Lowest QC_5 | 3623.4 | 39589 | 3921.8 | 93648.3 | 227884.5 | 3433738 | 71556.3 | 1017171.8
Lowest QC_5 | 3724.2 | 40184.1 | 3990.4 | 94356.1 | 227222.1 | 3510024.5 | 70531 | 1023287.6
Lowest QC_6 | 2631.1 | 40885.8 | 3264.3 | 104310.2 | 185509.6 | 3832305.8 | 58258.7 | 1143130.3
Lowest QC_6 | 2655 | 41256.9 | 3291.9 | 104879.9 | 189813 | 3867651.8 | 58775.2 | 1153637.5
Lowest QC_6 | 2796.9 | 40957.8 | 3191.8 | 106040.6 | 194429 | 3919123.8 | 58095.9 | 1137494.9
Low QC_1 | 2683.7 | 17262 | 2829.7 | 47128.9 | 126973.7 | 1751138.9 | 36833.8 | 513529.2
Low QC_1 | 2161.6 | 17723.8 | 2768.6 | 47195.7 | 124724.8 | 1732633.5 | 37397.5 | 514645.4
Low QC_1 | 2158.7 | 17818.5 | 2606.2 | 46909.9 | 124605.3 | 1746793.3 | 36703.4 | 514568.3
Low QC_2 | 1285.9 | 16423.4 | 2008.1 | 42540.3 | 114605.8 | 1560468.3 | 35995 | 454617.9
Low QC_2 | 1550.5 | 16068.1 | 1826.4 | 41503.2 | 111471 | 1557149 | 36337.3 | 452970.4
Low QC_2 | 1380.5 | 16320.7 | 1853.6 | 41572.8 | 112375 | 1555178.9 | 36419.9 | 455091.3
Low QC_3 | 1293.7 | 13658.9 | 1755.7 | 36825.4 | 119564.3 | 1386879 | 36502.1 | 408236.3
Low QC_3 | 1327.6 | 13614.7 | 1819.8 | 37030.7 | 119973.5 | 1407185.3 | 37414.7 | 404321.7
Low QC_3 | 1105.4 | 13634.6 | 1911.7 | 36888.8 | 119414.4 | 1398082.4 | 37456.9 | 406196.9
Low QC_4 | 1738.2 | 20363.1 | 2603.9 | 51974.6 | 138628.1 | 1984280.9 | 42689.1 | 581385.3
Low QC_4 | 1958.2 | 20270 | 2267.5 | 52737.5 | 143235.4 | 1988680.3 | 43039.6 | 579378.2
Low QC_4 | 2099.3 | 20075.5 | 2384.9 | 51645.2 | 142828.9 | 2011022 | 42849.8 | 585329.9
Low QC_5 | 2049.1 | 16663.1 | 3025 | 49650.6 | 166018.1 | 1950936.1 | 48094.7 | 556536.7
Low QC_5 | 2041.4 | 16849.1 | 3034.1 | 50119.4 | 180115.2 | 2109272.8 | 52511.1 | 588172.8
Low QC_5 | 2228.9 | 17125.1 | 3531.8 | 50371.9 | 183743.8 | 2133546.3 | 53006 | 594159.9
Low QC_6 | 2026.5 | 22063.8 | 3191.5 | 64526 | 168603.7 | 2675461.3 | 50821.1 | 754791.4
Low QC_6 | 2043.9 | 21688.3 | 2835.4 | 64369.2 | 164554.3 | 2645656 | 48586.5 | 743783.3
Low QC_6 | 1834.5 | 18508.2 | 2682.4 | 53028.2 | 135256.4 | 2134449.3 | 40859.3 | 615701.5
Middle QC_1 | 36513.3 | 29825 | 84409.1 | 76301 | 2340367.3 | 2759841.5 | 1433455.1 | 786586.2
Middle QC_1 | 36350.9 | 29258.5 | 85377.1 | 75771.8 | 2208925.5 | 2574213 | 1373040.9 | 759448.7
Middle QC_1 | 35932.2 | 28385.2 | 83565 | 75081.7 | 2185714.5 | 2580786.8 | 1354160.5 | 753951.9
Middle QC_2 | 25289.7 | 20538.3 | 67902.7 | 54399.3 | 1611430.9 | 1918545.1 | 1113457.5 | 571559.3
Middle QC_2 | 26402.9 | 20916 | 66479.1 | 54076.3 | 1634735.5 | 1942087.1 | 1142341 | 571835.6
Middle QC_2 | 26592.6 | 20963.5 | 67465.2 | 53686.5 | 1659031.1 | 1964101.9 | 1123257.1 | 571250.9
Middle QC_3 | 29573.1 | 24186.5 | 71474.2 | 59186.9 | 1785846.9 | 2102429.8 | 1193764.9 | 621209
Middle QC_3 | 28924.1 | 24535.8 | 72290.3 | 59610.6 | 1753614.3 | 2082044.1 | 1201713.3 | 607949
Middle QC_3 | 28608.2 | 23989.9 | 72605.2 | 59739.4 | 1732497.8 | 2058483 | 1183610.1 | 621999.4
Middle QC_4 | 30872.6 | 22938.3 | 77308.6 | 58208.4 | 1835577.9 | 2029118.1 | 1263447.8 | 590606.5
Middle QC_4 | 31341.5 | 23180 | 77207.8 | 57641.4 | 1838939.6 | 2046651.4 | 1287794.4 | 604362.1
Middle QC_4 | 29179.7 | 21688.4 | 73126.5 | 55564.1 | 1756278.1 | 1982115.9 | 1220977.4 | 574434
Middle QC_5 | 30733.8 | 23019.8 | 78400.2 | 56462 | 1814824.8 | 1986076.4 | 1355283.4 | 596193.9
Middle QC_5 | 29077.7 | 20315.4 | 71896.5 | 51593.6 | 1687426.5 | 1850769.9 | 1260689 | 542905.4
Middle QC_5
Middle QC_6 | 16026.7 | 15599 | 42757.5 | 38573.5 | 945217.7 | 1253412.8 | 705361.4 | 384181.8
Middle QC_6 | 16648.8 | 15481.3 | 42801.1 | 39550 | 1064119.6 | 1410807.1 | 777467.9 | 413300.7
Middle QC_6 | 14896.3 | 12937.4 | 36702.5 | 33471.5 | 930990.4 | 1218711.8 | 634866.9 | 357559
High QC_1 | 9798.9 | 13856.2 | 25841.4 | 39348 | 713582.9 | 1415767.6 | 439037.3 | 419849.2
High QC_1 | 10504.5 | 14062.2 | 26601 | 39201.1 | 779505.6 | 1561599.9 | 491436.5 | 451913.3
High QC_1 | 9815.2 | 13301.9 | 24717.5 | 37530.9 | 751134.1 | 1504148.1 | 467768.1 | 429206.8
High QC_2 | 12419.9 | 18103.9 | 30578.1 | 47615.3 | 830325.4 | 1814260.8 | 546305.6 | 541037.9
High QC_2 | 13430.5 | 18763.1 | 30205.5 | 47445.1 | 799995.4 | 1739441.4 | 537443.2 | 516917.2
High QC_2 | 12003.9 | 17034.4 | 26367.1 | 40847 | 701128.4 | 1516333.5 | 466493.8 | 457435.8
High QC_3 | 17942.5 | 23768.8 | 41125.5 | 55494 | 1028188.6 | 1999256.9 | 695803.9 | 598283.4
High QC_3 | 18560.3 | 23555.7 | 41084.7 | 56145.4 | 1042308.1 | 1974953.8 | 695467.6 | 588253.4
High QC_3 | 18554.4 | 24024.9 | 40571.7 | 55749.2 | 1024119.3 | 1986003.3 | 702119.5 | 594084.3
High QC_4 | 21474.6 | 25205.9 | 48272.8 | 62940.2 | 1200430.9 | 2230454.8 | 811255 | 665109.3
High QC_4 | 21853.4 | 25243 | 48540 | 62500.3 | 1231054.6 | 2287331 | 839889.6 | 672355.3
High QC_4 | 21812.4 | 25118.1 | 48251.2 | 63257.1 | 1215460.1 | 2302493.8 | 838309.8 | 681843.7
High QC_5 | 22045 | 26876.1 | 53199 | 66379.8 | 1379549.3 | 2452869.3 | 875940.5 | 726077.6
High QC_5 | 22513.1 | 27104.5 | 52351.9 | 66553.7 | 1375388.4 | 2476920.5 | 879792.7 | 718171
High QC_5
High QC_6 | 24103.1 | 33216.9 | 51207.2 | 81521.4 | 1263699.4 | 2607157.3 | 822254.8 | 794317.9
High QC_6 | 24408.3 | 33870.6 | 52791.3 | 82090.8 | 1422036 | 2896063.3 | 879417.4 | 856224.3
High QC_6
Highest QC_1 | 9534.7 | 50121.7 | 9799.4 | 128790.5 | 503618.2 | 4003685.8 | 158729.9 | 1249956
Highest QC_1 | 9393.3 | 50136.3 | 10266.8 | 131771.3 | 570447.7 | 4498481.5 | 172142.7 | 1326992.1
Highest QC_1 | 7064.7 | 35836.5 | 6164.4 | 83954.3 | 350048.8 | 2783713.3 | 111234.1 | 834866
Highest QC_2 | 11924.4 | 38737.4 | 15334.6 | 118442.5 | 871911.5 | 4004380.8 | 263201.2 | 1180325.6
Highest QC_2 | 11884.8 | 39895.3 | 15454.3 | 122245 | 988787.8 | 4576019.5 | 287914.6 | 1301531.1
Highest QC_2 | 9990.6 | 30891.3 | 11712.9 | 87648.7 | 714292.9 | 3239274.8 | 210687.2 | 962100
Highest QC_3 | 9810.7 | 42335.9 | 11584.8 | 111030.4 | 715638.8 | 4018199.3 | 208854.4 | 1169204.8
Highest QC_3 | 10834.6 | 42203.2 | 11375.9 | 114339.2 | 716498 | 3988223.3 | 207620.9 | 1174750.6
Highest QC_3
Highest QC_4 | 11699.4 | 73304.7 | 12195.6 | 194745.2 | 603888 | 6045185.5 | 193897.2 | 1896909.1
Highest QC_4 | 12136.3 | 73030.3 | 12250.6 | 198163.6 | 683221.9 | 6761572.5 | 210226.5 | 2044145.1
Highest QC_4
Highest QC_5 | 6499.5 | 47282.8 | 7064.7 | 117587.9 | 370954.6 | 3694010.5 | 118790.2 | 1154848.1
Highest QC_5 | 6861.7 | 48722.2 | 7257.3 | 122565 | 428804.9 | 4210245.5 | 132795.6 | 1277959.9
Highest QC_5 | 5455.5 | 36337.9 | 5253.6 | 85556 | 302555.7 | 2954733.5 | 94779.6 | 916354.7
Highest QC_6 | 10718.4 | 59028.1 | 15045.3 | 191564.2 | 904989.3 | 6485200.5 | 263031.6 | 1938201.6
Highest QC_6 | 11134.5 | 59158.5 | 15573 | 197235.8 | 1042074.3 | 7531476 | 298305.7 | 2178066.3
Highest QC_6
Matrix blank | 10548.9 |  | 10758.3 |  | 560258.2 | 372.3 | 175220.8 | 194.6
Total blank 3 |  |  |  |  | 16548.4 | 265.5 | 9890 | 197.9