## Participant Information
Participant Information
Participant Information |  | Group Name
|  | Lab code | 3
| Institution | Institution name
Street
City
Postal code
Country
| Contact person | Name
Address
Email
Phone

## Summary
Summary
Summary | This protocol describes the extraction and LC-MRM procedure for absolute quantification of Cer d18:1/16:0, Cer d18:1/18:0, Cer d18:1/24:0 and Cer d18:1/24:1 in human plasma according to Kauhanen et al, Anal Bioanal Chem (2016) 408:3475–3483.
A calibration curve is built by injecting several dilutions of the non-labelled standards mixed with a fixed concentration of labelled standards. Each calibration point is prepared and analysed in triplicates (n=3).
To establish the endogenous concentration of the four selected ceramides in the study samples (NIST SRM1950, NIST high TAG, NIST T1D, and NIST young A-A) each sample is extracted and analyzed in six replicates (n =6).
Prepare the pooled QC by mixing 100 mL each of NIST SRM1950, NIST high TAG, NIST T1D, and NIST young AA (total 400 µL).
We suggest proceeding first to measure the ceramides concentration in the four study samples and then continue with the intra-assay validation to avoid wasting material.

## Extraction - Overview
Extraction Overview
Extraction Overview | Table 1 - Extraction - Materials
Chemicals, solvents and materials
Ethyl acetate (EtAc)
Isopropanol (IPA)
5% Bovine Serum Albumin (BSA)
250 µL Glass vials
Table 2 - Extraction - Sample Types
| Sample type | Components
| Pooled Quality Control samples (QC) | 10 µL pooled plasma QC |  |  | 20 µL Labelled IS |  | 570 µL EtAc:IPA 2:8
| Matrix Blank (MB) | 10 µL Matrix (pooled plasma QC) only |  |  |  |  | 590 µL EtAc:IPA 2:8
| Study Samples (S) | 10 µL Human plasma (NIST SRM1950, high TAG, T1D, young AA) |  |  | 20 µL Labelled IS |  | 570 µL EtAc:IPA 2:8
| Total Blank (TB) | 10 µL 5% BSA |  |  |  |  | 590 µL EtAc:IPA 2:8
| Calibration line standards (STD) | 10 µL 5% BSA + 20 µL non-labelled STDx mixture (Table 5) |  |  | 20 µL Labelled IS |  | 550 µL EtAc:IPA 2:8
Table 3 - Extraction - Labelled Internal Standard Mixture Solution (provided)
| Labelled IS mixture in EtAc:IPA 2:8 (v/v) | Concentration |  |  |  |  | Unit | Vendor
| D7-Cer d18:1/16:0 | 0.125 |  |  |  |  | pmol/µL [µM] | Avanti Polar Lipids
| D7-Cer d18:1/18:0 | 0.050
| D7-Cer d18:1/24:0 | 1500
| D7-Cer d18:1/24:1 | 0.500
Table 4 - Extraction - Non-Labelled Standards - for STD1 Mixture Solution (provided)
| Non-labelled standards in EtAc:IPA 2:8 (v/v) | Concentration |  |  |  |  | Unit | Vendor
| Cer d18:1/16:0 | 2.0 |  |  |  |  | pmol/µL [µM] | Avanti Polar Lipids
| Cer d18:1/18:0 | 2.0
| Cer d18:1/24:0 | 20.0
| Cer d18:1/24:1 | 20.0
Table 5 - Serial dilutions of Non-Labelled Standards from STD1 Mixture Solution for calibration line
| Add 20 µL of | STD1 Conc. | STD2 Conc. | STD3 Conc. | STD4 Conc. | STD5 Conc. | STD6 Conc. | Unit
| Cer d18:1/16:0 | 2.0 | 1.0 | 0.10 | 0.02 | 0.01 | 0.008 | pmol/µL [µM]
| Cer d18:1/18:0 | 2.0 | 1.0 | 0.10 | 0.02 | 0.01 | 0.008
| Cer d18:1/24:0 | 20.0 | 10.0 | 1.0 | 0.2 | 0.1 | 0.08
| Cer d18:1/24:1 | 20.0 | 10.0 | 1.0 | 0.2 | 0.1 | 0.08

## Extraction - Procedure
Extraction Procedure
Extraction Procedure | Extraction Procedure
| Step | Description | Reference Table
| 1 | 1) Prepare different dilutions of STD1 by diluting the provided mixture with EtAc:IPA 2:8 | Extraction - Overview'!B34
| 2 | 2) Thaw plasma samples on ice (+4°C) and bring them to room temperature prior to extraction. Mix by pipetting up and down several times before use
| 3 | 3) Prepare samples in plates/tubes manually or with a robotic liquid handler, according to specifications in Column 2 of Table 2 | Extraction - Overview'!C14
| 4 | 4) Add 20 µL of labelled internal standard mixture to samples (Column 3 Table 2). | Extraction - Overview'!F14
| 5 | 5) For the calibration line, add 20 µL of different dilutions of the non-labelled mixture (STD1-6).
| 6 | 6) Add 590 µL of ethyl acetate:isopropanol (2:8, vol/vol) to Total Blank and Matrix Blank; 550 µL to calibration line (STD); 570 µL to study samples and QCs (Column 4 Table 2) | Extraction - Overview'!H14
| 7 | 7) Mix samples 10 min by automated pipetting or vortexing
| 8 | 8) Centrifuge samples for 10 min at 3000 x g at room temperature
| 9 | 9) Transfer 250 µL of the clear supernatant to MS vials, evaporate and add 100 µL EtAc:IPA 2:8 store samples at -20°C prior to LC-MRM analysis.

## LC-MRM - Overview
LC-MRM Overview
LC-MRM Overview | Table 6 Chemicals and solvents
2-Propanol (IPA), Chromasolv LC-MS, Merck
ACN, Chromasolv LC-MS, Honeywell
Ethyl acetate (EtAc), Sigma-Aldrich, Chromalsolv Plus 99%
Formic Acid, THP, LC-MS Grade
Table 7 LC-MRM LC-MS Details
| Property | Specifications
| Mass spectrometry | Bruker EVOQ Elite
| LC | Bruker Elute
| Injection volume (indicative) | 2 µL
| Weak wash | ACN/IPA 4:3
| Strong wash | IPA
| Pre-column | -
| Column | Phenomenex Kinetex C18 150 x 2.1 mm, 2.6 µm
| Column temperature | 60 °C
| Mobile phase A | ACN/IPA 4:3, 10 mM ammonium acetate, 0.1 % formic acid
| Mobile phase B | H2O deionized, 10 mM ammonium acetate, 0.1 % formic acid
Table 8 LC-MRM LC gradient
| Time (minutes) | A% | B% |  | Flow (µL/min)
| 0 | 15 | 85 |  | 500
| 1 | 15 | 85 |  | 500
| 3 | 0 | 100 |  | 500
| 5 | 0 | 100 |  | 500
| 6 | 15 | 85 |  | 500
| 7 | 15 | 85 |  | 500
Table 9 LC-MRM MS conditions (EXAMPLE: this only applies to Sciex QTRAP, please adapt)
| Polarity mode | Positive
| Ion spray | 3900 V
| Source temperature | 350 °C
| Heated Probe Gas | 50
| Nebulizer Gas | 50
| Cone Temp | 350 °C
| Cone Gas | 20
Table 10 LC-MRM Mass transitions used for absolute quantification. On different systems, RT and CE might be different!
| Analyte | Mass transitions | IS | Mass transitions | RT (min) | CE
| Cer d18:1/16:0 | 538.5 – 264.3 | D7- Cer d18:1/16:0 | 545.6 - 271.3 | 2.67 | 21
C34H68NO3 – C18H34N
| Cer d18:1/18:0 | 566.6 – 264.3 | D7- Cer d18:1/18:0 | 573.6 – 271.3 | 2.99 | 22
C36H72NO3 – C18H34N
| Cer d18:1/24:0 | 650.6 – 264.3 | D7- Cer d18:1/24:0 | 657.7 – 271.3 | 3.67 | 25
C42H84NO3 – C18H34N
| Cer d18:1/24:1 | 648.6 – 264.3 | D7- Cer d18:1/24:1 | 655.7 – 271.3 | 3.47 | 25
C42H82NO3 – C18H34N
Table 11 Data processing
| Processing software | Bruker MS Data Review V8.2

## LC-MRM - Sample Sequence
Sample Sequence (Example)
Sample Sequence (Example) | Sample Sequence 1
| Number | Sample Name | Remarks
| 1 | Total Blank 1
| 2 | Matrix blank
| 3 | Calibration line 1 from STD6 to 1 (n=3)
| 4 | […]
| 5 | Pooled QC
| 6 | Pooled QC
| 7 | SRM 1950-1
| 8 | SRM 1950-2
| 9 | SRM 1950-3
| 10 | SRM 1950-4
| 11 | SRM 1950-5
| 12 | SRM 1950-6
| 13 | Pooled QC
| 14 | Pooled QC
| 15 | T1D-1
| 16 | T1D-2
| 17 | T1D-3
| 18 | T1D-4
| 19 | T1D-5
| 20 | T1D-6
| 21 | Total Blank 2
| 22 | Pooled QC
| 23 | Pooled QC
| 24 | Young AA-1
| 25 | Young AA-2
| 26 | Young AA-3
| 27 | Young AA-4
| 28 | Young AA-5
| 29 | Young AA-6
| 30 | Pooled QC
| 31 | Pooled QC
| 32 | hTAG-1
| 33 | hTAG-2
| 34 | hTAG-3
| 35 | hTAG-4
| 36 | hTAG-5
| 37 | hTAG-6
| 38 | Pooled QC
| 39 | Pooled QC
| 40 | Calibration line 2 from STD1 to 6 (n=3)
| 41 | […]
| 42 | Matrix blank
| 43 | Total Blank 3

## LC-MRM - Intra Assay QC Res
INTRA ASSAY - QC RESULTS
INTRA ASSAY - QC RESULTS | Table 12 LC-MRM Calibration Line 1 Data (Beginning). Please report the values of each of the three replicates separately for each compound and standard.
| Calibration Line 1 | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| STD 1 | 158900000 | 173600000 | 175500000 | 168500000 | 155500000 | 193900000 | 1801000000 | 2390000000 | 2469000000 | 1843000000 | 2202000000 | 2320000000 | 6784000 | 7015000 | 7058000 | 5916000 | 5316000 | 6704000 | 138600000 | 143300000 | 152500000 | 45510000 | 50150000 | 53870000
| STD 2 | 85360000 | 78120000 | 84140000 | 85230000 | 90000000 | 84920000 | 1230000000 | 1184000000 | 1072000000 | 1020000000 | 1022000000 | 1022000000 | 7054000 | 6580000 | 7483000 | 6565000 | 6466000 | 7344000 | 142000000 | 153600000 | 166700000 | 54390000 | 51850000 | 55690000
| STD 3 | 6064000 | 7715000 | 8063000 | 6761000 | 8292000 | 6215000 | 95250000 | 105600000 | 98150000 | 84320000 | 96360000 | 86470000 | 5523000 | 7214000 | 7184000 | 5883000 | 6788000 | 6131000 | 128900000 | 150400000 | 136300000 | 38860000 | 52010000 | 49630000
| STD 4 | 1646000 | 1882000 | 6008000 | 1550000 | 1980000 | 2565000 | 22120000 | 24090000 | 26110000 | 18840000 | 19800000 | 23110000 | 7714000 | 7719000 | 6482000 | 6947000 | 7688000 | 6030000 | 142700000 | 167400000 | 149100000 | 49160000 | 51500000 | 50890000
| STD 5 | 1248000 | 1188000 | 817262 | 756977 | 1051000 | 595454 | 12690000 | 11370000 | 10390000 | 9646000 | 8508000 | 8013000 | 6760000 | 6898000 | 5695000 | 5668000 | 7181000 | 6026000 | 131200000 | 131700000 | 145700000 | 46190000 | 45980000 | 49270000
| STD 6 | 1023000 | 1008000 | 895349 | 853428 | 894241 | 874442 | 11220000 | 9926000 | 9495000 | 8239000 | 7610000 | 6959000 | 7811000 | 7938000 | 7606000 | 6572000 | 6615000 | 7154000 | 157500000 | 150300000 | 142900000 | 54400000 | 55410000 | 50500000
Table 13 LC-MRM Calibration Line 2 Data (End). Please report the values of each of the three replicates separately for each compound and standard.
| Calibration Line 2 | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| STD 1 | 178700000 | 194600000 | 176000000 | 185200000 | 207000000 | 186300000 | 1959000000 | 2550000000 | 2138000000 | 2071000000 | 2598000000 | 1917000000 | 6710000 | 8170000 | 6652000 | 7079000 | 6930000 | 6138000 | 147700000 | 155600000 | 141000000 | 55050000 | 62100000 | 50920000
| STD 2 | 84160000 | 97090000 | 96450000 | 95700000 | 107500000 | 99220000 | 1297000000 | 1393000000 | 1397000000 | 1045000000 | 1285000000 | 1248000000 | 8023000 | 8606000 | 7974000 | 7828000 | 8173000 | 7898000 | 171700000 | 174600000 | 164600000 | 56720000 | 56440000 | 55770000
| STD 3 | 8207000 | 8521000 | 7231000 | 8810000 | 9341000 | 7430000 | 113100000 | 118500000 | 108000000 | 104600000 | 111900000 | 86270000 | 8049000 | 8016000 | 6141000 | 7161000 | 7274000 | 6073000 | 165200000 | 164200000 | 123300000 | 57460000 | 59220000 | 46000000
| STD 4 | 5826000 | 1962000 | 1686000 | 2567000 | 1882000 | 1637000 | 24960000 | 22520000 | 22780000 | 22340000 | 21290000 | 19360000 | 7351000 | 7259000 | 7741000 | 7472000 | 6402000 | 7544000 | 155000000 | 161500000 | 155700000 | 51750000 | 56210000 | 55230000
| STD 5 | 969229 | 914288 | 1242000 | 841009 | 831870 | 886845 | 11650000 | 11110000 | 13740000 | 9520000 | 9242000 | 12080000 | 6786000 | 7138000 | 6928000 | 6568000 | 6872000 | 6656000 | 141700000 | 148600000 | 140700000 | 51640000 | 54830000 | 49610000
| STD 6 | 1080000 | 1038000 | 1187000 | 783157 | 666418 | 748435 | 10860000 | 11560000 | 10030000 | 8919000 | 7842000 | 7544000 | 8154000 | 8830000 | 8174000 | 6776000 | 7954000 | 7062000 | 164500000 | 178900000 | 174800000 | 63900000 | 57980000 | 55100000
Table 14 LC-MRM Sample Data. Please report the values of each of the three replicates separately for each compound and standard.
| NIST SRM | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| SRM 1 | 11440000 |  |  | 4579000 |  |  | 155500000 |  |  | 50950000 |  |  | 7765000 |  |  | 7313000 |  |  | 147500000 |  |  | 56590000
| SRM 2 | 10270000 |  |  | 4313000 |  |  | 151200000 |  |  | 46460000 |  |  | 6549000 |  |  | 6946000 |  |  | 149800000 |  |  | 56570000
| SRM 3 | 9504000 |  |  | 3528000 |  |  | 147600000 |  |  | 45930000 |  |  | 7096000 |  |  | 6836000 |  |  | 140600000 |  |  | 55100000
| SRM 4 | 10950000 |  |  | 3618000 |  |  | 160000000 |  |  | 48180000 |  |  | 7198000 |  |  | 6471000 |  |  | 149400000 |  |  | 59760000
| SRM 5 | 10500000 |  |  | 4534000 |  |  | 159600000 |  |  | 51430000 |  |  | 8123000 |  |  | 8330000 |  |  | 173600000 |  |  | 59350000
| SRM 6 | 10730000 |  |  | 4062000 |  |  | 148500000 |  |  | 49000000 |  |  | 7180000 |  |  | 7035000 |  |  | 152900000 |  |  | 56650000
| NIST hTAG | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| hTAG 1 | 16680000 |  |  | 6837000 |  |  | 313900000 |  |  | 98120000 |  |  | 9578000 |  |  | 9469000 |  |  | 190000000 |  |  | 69100000
| hTAG 2 | 17310000 |  |  | 5399000 |  |  | 316700000 |  |  | 89920000 |  |  | 8862000 |  |  | 8315000 |  |  | 199500000 |  |  | 68200000
| hTAG 3 | 18350000 |  |  | 6604000 |  |  | 307600000 |  |  | 101600000 |  |  | 10960000 |  |  | 9284000 |  |  | 206800000 |  |  | 71210000
| hTAG 4 | 13340000 |  |  | 5264000 |  |  | 241800000 |  |  | 71950000 |  |  | 7499000 |  |  | 7783000 |  |  | 160100000 |  |  | 56950000
| hTAG 5 | 17610000 |  |  | 6229000 |  |  | 283000000 |  |  | 89230000 |  |  | 10290000 |  |  | 8801000 |  |  | 184300000 |  |  | 61140000
| hTAG 6 | 15360000 |  |  | 5365000 |  |  | 258500000 |  |  | 82530000 |  |  | 10080000 |  |  | 8509000 |  |  | 179600000 |  |  | 61600000
| NIST T1D | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| T1D 1 | 10360000 |  |  | 4853000 |  |  | 163900000 |  |  | 52560000 |  |  | 8288000 |  |  | 8150000 |  |  | 179500000 |  |  | 61670000
| T1D 2 | 10370000 |  |  | 4724000 |  |  | 163100000 |  |  | 56430000 |  |  | 8620000 |  |  | 7956000 |  |  | 176100000 |  |  | 67160000
| T1D 3 | 11080000 |  |  | 5183000 |  |  | 148900000 |  |  | 54130000 |  |  | 7899000 |  |  | 7732000 |  |  | 166200000 |  |  | 57980000
| T1D 4 | 11750000 |  |  | 5753000 |  |  | 170000000 |  |  | 56280000 |  |  | 8398000 |  |  | 8556000 |  |  | 182000000 |  |  | 61310000
| T1D 5 | 10500000 |  |  | 5257000 |  |  | 179000000 |  |  | 51250000 |  |  | 8180000 |  |  | 8974000 |  |  | 176800000 |  |  | 62920000
| T1D 6 | 10720000 |  |  | 4855000 |  |  | 164700000 |  |  | 49870000 |  |  | 7790000 |  |  | 7335000 |  |  | 161700000 |  |  | 58080000
| NIST Young AA | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| yA-A 1 | 5647000 |  |  | 6837000 |  |  | 313900000 |  |  | 98120000 |  |  | 9578000 |  |  | 9469000 |  |  | 190000000 |  |  | 69100000
| yA-A 2 | 7331000 |  |  | 5399000 |  |  | 316700000 |  |  | 89920000 |  |  | 8862000 |  |  | 8315000 |  |  | 199500000 |  |  | 68200000
| yA-A 3 | 6621000 |  |  | 6604000 |  |  | 307600000 |  |  | 101600000 |  |  | 10960000 |  |  | 9284000 |  |  | 206800000 |  |  | 71210000
| yA-A 4 | 7209000 |  |  | 5264000 |  |  | 241800000 |  |  | 71950000 |  |  | 7499000 |  |  | 7783000 |  |  | 160100000 |  |  | 56950000
| yA-A 5 | 6825000 |  |  | 6229000 |  |  | 283000000 |  |  | 89230000 |  |  | 10290000 |  |  | 8801000 |  |  | 184300000 |  |  | 61140000
| yA-A 6 | 7286000 |  |  | 5365000 |  |  | 258500000 |  |  | 82530000 |  |  | 10080000 |  |  | 8509000 |  |  | 179600000 |  |  | 61600000
Table 15 LC-MRM Intra Assay QC Data. Please report the values of each of the three replicates separately for each compound and standard.
| Lowest Level QC | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| LLQC 1 | 2105000 | 3384000 | 2936000 | 645481 | 1349000 | 1069000 | 29190000 | 52240000 | 48910000 | 8149000 | 14200000 | 13550000 | 7427000 | 11140000 | 11220000 | 7436000 | 11210000 | 10450000 | 149200000 | 252700000 | 250600000 | 51510000 | 85450000 | 86600000
| LLQC 2 | 1817000 | 3652000 | 3048000 | 824097 | 1397000 | 1199000 | 30370000 | 55200000 | 41260000 | 9274000 | 16450000 | 12410000 | 7859000 | 12960000 | 11180000 | 7177000 | 12880000 | 10190000 | 145200000 | 325900000 | 225800000 | 50500000 | 105800000 | 82310000
| LLQC 3 | 2190000 | 3639000 | 2733000 | 699317 | 1344000 | 867319 | 25890000 | 48290000 | 40600000 | 7042000 | 14950000 | 10830000 | 7498000 | 13140000 | 11610000 | 6689000 | 12430000 | 10740000 | 143000000 | 294500000 | 244000000 | 50210000 | 92180000 | 84450000
| LLQC 4 | 1221000 | 2663000 | 2344000 | 642470 | 1331000 | 584288 | 24420000 | 46470000 | 34490000 | 7485000 | 13800000 | 11420000 | 7411000 | 14410000 | 10020000 | 7312000 | 13810000 | 8962000 | 143900000 | 300200000 | 238400000 | 49400000 | 102100000 | 75630000
| LLQC 5 | 1759000 | 3499000 | 2571000 | 892877 | 1251000 | 1034000 | 28770000 | 54720000 | 45590000 | 8734000 | 15760000 | 14220000 | 6522000 | 12120000 | 9996000 | 5679000 | 11090000 | 9483000 | 133900000 | 260200000 | 234200000 | 48360000 | 87410000 | 73450000
| LLQC 6 | 1545000 | 3297000 | 2651000 | 586365 | 1192000 | 1011000 | 23880000 | 50170000 | 40960000 | 6730000 | 13170000 | 12180000 | 7113000 | 14290000 | 10570000 | 6535000 | 13950000 | 10900000 | 133400000 | 313800000 | 240700000 | 49340000 | 100700000 | 79920000
| Low QC | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| LQC 1 | 5019000 | 9289000 | 7526000 | 2006000 | 3651000 | 2915000 | 78400000 | 146900000 | 115400000 | 22860000 | 48330000 | 36640000 | 6927000 | 13610000 | 10900000 | 6805000 | 13390000 | 9507000 | 140300000 | 312300000 | 235700000 | 49740000 | 100200000 | 74780000
| LQC 2 | 4795000 | 8588000 | 7666000 | 1855000 | 3021000 | 2591000 | 76070000 | 134800000 | 114700000 | 23580000 | 43770000 | 37740000 | 7432000 | 13190000 | 11910000 | 7481000 | 12460000 | 11520000 | 159300000 | 312400000 | 237500000 | 52690000 | 105400000 | 80220000
| LQC 3 | 5756000 | 8186000 | 5702000 | 1861000 | 3642000 | 2171000 | 74670000 | 142100000 | 105900000 | 24170000 | 43000000 | 31810000 | 8090000 | 12550000 | 10100000 | 7973000 | 13370000 | 8801000 | 153400000 | 280000000 | 207300000 | 53780000 | 92850000 | 67630000
| LQC 4 | 5025000 | 9543000 | 8091000 | 1770000 | 3517000 | 3515000 | 73500000 | 154700000 | 129800000 | 22830000 | 44080000 | 41050000 | 7420000 | 15280000 | 12150000 | 7349000 | 14360000 | 11790000 | 153400000 | 294900000 | 259000000 | 52100000 | 97430000 | 87620000
| LQC 5 | 5169000 | 8788000 | 7458000 | 1914000 | 3271000 | 2740000 | 81280000 | 150400000 | 121000000 | 26680000 | 45050000 | 37670000 | 7200000 | 11680000 | 9936000 | 6404000 | 11730000 | 10020000 | 140900000 | 256300000 | 214700000 | 48080000 | 79480000 | 69520000
| LQC 6 | 5389000 | 8659000 | 7442000 | 1782000 | 3689000 | 3258000 | 79620000 | 132000000 | 125700000 | 22800000 | 38760000 | 39550000 | 7271000 | 11350000 | 12200000 | 6495000 | 12080000 | 11460000 | 152500000 | 252800000 | 244500000 | 49270000 | 85480000 | 85060000
| Middle QC | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| MQC 1 | 10790000 | 15800000 | 16530000 | 3531000 | 7424000 | 6208000 | 162500000 | 271300000 | 281300000 | 47310000 | 81810000 | 85030000 | 7310000 | 12800000 | 12410000 | 6622000 | 11820000 | 12610000 | 138000000 | 251200000 | 269700000 | 49440000 | 87910000 | 89940000
| MQC 2 | 10170000 | 17230000 | 16220000 | 3664000 | 6129000 | 6258000 | 161400000 | 284200000 | 272100000 | 49170000 | 82140000 | 78020000 | 8028000 | 11060000 | 11200000 | 6854000 | 11020000 | 10930000 | 139400000 | 266000000 | 256900000 | 50150000 | 81620000 | 79660000
| MQC 3 | 9413000 | 20250000 | 17710000 | 3434000 | 7837000 | 7078000 | 143000000 | 315300000 | 275300000 | 44090000 | 91380000 | 83780000 | 6417000 | 14520000 | 13050000 | 6399000 | 13910000 | 12450000 | 156900000 | 308100000 | 268800000 | 46280000 | 104700000 | 90520000
| MQC 4 | 10710000 | 17370000 | 16640000 | 3621000 | 6405000 | 6465000 | 156300000 | 320700000 | 264300000 | 49180000 | 95240000 | 80910000 | 8067000 | 14060000 | 11730000 | 6741000 | 12980000 | 11670000 | 165600000 | 317200000 | 252000000 | 53710000 | 107300000 | 83710000
| MQC 5 | 11360000 | 16530000 | 16560000 | 3551000 | 6472000 | 5747000 | 147500000 | 277500000 | 276000000 | 51180000 | 80590000 | 77750000 | 7174000 | 12260000 | 12020000 | 6728000 | 11950000 | 10650000 | 152100000 | 262000000 | 241600000 | 54980000 | 82870000 | 81390000
| MQC 6 | 9381000 | 17130000 | 13340000 | 3367000 | 6293000 | 5637000 | 143800000 | 301900000 | 231600000 | 43800000 | 79450000 | 72010000 | 7171000 | 13130000 | 10990000 | 6642000 | 12010000 | 10180000 | 139700000 | 267000000 | 228900000 | 45930000 | 89940000 | 72410000
| High QC | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| HQC 1 | 25730000 | 22840000 | 26930000 | 11630000 | 13750000 | 13520000 | 370600000 | 420000000 | 422900000 | 157800000 | 177700000 | 160700000 | 10510000 | 8977000 | 10360000 | 8963000 | 9374000 | 11130000 | 195000000 | 222400000 | 228600000 | 64090000 | 76430000 | 73060000
| HQC 2 | 19980000 | 25510000 | 23510000 | 13710000 | 13220000 | 13870000 | 359200000 | 332200000 | 392700000 | 162800000 | 156200000 | 166100000 | 9893000 | 11080000 | 11300000 | 11210000 | 10480000 | 11540000 | 236100000 | 233200000 | 245600000 | 80630000 | 78660000 | 91360000
| HQC 3 | 19220000 | 24570000 | 23260000 | 13530000 | 12540000 | 13940000 | 357400000 | 373000000 | 404100000 | 173800000 | 172700000 | 173500000 | 10010000 | 12750000 | 11690000 | 11960000 | 10950000 | 12290000 | 260800000 | 241900000 | 277600000 | 91950000 | 91230000 | 88780000
| HQC 4 | 23600000 | 23910000 | 23520000 | 12570000 | 13860000 | 12970000 | 340400000 | 356000000 | 362600000 | 152200000 | 165100000 | 151500000 | 13120000 | 11820000 | 13100000 | 10360000 | 12160000 | 12040000 | 233300000 | 268900000 | 239500000 | 84480000 | 88680000 | 78020000
| HQC 5 | 28350000 | 25520000 | 24460000 | 13940000 | 12850000 | 13890000 | 363000000 | 386000000 | 374000000 | 158500000 | 161000000 | 171900000 | 14360000 | 13150000 | 12380000 | 12330000 | 12150000 | 11310000 | 250200000 | 238400000 | 246900000 | 84420000 | 84350000 | 93640000
| HQC 6 | 21090000 | 20360000 | 17960000 | 11710000 | 11200000 | 13110000 | 315800000 | 347400000 | 333800000 | 155400000 | 144100000 | 141200000 | 12010000 | 11960000 | 10540000 | 11810000 | 10950000 | 13180000 | 246400000 | 261000000 | 285300000 | 82280000 | 85260000 | 82220000
| Highest Level QC | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| HLQC 1 | 32690000 | 30700000 | 30670000 | 22100000 | 20000000 | 20720000 | 513400000 | 472500000 | 472900000 | 261900000 | 261500000 | 245600000 | 12490000 | 13120000 | 12170000 | 12230000 | 12110000 | 10960000 | 248000000 | 254800000 | 267700000 | 88860000 | 92370000 | 90670000
| HLQC 2 | 31180000 | 33000000 | 29950000 | 20590000 | 21810000 | 20990000 | 465500000 | 486400000 | 468700000 | 272400000 | 234400000 | 260600000 | 12630000 | 13690000 | 11680000 | 11770000 | 12500000 | 11300000 | 288200000 | 251100000 | 241900000 | 95720000 | 85180000 | 89580000
| HLQC 3 | 29310000 | 30780000 | 27440000 | 21480000 | 20230000 | 19720000 | 448700000 | 462400000 | 432100000 | 249500000 | 235800000 | 262300000 | 13160000 | 11950000 | 12010000 | 11860000 | 10900000 | 10180000 | 268600000 | 258600000 | 253800000 | 90230000 | 81390000 | 96020000
| HLQC 4 | 27890000 | 33400000 | 31620000 | 22270000 | 21710000 | 19100000 | 498600000 | 488700000 | 422100000 | 280400000 | 248500000 | 232200000 | 11500000 | 11330000 | 12250000 | 11700000 | 12030000 | 10480000 | 272900000 | 256000000 | 242600000 | 87530000 | 86030000 | 86650000
| HLQC 5 | 29120000 | 28780000 | 35980000 | 20980000 | 22910000 | 18700000 | 448200000 | 585400000 | 454900000 | 247700000 | 301700000 | 266100000 | 11890000 | 12980000 | 13050000 | 12570000 | 12760000 | 10530000 | 243400000 | 275800000 | 222600000 | 89970000 | 97060000 | 82740000
| HLQC 6 | 29280000 | 35260000 | 31560000 | 22460000 | 22730000 | 18530000 | 480000000 | 497400000 | 448200000 | 264000000 | 259000000 | 229200000 | 10380000 | 13580000 | 12230000 | 12880000 | 11880000 | 10480000 | 274200000 | 296500000 | 232100000 | 95380000 | 89150000 | 76820000
| Blank QC | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| Total Blank 1 | 0 | 0 | 1267000 | 0 | 0 | 0 | 181370 | 211433 | 760827 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| Total Blank 2 | 143990 | 0 | 1193000 | 0 | 0 | 0 | 630772 | 339171 | 53148 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| Total Blank 3 | 645222 | 1131000 | 572380 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0

## LC-MRM - Intra Assay QC Proc
Intra Assay Procedure
Intra Assay Procedure | Quality Controls for Intra-Assay Variation
| Step | Description | Reference Table
| 1 | Prepare the Pooled NIST QC sample containing equal volumes of the four NIST reference plasma you received.
| 2 | The pooled QC representing middle quality control (MQC), low QC (LQC), and lower limit of quantification QC (LLOQ) is prepared by diluting MQC (e.g. two- and three-fold) with water.
| 3 | High QC (HQC) and upper limit of quantification (ULOQ) QC are prepared by spiking in 10 µL and 20 µL respectively of STD3 endogenous standards from table | Extraction - Overview'!B34
| 4 | Internal standards are added to all QC samples.
Intra-Assay Variation
| 6 | The precision and accuracy of the assay is determined for each of the four ceramides measured in LLOQ, LQC, MQC, HQC, and HLOQ.
| 7 | Intra-assay variation, precision, and accuracy is calculated for each of the ceramide in replicates of six (n = 6) at each QC concentration independently on three separate occasions.
| 8 | The intra-assay precision (percentage coefficient variance, %CV) and accuracy (percentage accuracy, %Accuracy) are calculated from the nominal concentrations according to the formulas reported in corresponding sheets and tables of this document.
| 9 | Recovery estimation is not necessary for the published method but only for different ones. Stability estimation is not necessary for the published method but only for different ones. Calibration line is prepared in 5% BSA.

## Raw Peak Areas
Sample_Name | Cer d18:1/16:0 [M>264] | Cer d18:1/16:0-d7 [M>264] | Cer d18:1/18:0 [M>264] | Cer d18:1/18:0-d7 [M>264] | Cer d18:1/24:0 [M>264] | Cer d18:1/24:0-d7 [M>264] | Cer d18:1/24:1 [M>264] | Cer d18:1/24:1-d7 [M>264]
Total Blank 1 | 0 | 0 | 0 | 0 | 942636 | 0 | 0 | 0
Matrix blank | 9586000 | 0 | 3381000 | 0 | 154700000 | 0 | 43440000 | 0
Calibration line 1 from STD6 to 6 (n=3)
STD6_3 | 1023000 | 7811000 | 853428 | 6572000 | 11220000 | 157500000 | 8239000 | 54400000
STD6_2 | 1008000 | 7938000 | 894241 | 6615000 | 9926000 | 150300000 | 7610000 | 55410000
STD6 | 895349 | 7606000 | 874442 | 7154000 | 9495000 | 142900000 | 6959000 | 50500000
STD5_3 | 1248000 | 6760000 | 756977 | 5668000 | 12690000 | 131200000 | 9646000 | 46190000
STD5_2 | 1188000 | 6898000 | 1051000 | 7181000 | 11370000 | 131700000 | 8508000 | 45980000
STD5 | 817262 | 5695000 | 595454 | 6026000 | 10390000 | 145700000 | 8013000 | 49270000
STD4_3 | 1646000 | 7714000 | 1550000 | 6947000 | 22120000 | 142700000 | 18840000 | 49160000
STD4_2 | 1882000 | 7719000 | 1980000 | 7688000 | 24090000 | 167400000 | 19800000 | 51500000
STD4 | 6008000 | 6482000 | 2565000 | 6030000 | 26110000 | 149100000 | 23110000 | 50890000
STD3_3 | 6064000 | 5523000 | 6761000 | 5883000 | 95250000 | 128900000 | 84320000 | 38860000
STD3_2 | 7715000 | 7214000 | 8292000 | 6788000 | 105600000 | 150400000 | 96360000 | 52010000
STD3 | 8063000 | 7184000 | 6215000 | 6131000 | 98150000 | 136300000 | 86470000 | 49630000
STD2_3 | 85360000 | 7054000 | 85230000 | 6565000 | 1230000000 | 142000000 | 1020000000 | 54390000
STD2_2 | 78120000 | 6580000 | 90000000 | 6466000 | 1184000000 | 153600000 | 1022000000 | 51850000
STD2 | 84140000 | 7483000 | 84920000 | 7344000 | 1072000000 | 166700000 | 1022000000 | 55690000
STD1_3 | 158900000 | 6784000 | 168500000 | 5916000 | 1801000000 | 138600000 | 1843000000 | 45510000
STD1_2 | 173600000 | 7015000 | 155500000 | 5316000 | 2390000000 | 143300000 | 2202000000 | 50150000
STD1 | 175500000 | 7058000 | 193900000 | 6704000 | 2469000000 | 152500000 | 2320000000 | 53870000
Pooled QC | 11320000 | 8421000 | 4018000 | 8118000 | 173200000 | 183800000 | 54710000 | 64400000
Pooled QC | 11810000 | 8176000 | 4048000 | 7760000 | 179100000 | 177800000 | 56840000 | 60450000
SRM 1950-1 | 11440000 | 7765000 | 4579000 | 7313000 | 155500000 | 147500000 | 50950000 | 56590000
SRM 1950-2 | 10270000 | 6549000 | 4313000 | 6946000 | 151200000 | 149800000 | 46460000 | 56570000
SRM 1950-3 | 9504000 | 7096000 | 3528000 | 6836000 | 147600000 | 140600000 | 45930000 | 55100000
SRM 1950-4 | 10950000 | 7198000 | 3618000 | 6471000 | 160000000 | 149400000 | 48180000 | 59760000
SRM 1950-5 | 10500000 | 8123000 | 4534000 | 8330000 | 159600000 | 173600000 | 51430000 | 59350000
SRM 1950-6 | 10730000 | 7180000 | 4062000 | 7035000 | 148500000 | 152900000 | 49000000 | 56650000
Pooled QC | 10470000 | 7704000 | 3815000 | 8063000 | 160000000 | 143900000 | 46630000 | 55160000
Pooled QC | 9810000 | 7988000 | 3932000 | 7555000 | 163200000 | 156800000 | 53550000 | 56600000
T1D-1 | 10360000 | 8288000 | 4853000 | 8150000 | 163900000 | 179500000 | 52560000 | 61670000
T1D-2 | 10370000 | 8620000 | 4724000 | 7956000 | 163100000 | 176100000 | 56430000 | 67160000
T1D-3 | 11080000 | 7899000 | 5183000 | 7732000 | 148900000 | 166200000 | 54130000 | 57980000
T1D-4 | 11750000 | 8398000 | 5753000 | 8556000 | 170000000 | 182000000 | 56280000 | 61310000
T1D-5 | 10500000 | 8180000 | 5257000 | 8974000 | 179000000 | 176800000 | 51250000 | 62920000
T1D-6 | 10720000 | 7790000 | 4855000 | 7335000 | 164700000 | 161700000 | 49870000 | 58080000
Total Blank 2 | 424235 | 0 | 0 | 0 | 1324000 | 0 | 0 | 0
Pooled QC | 8973000 | 6273000 | 3495000 | 6616000 | 138200000 | 130900000 | 46280000 | 44880000
Pooled QC | 11280000 | 8000000 | 3637000 | 6974000 | 156400000 | 145800000 | 52710000 | 53750000
Young AA-1 | 5647000 | 7902000 | 2657000 | 8131000 | 101700000 | 156400000 | 31610000 | 63650000
Young AA-2 | 7331000 | 8418000 | 2526000 | 6912000 | 103800000 | 164100000 | 30980000 | 59400000
Young AA-3 | 6621000 | 7879000 | 2071000 | 6524000 | 95880000 | 144800000 | 30680000 | 54420000
Young AA-4 | 7209000 | 7664000 | 2446000 | 6907000 | 115600000 | 156700000 | 33190000 | 56000000
Young AA-5 | 6825000 | 7134000 | 2318000 | 7053000 | 106700000 | 145100000 | 29280000 | 52710000
Young AA-6 | 7286000 | 7241000 | 2568000 | 7221000 | 99390000 | 157800000 | 28560000 | 55810000
Pooled QC | 8347000 | 8524000 | 3178000 | 7855000 | 135100000 | 174500000 | 42200000 | 53430000
Pooled QC | 8941000 | 7925000 | 3023000 | 7364000 | 135000000 | 171300000 | 44390000 | 62860000
hTAG-1 | 16680000 | 9578000 | 6837000 | 9469000 | 313900000 | 190000000 | 98120000 | 69100000
hTAG-2 | 17310000 | 8862000 | 5399000 | 8315000 | 316700000 | 199500000 | 89920000 | 68200000
hTAG-3 | 18350000 | 10960000 | 6604000 | 9284000 | 307600000 | 206800000 | 101600000 | 71210000
hTAG-4 | 13340000 | 7499000 | 5264000 | 7783000 | 241800000 | 160100000 | 71950000 | 56950000
hTAG-5 | 17610000 | 10290000 | 6229000 | 8801000 | 283000000 | 184300000 | 89230000 | 61140000
hTAG-6 | 15360000 | 10080000 | 5365000 | 8509000 | 258500000 | 179600000 | 82530000 | 61600000
Pooled QC | 10830000 | 8401000 | 4107000 | 7578000 | 153200000 | 158600000 | 48740000 | 56530000
Pooled QC | 9811000 | 8027000 | 3845000 | 7180000 | 160200000 | 158500000 | 49050000 | 56610000
Calibration line 2 from STD1 to 6 (n=3)
STD1 | 178700000 | 6710000 | 185200000 | 7079000 | 1959000000 | 147700000 | 2071000000 | 55050000
STD1_2 | 194600000 | 8170000 | 207000000 | 6930000 | 2550000000 | 155600000 | 2598000000 | 62100000
STD1_3 | 176000000 | 6652000 | 186300000 | 6138000 | 2138000000 | 141000000 | 1917000000 | 50920000
STD2 | 84160000 | 8023000 | 95700000 | 7828000 | 1297000000 | 171700000 | 1045000000 | 56720000
STD2_2 | 97090000 | 8606000 | 107500000 | 8173000 | 1393000000 | 174600000 | 1285000000 | 56440000
STD2_3 | 96450000 | 7974000 | 99220000 | 7898000 | 1397000000 | 164600000 | 1248000000 | 55770000
STD3 | 8207000 | 8049000 | 8810000 | 7161000 | 113100000 | 165200000 | 104600000 | 57460000
STD3_2 | 8521000 | 8016000 | 9341000 | 7274000 | 118500000 | 164200000 | 111900000 | 59220000
STD3_3 | 7231000 | 6141000 | 7430000 | 6073000 | 108000000 | 123300000 | 86270000 | 46000000
STD4 | 5826000 | 7351000 | 2567000 | 7472000 | 24960000 | 155000000 | 22340000 | 51750000
STD4_2 | 1962000 | 7259000 | 1882000 | 6402000 | 22520000 | 161500000 | 21290000 | 56210000
STD4_3 | 1686000 | 7741000 | 1637000 | 7544000 | 22780000 | 155700000 | 19360000 | 55230000
STD5 | 969229 | 6786000 | 841009 | 6568000 | 11650000 | 141700000 | 9520000 | 51640000
STD5_2 | 914288 | 7138000 | 831870 | 6872000 | 11110000 | 148600000 | 9242000 | 54830000
STD5_3 | 1242000 | 6928000 | 886845 | 6656000 | 13740000 | 140700000 | 12080000 | 49610000
STD6 | 1080000 | 8154000 | 783157 | 6776000 | 10860000 | 164500000 | 8919000 | 63900000
STD6_2 | 1038000 | 8830000 | 666418 | 7954000 | 11560000 | 178900000 | 7842000 | 57980000
STD6_3 | 1187000 | 8174000 | 748435 | 7062000 | 10030000 | 174800000 | 7544000 | 55100000
Matrix blank | 10510000 | 0 | 3929000 | 0 | 171600000 | 0 | 50440000 | 0
Total Blank 3 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0