## Participant Information
Participant Information
Participant Information |  | Group Name
|  | Lab code | 5
| Institution | Institution name
Street
City
Postal code
Country
| Contact person | Name
Address
Email
Phone

## Summary
Summary
Summary | This protocol describes the extraction and LC-MRM procedure for absolute quantification of Cer d18:1/16:0, Cer d18:1/18:0, Cer d18:1/24:0 and Cer d18:1/24:1 in human plasma according to Kauhanen et al, Anal Bioanal Chem (2016) 408:3475–3483.
A calibration curve is built by injecting several dilutions of the non-labelled standards mixed with a fixed concentration of labelled standards. Each calibration point is prepared and analysed in triplicates (n=3).
To establish the endogenous concentration of the four selected ceramides in the study samples (NIST SRM1950, NIST high TAG, NIST T1D, and NIST young A-A) each sample is extracted and analyzed in six replicates (n =6).
Prepare the pooled QC by mixing 100 mL each of NIST SRM1950, NIST high TAG, NIST T1D, and NIST young AA (total 400 mL).
We suggest proceeding first to measure the ceramides concentration in the four study samples and then continue with the intra-assay validation to avoid wasting material.

## Extraction - Overview
Extraction Overview
Extraction Overview | Table 1 - Extraction - Materials
Chemicals, solvents and materials
Ethyl acetate (EtAc)
2-propanol (IPA)
5% Bovine Serum Albumin (BSA)
96-well plates/tubes
Sealing foil
2 mL DeepWell plates/tubes/LC vials
Table 2 - Extraction - Sample Types
| Sample type | Components
| Pooled Quality Control samples (QC) | 10 µL pooled QC |  |  | 20 µL Labelled IS |  | 570 µL EtAc:IPA 2:8 (v/v)
| Matrix Blank (MB) | 10 µL Matrix (pooled QC) only |  |  | - |  | 590 µL EtAc:IPA 2:8
| Study Samples (S) | 10 µL Human plasma (NIST SRM1950, high TAG, T1D, young AA) |  |  | 20 µL Labelled IS |  | 570 µL EtAc:IPA 2:8
| Total Blank (TB) | 10 µL 5% BSA |  |  | - |  | 590 µL EtAc:IPA 2:8
| Calibration line standards (STD) | 10 µL 5% BSA + 20 µL non labelled STD mixture (Table 5) |  |  | 20 µL Labelled IS |  | 550 µL EtAc:IPA 2:8
Table 3 - Extraction - Labelled Internal Standard Mixture Solution (provided)
| Labelled IS mixture in EtAc:IPA 2:8 (v/v) | Concentration |  |  |  |  | Unit | Vendor
| D7-Cer d18:1/16:0 | 0.125 |  |  |  |  | pmol/µL [µM] | Avanti Polar Lipids
| D7-Cer d18:1/18:0 | 0.050
| D7-Cer d18:1/24:0 | 1500
| D7-Cer d18:1/24:1 | 0.500
Table 4 - Extraction - Non-Labelled Standards - for STD1 Mixture Solution (provided)
| Non-labelled standards in EtAc:IPA 2:8 (v/v) | Concentration |  |  |  |  | Unit | Vendor
| Cer d18:1/16:0 | 2.0 |  |  |  |  | pmol/µL [µM] | Avanti Polar Lipids
| Cer d18:1/18:0 | 2.0
| Cer d18:1/24:0 | 20.0
| Cer d18:1/24:1 | 20.0
Table 5 - Serial dilutions of Non-Labelled Standards from STD1 Mixture Solution for calibration line
| Add 20 µL of | STD1 Conc. | STD2 Conc. | STD3 Conc. | STD4 Conc. | STD5 Conc. | STD6 Conc. | Unit
| Cer d18:1/16:0 | 2.0 | 1.0 | 0.10 | 0.02 | 0.01 | 0.008 | pmol/µL [µM]
| Cer d18:1/18:0 | 2.0 | 1.0 | 0.10 | 0.02 | 0.01 | 0.008
| Cer d18:1/24:0 | 20.0 | 10.0 | 1.0 | 0.2 | 0.1 | 0.08
| Cer d18:1/24:1 | 20.0 | 10.0 | 1.0 | 0.2 | 0.1 | 0.08

## Extraction - Procedure
Extraction Procedure
Extraction Procedure | Extraction Procedure
| Step | Description | Reference Table
| 1 | Prepare different dilutions of STD1 by diluting the provided mixture (Extraction - Overview, Table 4) with EtAc:IPA 2:8 as suggested in table | Extraction - Overview'!B34
| 2 | Thaw plasma on ice (+4°C) and bring to room temperature prior to extraction. Mix by pipetting before use.
| 3 | Prepare samples in plates/tubes manually or with a robot, according to specifications in column C of table | Extraction - Overview'!C14
| 4 | Add 20 µL of labelled internal standard mixture to samples according to column F of table | Extraction - Overview'!F14
| 5 | For the calibration line, add 20 µL of different dilutions of the non-labelled mixture (STD1-6).
| 6 | Add 590 µL of ethyl acetate:isopropanol (2:8, vol/vol) to Total Blank (TB) and Matrix Blank (MB); 550 µL to calibration line (STD); 570 µL to study samples and QCs according to column H of table | Extraction - Overview'!H14
| 7 | Mix samples 10 min by pipetting or vortexing
| 8 | Centrifuge samples for 10 min at 3000 x gat room temperature
| 9 | Transfer 50 µL of the clear supernatant to MS-vials or MS plates, close/seal, store samples at -20 °C prior to LC-MRM analysis.

## LC-MRM - Overview
LC-MRM Overview
LC-MRM Overview | Table 6 Chemicals and solvents
Ammonium formate
Formic acid
Ultra-pure water
Acetonitrile
Table 7 LC-MRM LC-MS Details
| Property | Specifications
| Mass spectrometry | QQQ - SCIEX 6500+
| LC | UPLC - Shimadzu 20AD
| Injection volume (indicative) | 5 µL
| Weak wash | NA
| Strong wash | NA - wash in this system is just needle dip (1:1 MeOH:Water)
| Pre-column | None
| Column | Agilent Poroshell 120 EC-C18 2.7um 100x2.1mm (PN 895775-902)
| Column temperature | 60 °C
| Mobile phase A | 10 mM ammonium acetate in water with 0.1% formic acid
| Mobile phase B | 10 mM ammonium acetate in acetonitrile:2-propanol (4:3, vol/vol) with 0.1% formic acid
Table 8 LC-MRM LC gradient
| Time (minutes) | A% | B% |  | Flow (µL/min)
| 0.00 | 15 | 85 |  | 300
| 1 | 15 | 85 |  | 300
| 3 | 0 | 100 |  | 300
| 8 | 0 | 100 |  | 300
| 8.01 | 15 | 85 |  | 30
| 10 | 15 | 85 |  | 30
Table 9 LC-MRM MS conditions (EXAMPLE: this only applies to Sciex QTRAP, please adapt)
| Polarity mode | Positive
| Ion spray | 5000 V
| Curtain gas | 40 psi
| Source temperature | 300 °C
| Gas 1 | 50 psi
| Gas 2 | 35 psi
| DP | 30 V
| EP | 10 V
| Collision cell exit potential | 20 V
| Collision energy (CE) | as per below
Table 10 LC-MRM Mass transitions used for absolute quantification. On different systems, RT and CE might be different!
| Analyte | Mass transitions | IS | Mass transitions | RT (min) | CE
| Cer d18:1/16:0 | 538.5 – 264.3 - CE 40 | D7- Cer d18:1/16:0 | 545.6 - 271.3 | 1.9-2.0 | 40
C34H68NO3 – C18H34N
| Cer d18:1/18:0 | 566.6 – 264.3 - CE 40 | D7- Cer d18:1/18:0 | 573.6 – 271.3 | 2.1-2.2 | 40
C36H72NO3 – C18H34N
| Cer d18:1/24:0 | 650.6 – 264.3 - CE 65 | D7- Cer d18:1/24:0 | 657.7 – 271.3 | 2.4-2.5 | 40
C42H84NO3 – C18H34N
| Cer d18:1/24:1 | 648.6 – 264.3 - CE 60 | D7- Cer d18:1/24:1 | 655.7 – 271.3 | 2.3-2.4 | 40
C42H82NO3 – C18H34N
Table 11 Data processing
| Processing software | SCIEX Analyst MD 1.6

## LC-MRM - Sample Sequence
Sample Sequence (Example)
Sample Sequence (Example) | Sample Sequence 1
| Number | Sample Name | Remarks
| 1 | Total Blank 1
| 2 | Matrix blank
| 3 | Cal line 1 S6
| 4 | Cal line 1 S5
| 5 | Cal line 1 S4
| 6 | Cal line 1 S3
| 7 | Cal line 1 S2
| 8 | Cal line 1 S1
| 9 | Cal line 2 S6
| 10 | Cal line 2 S5
| 11 | Cal line 2 S4
| 12 | Cal line 2 S3
| 13 | Cal line 2 S2
| 14 | Cal line 2 S1
| 15 | Cal line 3 S6
| 16 | Cal line 3 S5
| 17 | Cal line 3 S4
| 18 | Cal line 3 S3
| 19 | Cal line 3 S2
| 20 | Cal line 3 S1
| 21 | Pooled QC
| 22 | Pooled QC
| 23 | MQC | Re-injected 2x for a total of 3 injections each
| 24 | MQC | Re-injected 2x for a total of 3 injections each
| 25 | LLQC | Re-injected 2x for a total of 3 injections each
| 26 | LLQC | Re-injected 2x for a total of 3 injections each
| 27 | LQC | Re-injected 2x for a total of 3 injections each
| 28 | LQC | Re-injected 2x for a total of 3 injections each
| 29 | HQC | Re-injected 2x for a total of 3 injections each
| 30 | HQC | Re-injected 2x for a total of 3 injections each
| 31 | HLQC | Re-injected 2x for a total of 3 injections each
| 32 | HLQC | Re-injected 2x for a total of 3 injections each
| 33 | SRM 1950-1 | Re-injected 2x for a total of 3 injections each
| 34 | SRM 1950-2 | Re-injected 2x for a total of 3 injections each
| 35 | SRM 1950-3 | Re-injected 2x for a total of 3 injections each
| 36 | SRM 1950-4 | Re-injected 2x for a total of 3 injections each
| 37 | SRM 1950-5 | Re-injected 2x for a total of 3 injections each
| 38 | SRM 1950-6 | Re-injected 2x for a total of 3 injections each
| 39 | Pooled QC
| 40 | Pooled QC
| 41 | MQC | Re-injected 2x for a total of 3 injections each
| 42 | MQC | Re-injected 2x for a total of 3 injections each
| 43 | LLQC | Re-injected 2x for a total of 3 injections each
| 44 | LLQC | Re-injected 2x for a total of 3 injections each
| 45 | LQC | Re-injected 2x for a total of 3 injections each
| 46 | LQC | Re-injected 2x for a total of 3 injections each
| 47 | HQC | Re-injected 2x for a total of 3 injections each
| 48 | HQC | Re-injected 2x for a total of 3 injections each
| 49 | HLQC | Re-injected 2x for a total of 3 injections each
| 50 | HLQC | Re-injected 2x for a total of 3 injections each
| 51 | T1D-1 | Re-injected 2x for a total of 3 injections each
| 52 | T1D-2 | Re-injected 2x for a total of 3 injections each
| 53 | T1D-3 | Re-injected 2x for a total of 3 injections each
| 54 | T1D-4 | Re-injected 2x for a total of 3 injections each
| 55 | T1D-5 | Re-injected 2x for a total of 3 injections each
| 56 | T1D-6 | Re-injected 2x for a total of 3 injections each
| 57 | Total Blank 2
| 58 | Pooled QC
| 59 | Pooled QC
| 60 | Young AA-1 | Re-injected 2x for a total of 3 injections each
| 61 | Young AA-2 | Re-injected 2x for a total of 3 injections each
| 62 | Young AA-3 | Re-injected 2x for a total of 3 injections each
| 63 | Young AA-4 | Re-injected 2x for a total of 3 injections each
| 64 | Young AA-5 | Re-injected 2x for a total of 3 injections each
| 65 | Young AA-6 | Re-injected 2x for a total of 3 injections each
| 66 | Pooled QC
| 67 | Pooled QC
| 68 | MQC | Re-injected 2x for a total of 3 injections each
| 69 | MQC | Re-injected 2x for a total of 3 injections each
| 70 | LLQC | Re-injected 2x for a total of 3 injections each
| 71 | LLQC | Re-injected 2x for a total of 3 injections each
| 72 | LQC | Re-injected 2x for a total of 3 injections each
| 73 | LQC | Re-injected 2x for a total of 3 injections each
| 74 | HQC | Re-injected 2x for a total of 3 injections each
| 75 | HQC | Re-injected 2x for a total of 3 injections each
| 76 | HLQC | Re-injected 2x for a total of 3 injections each
| 77 | HLQC | Re-injected 2x for a total of 3 injections each
| 78 | hTAG-1 | Re-injected 2x for a total of 3 injections each
| 79 | hTAG-2 | Re-injected 2x for a total of 3 injections each
| 80 | hTAG-3 | Re-injected 2x for a total of 3 injections each
| 81 | hTAG-4 | Re-injected 2x for a total of 3 injections each
| 82 | hTAG-5 | Re-injected 2x for a total of 3 injections each
| 83 | hTAG-6 | Re-injected 2x for a total of 3 injections each
| 84 | Pooled QC
| 85 | Pooled QC
| 86 | Cal line 4 S6
| 87 | Cal line 4 S5
| 88 | Cal line 4 S4
| 89 | Cal line 4 S3
| 90 | Cal line 4 S2
| 91 | Cal line 4 S1
| 92 | Cal line 5 S6
| 93 | Cal line 5 S5
| 94 | Cal line 5 S4
| 95 | Cal line 5 S3
| 96 | Cal line 5 S2
| 97 | Cal line 5 S1
| 98 | Cal line 6 S6
| 99 | Cal line 6 S5
| 100 | Cal line 6 S4
| 101 | Cal line 6 S3
| 102 | Cal line 6 S2
| 103 | Cal line 6 S1
| 104 | Matrix blank
| 105 | Total Blank 3

## LC-MRM - Intra Assay QC Proc
Intra Assay Procedure
Intra Assay Procedure | Quality Controls for Intra-Assay Variation
| Step | Description | Reference Table
| 1 | Prepare the Pooled NIST QC sample containing equal volumes of the four NIST reference plasma you received.
| 2 | The pooled QC representing middle quality control (MQC), low QC (LQC), and lower limit of quantification QC (LLOQ) is prepared by diluting MQC (e.g. two- and three-fold) with water.
| 3 | High QC (HQC) and upper limit of quantification (ULOQ) QC are prepared by spiking in 10 µL and 20 µL respectively of STD3 endogenous standards from table | Extraction - Overview'!B34
| 4 | Internal standards are added to all QC samples.
Intra-Assay Variation
| 6 | The precision and accuracy of the assay is determined for each of the four ceramides measured in LLOQ, LQC, MQC, HQC, and HLOQ.
| 7 | Intra-assay variation, precision, and accuracy is calculated for each of the ceramide in replicates of six (n = 6) at each QC concentration independently on three separate occasions.
| 8 | The intra-assay precision (percentage coefficient variance, %CV) and accuracy (percentage accuracy, %Accuracy) are calculated from the nominal concentrations according to the formulas reported in corresponding sheets and tables of this document.
| 9 | Recovery estimation is not necessary for the published method but only for different ones. Stability estimation is not necessary for the published method but only for different ones. Calibration line is prepared in 5% BSA.

## LC-MRM - Intra Assay QC Res
INTRA ASSAY - QC RESULTS
INTRA ASSAY - QC RESULTS | Table 12 LC-MRM Calibration Line 1 Data (Beginning). Please report the values of each of the three replicates separately for each compound and standard.
| Calibration Line 1 | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| STD 1 | 12082637 | 11385768 | 10429116 | 6676179 | 6620041 | 6384617 | 30652235 | 29085064 | 27392735 | 29661550 | 26214024 | 27071440 | 510261 | 491952 | 455176 | 299181 | 305588 | 300349 | 10586942 | 10604371 | 10048501 | 2593605 | 2551365 | 2557040
| STD 2 | 5797395 | 5595723 | 5954832 | 3097528 | 3346443 | 3468514 | 13560066 | 14313860 | 15348567 | 13555386 | 14645813 | 13586594 | 519396 | 502233 | 544164 | 301046 | 328405 | 313795 | 10808505 | 9319541 | 10335827 | 2701779 | 2900246 | 2669605
| STD 3 | 669182 | 673687 | 737218 | 355836 | 352583 | 367346 | 1749732 | 1567880 | 1596279 | 1626569 | 1634511 | 1565756 | 490545 | 506747 | 568130 | 306058 | 324902 | 320617 | 10275271 | 10143668 | 10232453 | 2815808 | 3010238 | 2831076
| STD 4 | 179134 | 174146 | 177579 | 64622 | 78980 | 77727 | 379016 | 335570 | 343861 | 332577 | 348740 | 334798 | 517969 | 511598 | 501976 | 312951 | 329220 | 314265 | 10790728 | 9691469 | 9968881 | 2959178 | 3115670 | 2882407
| STD 5 | 125621 | 143208 | 129018 | 42521 | 43883 | 41744 | 187213 | 187378 | 191492 | 182072 | 185785 | 164073 | 499732 | 504286 | 546662 | 317447 | 326054 | 336234 | 11246107 | 10125752 | 11658722 | 2865517 | 2949305 | 3146457
| STD 6 | 110813 | 121990 | 119846 | 34534 | 34032 | 38562 | 154632 | 155779 | 146181 | 153742 | 136045 | 121783 | 541013 | 523709 | 544391 | 357767 | 353704 | 327937 | 10586062 | 12488511 | 10948040 | 3694477 | 2854549 | 2716322
Table 13 LC-MRM Calibration Line 2 Data (End). Please report the values of each of the three replicates separately for each compound and standard.
| Calibration Line 2 | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| STD 1 | 9586001 | 9697098 | 8836843 | 7190498 | 6840343 | 6782790 | 27996349 | 27412441 | 26186426 | 28493233 | 29152699 | 27723440 | 440184 | 445014 | 434222 | 380143 | 339044 | 325012 | 9464889 | 9166301 | 9078214 | 2746527 | 2693014 | 2699043
| STD 2 | 4541696 | 4500370 | 4477437 | 3710950 | 3572211 | 3529956 | 13842245 | 14513356 | 13542426 | 15011034 | 14861746 | 14537508 | 439416 | 413715 | 442670 | 350098 | 326080 | 358874 | 9338455 | 9439707 | 8829206 | 2850886 | 2836059 | 2801136
| STD 3 | 540709 | 553048 | 582048 | 399593 | 377525 | 359197 | 1578172 | 1566274 | 1630274 | 1710730 | 1749875 | 1682758 | 429796 | 440539 | 454690 | 358146 | 351469 | 351848 | 10173599 | 10616561 | 10467091 | 2900546 | 3131775 | 3071388
| STD 4 | 163308 | 148930 | 158632 | 85430 | 80096 | 77220 | 344202 | 345272 | 346677 | 383620 | 338477 | 373774 | 454250 | 427395 | 463465 | 377882 | 352074 | 349471 | 10327035 | 10151367 | 10080590 | 2859165 | 2925397 | 2954301
| STD 5 | 108980 | 112972 | 105905 | 44211 | 44023 | 48439 | 175106 | 178041 | 179916 | 185735 | 159631 | 186763 | 453039 | 447382 | 444220 | 364760 | 343693 | 371383 | 10213456 | 9960270 | 10369821 | 2954896 | 2905798 | 3064654
| STD 6 | 107092 | 106708 | 116484 | 40000 | 35865 | 38501 | 149285 | 137834 | 144872 | 155047 | 148188 | 159262 | 478668 | 446221 | 515494 | 381046 | 355396 | 367372 | 10223169 | 10204947 | 10542342 | 3318859 | 3032690 | 3440082
Table 14 LC-MRM Sample Data. Please report the values of each of the three replicates separately for each compound and standard.
| NIST SRM | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| SRM 1 | 491249 | 505450 | 489983 | 158596 | 175383 | 158470 | 2039563 | 2072988 | 1794719 | 685363 | 713876 | 662490 | 352550 | 359480 | 348304 | 352843 | 327099 | 310487 | 10228412 | 10835536 | 9963402 | 2845790 | 2875756 | 2539866
| SRM 2 | 547423 | 545938 | 486645 | 159642 | 168300 | 148308 | 1874685 | 1819421 | 1850365 | 734455 | 723027 | 685033 | 354039 | 359495 | 346403 | 322236 | 295067 | 287208 | 9487562 | 8786314 | 9149947 | 2792725 | 2800455 | 2567650
| SRM 3 | 545957 | 539870 | 504467 | 162179 | 160240 | 150611 | 2057498 | 1939521 | 1810305 | 782254 | 771185 | 707466 | 362206 | 364970 | 353353 | 320943 | 314736 | 298844 | 9934964 | 9567280 | 9041989 | 3114746 | 3204664 | 2918913
| SRM 4 | 549249 | 538836 | 527327 | 162684 | 161570 | 157675 | 1911833 | 1797747 | 1660210 | 734175 | 735892 | 705615 | 361169 | 362595 | 351121 | 334697 | 317842 | 284787 | 9157789 | 9024310 | 8010264 | 2947739 | 2887513 | 2807436
| SRM 5 | 565010 | 537534 | 516180 | 163778 | 154283 | 160303 | 1950757 | 1717928 | 1630050 | 716703 | 711446 | 678797 | 371122 | 350776 | 341611 | 330044 | 310852 | 278710 | 9952394 | 8497921 | 8405061 | 3114227 | 2880370 | 2739768
| SRM 6 | 527512 | 527430 | 498655 | 158050 | 151157 | 148737 | 2026555 | 1652575 | 1663805 | 694529 | 680184 | 648768 | 357837 | 348358 | 342372 | 311371 | 294290 | 292522 | 10470569 | 7993921 | 8468636 | 2872764 | 2818410 | 2721965
| NIST hTAG | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| hTAG 1 | 489287 | 487761 | 482244 | 194276 | 195210 | 177631 | 2899401 | 2839706 | 2738870 | 1102006 | 1016374 | 1002354 | 307896 | 300396 | 299064 | 352249 | 344368 | 336543 | 9533561 | 9261549 | 9264558 | 2928251 | 2937011 | 2708110
| hTAG 2 | 551146 | 529149 | 572011 | 212449 | 207885 | 202200 | 3201589 | 2883444 | 3114922 | 1184925 | 1097688 | 1065857 | 307359 | 288237 | 321040 | 337820 | 337779 | 352472 | 9222398 | 8750912 | 9605376 | 2928336 | 2760975 | 2874241
| hTAG 3 | 531667 | 538284 | 529118 | 214801 | 193381 | 195577 | 3015091 | 2990589 | 3025155 | 1107274 | 969410 | 1035934 | 293721 | 302433 | 312616 | 358721 | 336212 | 326623 | 9325872 | 9416346 | 9686135 | 2902931 | 2679601 | 2667355
| hTAG 4 | 436522 | 458807 | 428466 | 178797 | 176206 | 157831 | 2421031 | 2357973 | 2533527 | 886000 | 857772 | 901269 | 295105 | 309506 | 306396 | 320728 | 339549 | 324461 | 9424959 | 9222578 | 9297760 | 2774716 | 2816216 | 2830675
| hTAG 5 | 516969 | 555284 | 537447 | 206878 | 219454 | 198323 | 3088499 | 3236191 | 2821720 | 1152412 | 1135739 | 1124023 | 286590 | 297684 | 304178 | 348897 | 354916 | 356627 | 9203468 | 9666998 | 8887549 | 2922459 | 2869172 | 2813086
| hTAG 6 | 488107 | 533239 | 509521 | 204814 | 206442 | 191331 | 2783164 | 2999804 | 2919982 | 1063344 | 1024629 | 1047173 | 274123 | 310729 | 295670 | 343407 | 358321 | 326321 | 8278806 | 9455890 | 9310229 | 2947674 | 2916231 | 2824466
| NIST T1D | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| T1D 1 | 453817 | 451271 | 433287 | 198945 | 183421 | 178286 | 1963515 | 1945715 | 1868393 | 655214 | 661884 | 674846 | 331129 | 328993 | 326648 | 312421 | 304989 | 303765 | 9811535 | 9806792 | 9743589 | 2512041 | 2646049 | 2616228
| T1D 2 | 460958 | 433336 | 438933 | 174959 | 168710 | 172003 | 1632920 | 1764701 | 1656978 | 675335 | 641105 | 682629 | 355938 | 343929 | 331094 | 294322 | 272871 | 287793 | 8058048 | 8687714 | 8315848 | 2908129 | 2771504 | 2837089
| T1D 3 | 475247 | 433888 | 418243 | 180499 | 170591 | 174460 | 1597088 | 1838044 | 1606570 | 693037 | 641647 | 666843 | 345730 | 326839 | 331256 | 288978 | 272510 | 275876 | 7596632 | 9328625 | 8303481 | 2765462 | 2465269 | 2739311
| T1D 4 | 473246 | 437321 | 437358 | 174403 | 161503 | 171348 | 1566450 | 1781352 | 1810733 | 667765 | 634224 | 656027 | 358329 | 336848 | 339464 | 290001 | 267851 | 267677 | 7875384 | 8756750 | 9312334 | 2720465 | 2417464 | 2661627
| T1D 5 | 476996 | 440679 | 426115 | 179680 | 158549 | 161616 | 1531738 | 1709255 | 1727824 | 689513 | 649573 | 635415 | 361014 | 343450 | 340229 | 279585 | 286524 | 270619 | 8126231 | 8806969 | 9142735 | 2774101 | 2654921 | 2494525
| T1D 6 | 475616 | 456482 | 456904 | 184415 | 165804 | 176553 | 1722795 | 1796716 | 1865560 | 681813 | 654254 | 661639 | 355067 | 319275 | 337666 | 287772 | 271643 | 270283 | 8245324 | 8796504 | 9458340 | 2787101 | 2515949 | 2691717
| NIST Young AA | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| yA-A 1 | 349202 | 345419 | 352923 | 83239 | 81969 | 77404 | 1173867 | 1248802 | 1295002 | 425455 | 382959 | 417723 | 383156 | 348609 | 349348 | 279877 | 269648 | 263100 | 7816087 | 8773850 | 8359046 | 2784164 | 2404104 | 2532306
| yA-A 2 | 337630 | 333039 | 346006 | 87839 | 82411 | 85283 | 1183429 | 1258396 | 1228443 | 409532 | 400186 | 410268 | 354420 | 358039 | 345427 | 289161 | 269932 | 278525 | 8434410 | 8487522 | 8353680 | 2688305 | 2475915 | 2682150
| yA-A 3 | 356496 | 333309 | 346060 | 86292 | 80001 | 83207 | 1224507 | 1234310 | 1260194 | 433211 | 403330 | 426114 | 361856 | 340563 | 365877 | 286493 | 269900 | 287978 | 8552359 | 8857375 | 8500302 | 2765145 | 2312923 | 2581643
| yA-A 4 | 347447 | 337932 | 323992 | 89151 | 84516 | 90040 | 1255201 | 1280206 | 1148760 | 423455 | 412915 | 414507 | 370121 | 369813 | 354539 | 297388 | 266169 | 288992 | 8191414 | 8655451 | 8382411 | 2634314 | 2601143 | 2561211
| yA-A 5 | 329375 | 328993 | 336568 | 82641 | 80357 | 85522 | 1189742 | 1238937 | 1243890 | 392110 | 389388 | 398915 | 350256 | 356933 | 354470 | 275720 | 280476 | 269123 | 8571515 | 8546570 | 8355498 | 2509309 | 2506488 | 2467177
| yA-A 6 | 332903 | 324008 | 331790 | 79661 | 79682 | 80859 | 1167745 | 1224914 | 1100175 | 376687 | 391457 | 369190 | 359368 | 366472 | 360022 | 263127 | 264093 | 271898 | 8285415 | 8564793 | 8139113 | 2345759 | 2526155 | 2363585
Table 15 LC-MRM Intra Assay QC Data. Please report the values of each of the three replicates separately for each compound and standard.
| Lowest Level QC | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| LLQC 1 | 230660 | 226918 | 212987 | 58178 | 57739 | 54217 | 764606 | 727574 | 679530 | 242425 | 231327 | 226000 | 492862 | 472082 | 462079 | 322116 | 346005 | 327166 | 10527435 | 11084414 | 10126929 | 2811131 | 2865552 | 2655057
| LLQC 2 | 224790 | 213781 | 218616 | 55526 | 55593 | 55393 | 690115 | 730239 | 626064 | 258261 | 236530 | 237083 | 465943 | 466175 | 456093 | 339668 | 316783 | 294624 | 9350463 | 9190032 | 8142840 | 2925677 | 2883111 | 2659468
| LLQC 3 | 190805 | 190762 | 172128 | 45780 | 45022 | 41231 | 588016 | 576949 | 538928 | 197606 | 194207 | 184806 | 441318 | 479236 | 467591 | 312216 | 308838 | 301563 | 10171532 | 10077119 | 9542808 | 2664421 | 2598251 | 2504124
| LLQC 4 | 221301 | 214640 | 211666 | 54407 | 46865 | 51704 | 613784 | 611901 | 660792 | 224021 | 224341 | 230980 | 449196 | 453123 | 428456 | 294686 | 300964 | 289570 | 8201792 | 8307419 | 9118485 | 2537225 | 2534236 | 2449425
| LLQC 5 | 200758 | 205016 | 190512 | 51588 | 59727 | 56992 | 721247 | 800814 | 737760 | 255586 | 296717 | 296807 | 422121 | 430172 | 411168 | 361153 | 379903 | 386594 | 10085311 | 10784187 | 10539413 | 3243167 | 3444204 | 3364755
| LLQC 6 | 194607 | 185662 | 186026 | 56333 | 60658 | 54797 | 742768 | 814766 | 793592 | 268609 | 296404 | 276870 | 384060 | 412526 | 407425 | 349141 | 390061 | 369997 | 10544147 | 10514035 | 10366881 | 3038005 | 3241601 | 3145699
| Low QC | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| LQC 1 | 319118 | 305635 | 290642 | 83328 | 80805 | 72166 | 1031594 | 1030135 | 922677 | 383122 | 359106 | 344687 | 439416 | 437639 | 432204 | 324724 | 318526 | 302277 | 9178971 | 9361131 | 8661986 | 3020282 | 2873916 | 2752403
| LQC 2 | 313147 | 299409 | 295831 | 79369 | 74591 | 82051 | 1037966 | 988572 | 940445 | 367297 | 345831 | 342106 | 434848 | 425009 | 442006 | 328467 | 293044 | 299630 | 9487505 | 9033414 | 8730489 | 2778416 | 2673163 | 2605245
| LQC 3 | 272245 | 272900 | 268882 | 76150 | 76829 | 69479 | 834860 | 826904 | 902929 | 321156 | 320217 | 325186 | 408001 | 418685 | 393410 | 290051 | 301805 | 301439 | 8124406 | 8022183 | 8700526 | 2472565 | 2574348 | 2740136
| LQC 4 | 289362 | 283534 | 277034 | 74066 | 65484 | 69316 | 841758 | 877751 | 909048 | 321902 | 331299 | 343164 | 443702 | 420146 | 405688 | 298646 | 288446 | 298914 | 8073743 | 7734364 | 8152768 | 2565503 | 2677063 | 2575637
| LQC 5 | 241664 | 240634 | 258015 | 83616 | 80693 | 75952 | 1059734 | 1103162 | 1074163 | 380513 | 401025 | 371840 | 383464 | 363273 | 386197 | 357352 | 376943 | 358545 | 10419596 | 10887221 | 10457075 | 3099273 | 3253324 | 3040456
| LQC 6 | 236095 | 254901 | 257638 | 79741 | 87962 | 87286 | 1050095 | 1149712 | 1111564 | 381540 | 401158 | 377121 | 362065 | 376865 | 382374 | 349882 | 368510 | 359141 | 9927272 | 10797149 | 10097249 | 2831396 | 3170648 | 2831090
| Middle QC | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| MQC 1 | 550732 | 503156 | 495438 | 158396 | 161002 | 149602 | 2024187 | 2028549 | 1850816 | 742550 | 744031 | 746145 | 389119 | 387485 | 367285 | 333963 | 321518 | 294537 | 9341429 | 8792916 | 8636169 | 3240775 | 2842105 | 2947193
| MQC 2 | 538791 | 526582 | 527269 | 156951 | 151448 | 159387 | 2052030 | 2015763 | 1919648 | 775349 | 729354 | 727860 | 384658 | 378822 | 371638 | 323667 | 326321 | 305174 | 8999023 | 8687966 | 8414710 | 3102620 | 3118762 | 2897937
| MQC 3 | 500606 | 506458 | 477532 | 160387 | 176606 | 151304 | 1797055 | 1879108 | 1959718 | 715640 | 750069 | 685584 | 350275 | 364338 | 326745 | 275061 | 304468 | 270877 | 8546826 | 8194412 | 9123131 | 2529550 | 2751227 | 2461815
| MQC 4 | 499164 | 475327 | 469103 | 148511 | 148401 | 141331 | 1880137 | 1662419 | 1745747 | 702277 | 676951 | 652257 | 361960 | 362760 | 335564 | 295320 | 289188 | 262125 | 8824083 | 7639021 | 8731692 | 2586099 | 2650314 | 2499997
| MQC 5 | 444485 | 451716 | 437541 | 155382 | 156325 | 174959 | 2092005 | 2157051 | 2146195 | 810235 | 881644 | 843486 | 308369 | 339696 | 311208 | 334359 | 354595 | 351685 | 9250851 | 9473711 | 10002930 | 2937653 | 3299293 | 3028334
| MQC 6 | 428203 | 415170 | 430769 | 162384 | 172459 | 162958 | 2076417 | 2208160 | 2087912 | 763018 | 840215 | 764507 | 302670 | 304114 | 315360 | 349099 | 384823 | 357592 | 9395643 | 10178530 | 9819841 | 2934584 | 3223801 | 3017817
| High QC | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| HQC 1 | 728255 | 737445 | 687528 | 361150 | 349696 | 326847 | 2877757 | 2886193 | 2713501 | 1644203 | 1516964 | 1454484 | 374029 | 363577 | 342484 | 340014 | 317177 | 289545 | 9311818 | 9372810 | 8641246 | 3022731 | 2804306 | 2533534
| HQC 2 | 775451 | 723191 | 730601 | 350003 | 345571 | 328853 | 2947548 | 2664412 | 2639982 | 1661785 | 1542521 | 1477115 | 394288 | 373951 | 363558 | 323252 | 329849 | 285057 | 9223803 | 8760686 | 8511167 | 3092385 | 2743974 | 2763111
| HQC 3 | 687988 | 656888 | 661559 | 337254 | 314087 | 317862 | 2498228 | 2441628 | 2327837 | 1378244 | 1379261 | 1490290 | 338448 | 342040 | 337713 | 284571 | 276103 | 286090 | 8304003 | 7956858 | 7918156 | 2597547 | 2575055 | 2641101
| HQC 4 | 705806 | 684680 | 666885 | 330841 | 325264 | 317873 | 2465369 | 2405498 | 2340615 | 1429101 | 1446041 | 1443867 | 358431 | 346251 | 352119 | 270099 | 266600 | 283746 | 8388924 | 7920406 | 7271555 | 2625073 | 2803845 | 2813251
| HQC 5 | 585758 | 571561 | 639180 | 361032 | 373245 | 391086 | 2775350 | 2887888 | 2869271 | 1718877 | 1672891 | 1741790 | 297027 | 306532 | 319422 | 351880 | 329598 | 337876 | 9563625 | 9868959 | 9679397 | 2864953 | 2952529 | 2983015
| HQC 6 | 625971 | 647709 | 632882 | 388589 | 396589 | 384187 | 3019016 | 2846794 | 2957167 | 1716962 | 1757665 | 1801417 | 314586 | 325640 | 317900 | 366445 | 381646 | 345033 | 9455789 | 9494004 | 9759203 | 3021799 | 2980296 | 2936207
| Highest Level QC | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| HLQC 1 | 978814 | 953291 | 858544 | 494096 | 495867 | 474773 | 3501263 | 3233944 | 3221059 | 2424156 | 2258300 | 2053305 | 385459 | 354383 | 355502 | 323903 | 305399 | 281100 | 8969702 | 8760380 | 8589472 | 3140821 | 2949089 | 2774278
| HLQC 2 | 1008098 | 944220 | 920290 | 530310 | 530833 | 511325 | 3568701 | 3393570 | 3145160 | 2359197 | 2336476 | 2113027 | 376742 | 378854 | 362585 | 325372 | 296366 | 292553 | 9130522 | 8883488 | 8341917 | 3121625 | 2982941 | 2691538
| HLQC 3 | 848395 | 838741 | 819791 | 479096 | 483945 | 441390 | 3091614 | 2850451 | 2958615 | 1955436 | 1981400 | 2051541 | 348521 | 351308 | 321560 | 284655 | 285699 | 276125 | 8289276 | 7282581 | 7860131 | 2641906 | 2709209 | 2524307
| HLQC 4 | 865103 | 835595 | 843885 | 476195 | 476480 | 472801 | 3111279 | 2891496 | 3162010 | 2077741 | 2146255 | 1978214 | 358694 | 342459 | 335839 | 273614 | 285349 | 264762 | 8026201 | 7547113 | 8294792 | 2732883 | 2661145 | 2668199
| HLQC 5 | 738301 | 738905 | 770222 | 547579 | 537129 | 543506 | 3538035 | 3700536 | 3651628 | 2456610 | 2567406 | 2508691 | 305576 | 296237 | 298914 | 333348 | 365059 | 362229 | 9837947 | 9472870 | 9672053 | 2989201 | 3071109 | 2864546
| HLQC 6 | 762829 | 742047 | 752462 | 537078 | 553460 | 549143 | 3568856 | 3552982 | 3609652 | 2521038 | 2630476 | 2501564 | 297350 | 300690 | 308913 | 362672 | 333381 | 343819 | 9949731 | 9654569 | 8964731 | 3095600 | 2986052 | 2923361
| Blank QC | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| Blank | 585708 |  |  | 175621 |  |  | 2527364 |  |  | 829612
| Blank | 453531 |  |  | 156425 |  |  | 1954704 |  |  | 774636
Blank

## Raw Peak Areas
Sample No | Sample Name | Note | Cer d18:1/16:0 [M>264] | Cer d18:1/16:0-d7 [M>264] | Cer d18:1/18:0 [M>264] | Cer d18:1/18:0-d7 [M>264] | Cer d18:1/24:0 [M>264] | Cer d18:1/24:0-d7 [M>264] | Cer d18:1/24:1 [M>264] | Cer d18:1/24:1-d7 [M>264]
1 | Total Blank 1 |  | 71335 | 0 | 5918 | 0 | 7012 | 0 | 2793 | 0
2 | Matrix blank |  | 585708 | 0 | 175621 | 0 | 2527364 | 0 | 829612 | 0
3 | STD6 rep1 |  | 110813 | 541013 | 34534 | 357767 | 154632 | 10586062 | 153742 | 3694477
4 | STD5 rep1 |  | 125621 | 499732 | 42521 | 317447 | 187213 | 11246107 | 182072 | 2865517
5 | STD4 rep1 |  | 179134 | 517969 | 64622 | 312951 | 379016 | 10790728 | 332577 | 2959178
6 | STD3 rep1 |  | 669182 | 490545 | 355836 | 306058 | 1749732 | 10275271 | 1626569 | 2815808
7 | STD2 rep1 |  | 5797395 | 519396 | 3097528 | 301046 | 13560066 | 10808505 | 13555386 | 2701779
8 | STD1 rep1 |  | 12082637 | 510261 | 6676179 | 299181 | 30652235 | 10586942 | 29661550 | 2593605
9 | STD6 rep2 |  | 121990 | 523709 | 34032 | 353704 | 155779 | 12488511 | 136045 | 2854549
10 | STD5 rep2 |  | 143208 | 504286 | 43883 | 326054 | 187378 | 10125752 | 185785 | 2949305
11 | STD4 rep2 |  | 174146 | 511598 | 78980 | 329220 | 335570 | 9691469 | 348740 | 3115670
12 | STD3 rep2 |  | 673687 | 506747 | 352583 | 324902 | 1567880 | 10143668 | 1634511 | 3010238
13 | STD2 rep2 |  | 5595723 | 502233 | 3346443 | 328405 | 14313860 | 9319541 | 14645813 | 2900246
14 | STD1 rep2 |  | 11385768 | 491952 | 6620041 | 305588 | 29085064 | 10604371 | 26214024 | 2551365
15 | STD6 rep3 |  | 119846 | 544391 | 38562 | 327937 | 146181 | 10948040 | 121783 | 2716322
16 | STD5 rep3 |  | 129018 | 546662 | 41744 | 336234 | 191492 | 11658722 | 164073 | 3146457
17 | STD4 rep3 |  | 177579 | 501976 | 77727 | 314265 | 343861 | 9968881 | 334798 | 2882407
18 | STD3 rep3 |  | 737218 | 568130 | 367346 | 320617 | 1596279 | 10232453 | 1565756 | 2831076
19 | STD2 rep3 |  | 5954832 | 544164 | 3468514 | 313795 | 15348567 | 10335827 | 13586594 | 2669605
20 | STD1 rep3 |  | 10429116 | 455176 | 6384617 | 300349 | 27392735 | 10048501 | 27071440 | 2557040
21 | Pooled QC |  | 526187 | 365872 | 168773 | 339174 | 2249138 | 10920963 | 746102 | 2881806
22 | Pooled QC |  | 543992 | 395555 | 172707 | 317957 | 2349877 | 10152335 | 821603 | 3173626
23 | MQC |  | 550732 | 389119 | 158396 | 333963 | 2024187 | 9341429 | 742550 | 3240775
24 | MQC |  | 538791 | 384658 | 156951 | 323667 | 2052030 | 8999023 | 775349 | 3102620
25 | LLQC |  | 230660 | 492862 | 58178 | 322116 | 764606 | 10527435 | 242425 | 2811131
26 | LLQC |  | 224790 | 465943 | 55526 | 339668 | 690115 | 9350463 | 258261 | 2925677
27 | LQC |  | 319118 | 439416 | 83328 | 324724 | 1031594 | 9178971 | 383122 | 3020282
28 | LQC |  | 313147 | 434848 | 79369 | 328467 | 1037966 | 9487505 | 367297 | 2778416
29 | HQC |  | 728255 | 374029 | 361150 | 340014 | 2877757 | 9311818 | 1644203 | 3022731
30 | HQC |  | 775451 | 394288 | 350003 | 323252 | 2947548 | 9223803 | 1661785 | 3092385
31 | HLQC |  | 978814 | 385459 | 494096 | 323903 | 3501263 | 8969702 | 2424156 | 3140821
32 | HLQC |  | 1008098 | 376742 | 530310 | 325372 | 3568701 | 9130522 | 2359197 | 3121625
33 | SRM 1950-1 |  | 491249 | 352550 | 158596 | 352843 | 2039563 | 10228412 | 685363 | 2845790
34 | SRM 1950-2 |  | 547423 | 354039 | 159642 | 322236 | 1874685 | 9487562 | 734455 | 2792725
35 | SRM 1950-3 |  | 545957 | 362206 | 162179 | 320943 | 2057498 | 9934964 | 782254 | 3114746
36 | SRM 1950-4 |  | 549249 | 361169 | 162684 | 334697 | 1911833 | 9157789 | 734175 | 2947739
37 | SRM 1950-5 |  | 565010 | 371122 | 163778 | 330044 | 1950757 | 9952394 | 716703 | 3114227
38 | SRM 1950-6 |  | 527512 | 357837 | 158050 | 311371 | 2026555 | 10470569 | 694529 | 2872764
39 | MQC | reinj 1 of #23 above | 503156 | 387485 | 161002 | 321518 | 2028549 | 8792916 | 744031 | 2842105
40 | MQC | reinj 1 of #24 above | 526582 | 378822 | 151448 | 326321 | 2015763 | 8687966 | 729354 | 3118762
41 | LLQC | reinj 1 of #25 above | 226918 | 472082 | 57739 | 346005 | 727574 | 11084414 | 231327 | 2865552
42 | LLQC | reinj 1 of #26 above | 213781 | 466175 | 55593 | 316783 | 730239 | 9190032 | 236530 | 2883111
43 | LQC | reinj 1 of #27 above | 305635 | 437639 | 80805 | 318526 | 1030135 | 9361131 | 359106 | 2873916
44 | LQC | reinj 1 of #28 above | 299409 | 425009 | 74591 | 293044 | 988572 | 9033414 | 345831 | 2673163
45 | HQC | reinj 1 of #29 above | 737445 | 363577 | 349696 | 317177 | 2886193 | 9372810 | 1516964 | 2804306
46 | HQC | reinj 1 of #30 above | 723191 | 373951 | 345571 | 329849 | 2664412 | 8760686 | 1542521 | 2743974
47 | HLQC | reinj 1 of #31 above | 953291 | 354383 | 495867 | 305399 | 3233944 | 8760380 | 2258300 | 2949089
48 | HLQC | reinj 1 of #32 above | 944220 | 378854 | 530833 | 296366 | 3393570 | 8883488 | 2336476 | 2982941
49 | SRM 1950-1 | reinj 1 of #33 above | 505450 | 359480 | 175383 | 327099 | 2072988 | 10835536 | 713876 | 2875756
50 | SRM 1950-2 | reinj 1 of #34 above | 545938 | 359495 | 168300 | 295067 | 1819421 | 8786314 | 723027 | 2800455
51 | SRM 1950-3 | reinj 1 of #35 above | 539870 | 364970 | 160240 | 314736 | 1939521 | 9567280 | 771185 | 3204664
52 | SRM 1950-4 | reinj 1 of #36 above | 538836 | 362595 | 161570 | 317842 | 1797747 | 9024310 | 735892 | 2887513
53 | SRM 1950-5 | reinj 1 of #37 above | 537534 | 350776 | 154283 | 310852 | 1717928 | 8497921 | 711446 | 2880370
54 | SRM 1950-6 | reinj 1 of #38 above | 527430 | 348358 | 151157 | 294290 | 1652575 | 7993921 | 680184 | 2818410
55 | MQC | reinj 2 of #23 above | 495438 | 367285 | 149602 | 294537 | 1850816 | 8636169 | 746145 | 2947193
56 | MQC | reinj 2 of #24 above | 527269 | 371638 | 159387 | 305174 | 1919648 | 8414710 | 727860 | 2897937
57 | LLQC | reinj 2 of #25 above | 212987 | 462079 | 54217 | 327166 | 679530 | 10126929 | 226000 | 2655057
58 | LLQC | reinj 2 of #26 above | 218616 | 456093 | 55393 | 294624 | 626064 | 8142840 | 237083 | 2659468
59 | LQC | reinj 2 of #27 above | 290642 | 432204 | 72166 | 302277 | 922677 | 8661986 | 344687 | 2752403
60 | LQC | reinj 2 of #28 above | 295831 | 442006 | 82051 | 299630 | 940445 | 8730489 | 342106 | 2605245
61 | HQC | reinj 2 of #29 above | 687528 | 342484 | 326847 | 289545 | 2713501 | 8641246 | 1454484 | 2533534
62 | HQC | reinj 2 of #30 above | 730601 | 363558 | 328853 | 285057 | 2639982 | 8511167 | 1477115 | 2763111
63 | HLQC | reinj 2 of #31 above | 858544 | 355502 | 474773 | 281100 | 3221059 | 8589472 | 2053305 | 2774278
64 | HLQC | reinj 2 of #32 above | 920290 | 362585 | 511325 | 292553 | 3145160 | 8341917 | 2113027 | 2691538
65 | SRM 1950-1 | reinj 2 of #33 above | 489983 | 348304 | 158470 | 310487 | 1794719 | 9963402 | 662490 | 2539866
66 | SRM 1950-2 | reinj 2 of #34 above | 486645 | 346403 | 148308 | 287208 | 1850365 | 9149947 | 685033 | 2567650
67 | SRM 1950-3 | reinj 2 of #35 above | 504467 | 353353 | 150611 | 298844 | 1810305 | 9041989 | 707466 | 2918913
68 | SRM 1950-4 | reinj 2 of #36 above | 527327 | 351121 | 157675 | 284787 | 1660210 | 8010264 | 705615 | 2807436
69 | SRM 1950-5 | reinj 2 of #37 above | 516180 | 341611 | 160303 | 278710 | 1630050 | 8405061 | 678797 | 2739768
70 | SRM 1950-6 | reinj 2 of #38 above | 498655 | 342372 | 148737 | 292522 | 1663805 | 8468636 | 648768 | 2721965
71 | Pooled QC |  | 461769 | 356613 | 157375 | 299048 | 1920213 | 9755097 | 661275 | 2552955
72 | Pooled QC |  | 507546 | 345797 | 158730 | 275315 | 1872219 | 8645030 | 689745 | 2638321
73 | MQC |  | 500606 | 350275 | 160387 | 275061 | 1797055 | 8546826 | 715640 | 2529550
74 | MQC |  | 499164 | 361960 | 148511 | 295320 | 1880137 | 8824083 | 702277 | 2586099
75 | LLQC |  | 190805 | 441318 | 45780 | 312216 | 588016 | 10171532 | 197606 | 2664421
76 | LLQC |  | 221301 | 449196 | 54407 | 294686 | 613784 | 8201792 | 224021 | 2537225
77 | LQC |  | 272245 | 408001 | 76150 | 290051 | 834860 | 8124406 | 321156 | 2472565
78 | LQC |  | 289362 | 443702 | 74066 | 298646 | 841758 | 8073743 | 321902 | 2565503
79 | HQC |  | 687988 | 338448 | 337254 | 284571 | 2498228 | 8304003 | 1378244 | 2597547
80 | HQC |  | 705806 | 358431 | 330841 | 270099 | 2465369 | 8388924 | 1429101 | 2625073
81 | HLQC |  | 848395 | 348521 | 479096 | 284655 | 3091614 | 8289276 | 1955436 | 2641906
82 | HLQC |  | 865103 | 358694 | 476195 | 273614 | 3111279 | 8026201 | 2077741 | 2732883
83 | T1D-1 |  | 453817 | 331129 | 198945 | 312421 | 1963515 | 9811535 | 655214 | 2512041
84 | T1D-2 |  | 460958 | 355938 | 174959 | 294322 | 1632920 | 8058048 | 675335 | 2908129
85 | T1D-3 |  | 475247 | 345730 | 180499 | 288978 | 1597088 | 7596632 | 693037 | 2765462
86 | T1D-4 |  | 473246 | 358329 | 174403 | 290001 | 1566450 | 7875384 | 667765 | 2720465
87 | T1D-5 |  | 476996 | 361014 | 179680 | 279585 | 1531738 | 8126231 | 689513 | 2774101
88 | T1D-6 |  | 475616 | 355067 | 184415 | 287772 | 1722795 | 8245324 | 681813 | 2787101
89 | MQC | reinj 1 of #73 above | 506458 | 364338 | 176606 | 304468 | 1879108 | 8194412 | 750069 | 2751227
90 | MQC | reinj 1 of #74 above | 475327 | 362760 | 148401 | 289188 | 1662419 | 7639021 | 676951 | 2650314
91 | LLQC | reinj 1 of #75 above | 190762 | 479236 | 45022 | 308838 | 576949 | 10077119 | 194207 | 2598251
92 | LLQC | reinj 1 of #76 above | 214640 | 453123 | 46865 | 300964 | 611901 | 8307419 | 224341 | 2534236
93 | LQC | reinj 1 of #77 above | 272900 | 418685 | 76829 | 301805 | 826904 | 8022183 | 320217 | 2574348
94 | LQC | reinj 1 of #78 above | 283534 | 420146 | 65484 | 288446 | 877751 | 7734364 | 331299 | 2677063
95 | HQC | reinj 1 of #79 above | 656888 | 342040 | 314087 | 276103 | 2441628 | 7956858 | 1379261 | 2575055
96 | HQC | reinj 1 of #80 above | 684680 | 346251 | 325264 | 266600 | 2405498 | 7920406 | 1446041 | 2803845
97 | HLQC | reinj 1 of #81 above | 838741 | 351308 | 483945 | 285699 | 2850451 | 7282581 | 1981400 | 2709209
98 | HLQC | reinj 1 of #82 above | 835595 | 342459 | 476480 | 285349 | 2891496 | 7547113 | 2146255 | 2661145
99 | T1D-1 | reinj 1 of #83 above | 451271 | 328993 | 183421 | 304989 | 1945715 | 9806792 | 661884 | 2646049
100 | T1D-2 | reinj 1 of #84 above | 433336 | 343929 | 168710 | 272871 | 1764701 | 8687714 | 641105 | 2771504
101 | T1D-3 | reinj 1 of #85 above | 433888 | 326839 | 170591 | 272510 | 1838044 | 9328625 | 641647 | 2465269
102 | T1D-4 | reinj 1 of #86 above | 437321 | 336848 | 161503 | 267851 | 1781352 | 8756750 | 634224 | 2417464
103 | T1D-5 | reinj 1 of #87 above | 440679 | 343450 | 158549 | 286524 | 1709255 | 8806969 | 649573 | 2654921
104 | T1D-6 | reinj 1 of #88 above | 456482 | 319275 | 165804 | 271643 | 1796716 | 8796504 | 654254 | 2515949
105 | MQC | reinj 2 of #73 above | 477532 | 326745 | 151304 | 270877 | 1959718 | 9123131 | 685584 | 2461815
106 | MQC | reinj 2 of #74 above | 469103 | 335564 | 141331 | 262125 | 1745747 | 8731692 | 652257 | 2499997
107 | LLQC | reinj 2 of #75 above | 172128 | 467591 | 41231 | 301563 | 538928 | 9542808 | 184806 | 2504124
108 | LLQC | reinj 2 of #76 above | 211666 | 428456 | 51704 | 289570 | 660792 | 9118485 | 230980 | 2449425
109 | LQC | reinj 2 of #77 above | 268882 | 393410 | 69479 | 301439 | 902929 | 8700526 | 325186 | 2740136
110 | LQC | reinj 2 of #78 above | 277034 | 405688 | 69316 | 298914 | 909048 | 8152768 | 343164 | 2575637
111 | HQC | reinj 2 of #79 above | 661559 | 337713 | 317862 | 286090 | 2327837 | 7918156 | 1490290 | 2641101
112 | HQC | reinj 2 of #80 above | 666885 | 352119 | 317873 | 283746 | 2340615 | 7271555 | 1443867 | 2813251
113 | HLQC | reinj 2 of #81 above | 819791 | 321560 | 441390 | 276125 | 2958615 | 7860131 | 2051541 | 2524307
114 | HLQC | reinj 2 of #82 above | 843885 | 335839 | 472801 | 264762 | 3162010 | 8294792 | 1978214 | 2668199
115 | T1D-1 | reinj 2 of #83 above | 433287 | 326648 | 178286 | 303765 | 1868393 | 9743589 | 674846 | 2616228
116 | T1D-2 | reinj 2 of #84 above | 438933 | 331094 | 172003 | 287793 | 1656978 | 8315848 | 682629 | 2837089
117 | T1D-3 | reinj 2 of #85 above | 418243 | 331256 | 174460 | 275876 | 1606570 | 8303481 | 666843 | 2739311
118 | T1D-4 | reinj 2 of #86 above | 437358 | 339464 | 171348 | 267677 | 1810733 | 9312334 | 656027 | 2661627
119 | T1D-5 | reinj 2 of #87 above | 426115 | 340229 | 161616 | 270619 | 1727824 | 9142735 | 635415 | 2494525
120 | T1D-6 | reinj 2 of #88 above | 456904 | 337666 | 176553 | 270283 | 1865560 | 9458340 | 661639 | 2691717
121 | Total Blank 2 |  | 65351 | 0 | 0 | 0 | 0 | 0 | 711 | 0
122 | Pooled QC |  | 416345 | 327337 | 142923 | 278193 | 1795474 | 7920176 | 650964 | 2542328
123 | Pooled QC |  | 472235 | 343046 | 141240 | 281598 | 1848712 | 7814067 | 693457 | 2839173
124 | Young AA-1 |  | 349202 | 383156 | 83239 | 279877 | 1173867 | 7816087 | 425455 | 2784164
125 | Young AA-2 |  | 337630 | 354420 | 87839 | 289161 | 1183429 | 8434410 | 409532 | 2688305
126 | Young AA-3 |  | 356496 | 361856 | 86292 | 286493 | 1224507 | 8552359 | 433211 | 2765145
127 | Young AA-4 |  | 347447 | 370121 | 89151 | 297388 | 1255201 | 8191414 | 423455 | 2634314
128 | Young AA-5 |  | 329375 | 350256 | 82641 | 275720 | 1189742 | 8571515 | 392110 | 2509309
129 | Young AA-6 |  | 332903 | 359368 | 79661 | 263127 | 1167745 | 8285415 | 376687 | 2345759
130 | Young AA-1 | reinj 1 of #124 above | 345419 | 348609 | 81969 | 269648 | 1248802 | 8773850 | 382959 | 2404104
131 | Young AA-2 | reinj 1 of #125 above | 333039 | 358039 | 82411 | 269932 | 1258396 | 8487522 | 400186 | 2475915
132 | Young AA-3 | reinj 1 of #126 above | 333309 | 340563 | 80001 | 269900 | 1234310 | 8857375 | 403330 | 2312923
133 | Young AA-4 | reinj 1 of #127 above | 337932 | 369813 | 84516 | 266169 | 1280206 | 8655451 | 412915 | 2601143
134 | Young AA-5 | reinj 1 of #128 above | 328993 | 356933 | 80357 | 280476 | 1238937 | 8546570 | 389388 | 2506488
135 | Young AA-6 | reinj 1 of #129 above | 324008 | 366472 | 79682 | 264093 | 1224914 | 8564793 | 391457 | 2526155
136 | Young AA-1 | reinj 2 of #124 above | 352923 | 349348 | 77404 | 263100 | 1295002 | 8359046 | 417723 | 2532306
137 | Young AA-2 | reinj 2 of #125 above | 346006 | 345427 | 85283 | 278525 | 1228443 | 8353680 | 410268 | 2682150
138 | Young AA-3 | reinj 2 of #126 above | 346060 | 365877 | 83207 | 287978 | 1260194 | 8500302 | 426114 | 2581643
139 | Young AA-4 | reinj 2 of #127 above | 323992 | 354539 | 90040 | 288992 | 1148760 | 8382411 | 414507 | 2561211
140 | Young AA-5 | reinj 2 of #128 above | 336568 | 354470 | 85522 | 269123 | 1243890 | 8355498 | 398915 | 2467177
141 | Young AA-6 | reinj 2 of #129 above | 331790 | 360022 | 80859 | 271898 | 1100175 | 8139113 | 369190 | 2363585
142 | Pooled QC |  | 460543 | 346638 | 158333 | 353978 | 2026464 | 9861908 | 787480 | 3162784
143 | Pooled QC |  | 424360 | 313414 | 150684 | 333531 | 1987101 | 9715748 | 788153 | 2994272
144 | MQC |  | 444485 | 308369 | 155382 | 334359 | 2092005 | 9250851 | 810235 | 2937653
145 | MQC |  | 428203 | 302670 | 162384 | 349099 | 2076417 | 9395643 | 763018 | 2934584
146 | LLQC |  | 200758 | 422121 | 51588 | 361153 | 721247 | 10085311 | 255586 | 3243167
147 | LLQC |  | 194607 | 384060 | 56333 | 349141 | 742768 | 10544147 | 268609 | 3038005
148 | LQC |  | 241664 | 383464 | 83616 | 357352 | 1059734 | 10419596 | 380513 | 3099273
149 | LQC |  | 236095 | 362065 | 79741 | 349882 | 1050095 | 9927272 | 381540 | 2831396
150 | HQC |  | 585758 | 297027 | 361032 | 351880 | 2775350 | 9563625 | 1718877 | 2864953
151 | HQC |  | 625971 | 314586 | 388589 | 366445 | 3019016 | 9455789 | 1716962 | 3021799
152 | HLQC |  | 738301 | 305576 | 547579 | 333348 | 3538035 | 9837947 | 2456610 | 2989201
153 | HLQC |  | 762829 | 297350 | 537078 | 362672 | 3568856 | 9949731 | 2521038 | 3095600
154 | MQC | reinj 1 of #144 above | 451716 | 339696 | 156325 | 354595 | 2157051 | 9473711 | 881644 | 3299293
155 | MQC | reinj 1 of #145 above | 415170 | 304114 | 172459 | 384823 | 2208160 | 10178530 | 840215 | 3223801
156 | LLQC | reinj 1 of #146 above | 205016 | 430172 | 59727 | 379903 | 800814 | 10784187 | 296717 | 3444204
157 | LLQC | reinj 1 of #147 above | 185662 | 412526 | 60658 | 390061 | 814766 | 10514035 | 296404 | 3241601
158 | LQC | reinj 1 of #148 above | 240634 | 363273 | 80693 | 376943 | 1103162 | 10887221 | 401025 | 3253324
159 | LQC | reinj 1 of #149 above | 254901 | 376865 | 87962 | 368510 | 1149712 | 10797149 | 401158 | 3170648
160 | HQC | reinj 1 of #150 above | 571561 | 306532 | 373245 | 329598 | 2887888 | 9868959 | 1672891 | 2952529
161 | HQC | reinj 1 of #151 above | 647709 | 325640 | 396589 | 381646 | 2846794 | 9494004 | 1757665 | 2980296
162 | HLQC | reinj 1 of #152 above | 738905 | 296237 | 537129 | 365059 | 3700536 | 9472870 | 2567406 | 3071109
163 | HLQC | reinj 1 of #153 above | 742047 | 300690 | 553460 | 333381 | 3552982 | 9654569 | 2630476 | 2986052
164 | MQC | reinj 2 of #144 above | 437541 | 311208 | 174959 | 351685 | 2146195 | 10002930 | 843486 | 3028334
165 | MQC | reinj 2 of #145 above | 430769 | 315360 | 162958 | 357592 | 2087912 | 9819841 | 764507 | 3017817
166 | LLQC | reinj 2 of #146 above | 190512 | 411168 | 56992 | 386594 | 737760 | 10539413 | 296807 | 3364755
167 | LLQC | reinj 2 of #147 above | 186026 | 407425 | 54797 | 369997 | 793592 | 10366881 | 276870 | 3145699
168 | LQC | reinj 2 of #148 above | 258015 | 386197 | 75952 | 358545 | 1074163 | 10457075 | 371840 | 3040456
169 | LQC | reinj 2 of #149 above | 257638 | 382374 | 87286 | 359141 | 1111564 | 10097249 | 377121 | 2831090
170 | HQC | reinj 2 of #150 above | 639180 | 319422 | 391086 | 337876 | 2869271 | 9679397 | 1741790 | 2983015
171 | HQC | reinj 2 of #151 above | 632882 | 317900 | 384187 | 345033 | 2957167 | 9759203 | 1801417 | 2936207
172 | HLQC | reinj 2 of #152 above | 770222 | 298914 | 543506 | 362229 | 3651628 | 9672053 | 2508691 | 2864546
173 | HLQC | reinj 2 of #153 above | 752462 | 308913 | 549143 | 343819 | 3609652 | 8964731 | 2501564 | 2923361
174 | hTAG-1 |  | 489287 | 307896 | 194276 | 352249 | 2899401 | 9533561 | 1102006 | 2928251
175 | hTAG-2 |  | 551146 | 307359 | 212449 | 337820 | 3201589 | 9222398 | 1184925 | 2928336
176 | hTAG-3 |  | 531667 | 293721 | 214801 | 358721 | 3015091 | 9325872 | 1107274 | 2902931
177 | hTAG-4 |  | 436522 | 295105 | 178797 | 320728 | 2421031 | 9424959 | 886000 | 2774716
178 | hTAG-5 |  | 516969 | 286590 | 206878 | 348897 | 3088499 | 9203468 | 1152412 | 2922459
179 | hTAG-6 |  | 488107 | 274123 | 204814 | 343407 | 2783164 | 8278806 | 1063344 | 2947674
180 | hTAG-1 | reinj 1 of #174 above | 487761 | 300396 | 195210 | 344368 | 2839706 | 9261549 | 1016374 | 2937011
181 | hTAG-2 | reinj 1 of #175 above | 529149 | 288237 | 207885 | 337779 | 2883444 | 8750912 | 1097688 | 2760975
182 | hTAG-3 | reinj 1 of #176 above | 538284 | 302433 | 193381 | 336212 | 2990589 | 9416346 | 969410 | 2679601
183 | hTAG-4 | reinj 1 of #177 above | 458807 | 309506 | 176206 | 339549 | 2357973 | 9222578 | 857772 | 2816216
184 | hTAG-5 | reinj 1 of #178 above | 555284 | 297684 | 219454 | 354916 | 3236191 | 9666998 | 1135739 | 2869172
185 | hTAG-6 | reinj 1 of #179 above | 533239 | 310729 | 206442 | 358321 | 2999804 | 9455890 | 1024629 | 2916231
186 | hTAG-1 | reinj 2 of #174 above | 482244 | 299064 | 177631 | 336543 | 2738870 | 9264558 | 1002354 | 2708110
187 | hTAG-2 | reinj 2 of #175 above | 572011 | 321040 | 202200 | 352472 | 3114922 | 9605376 | 1065857 | 2874241
188 | hTAG-3 | reinj 2 of #176 above | 529118 | 312616 | 195577 | 326623 | 3025155 | 9686135 | 1035934 | 2667355
189 | hTAG-4 | reinj 2 of #177 above | 428466 | 306396 | 157831 | 324461 | 2533527 | 9297760 | 901269 | 2830675
190 | hTAG-5 | reinj 2 of #178 above | 537447 | 304178 | 198323 | 356627 | 2821720 | 8887549 | 1124023 | 2813086
191 | hTAG-6 | reinj 2 of #179 above | 509521 | 295670 | 191331 | 326321 | 2919982 | 9310229 | 1047173 | 2824466
192 | Pooled QC |  | 388401 | 300887 | 155248 | 325940 | 1915974 | 9265042 | 687980 | 2791300
193 | Pooled QC |  | 416833 | 302305 | 171025 | 317609 | 2027185 | 8405918 | 781748 | 2925033
194 | STD6 rep1 |  | 107092 | 478668 | 40000 | 381046 | 149285 | 10223169 | 155047 | 3318859
195 | STD5 rep1 |  | 108980 | 453039 | 44211 | 364760 | 175106 | 10213456 | 185735 | 2954896
196 | STD4 rep1 |  | 163308 | 454250 | 85430 | 377882 | 344202 | 10327035 | 383620 | 2859165
197 | STD3 rep1 |  | 540709 | 429796 | 399593 | 358146 | 1578172 | 10173599 | 1710730 | 2900546
198 | STD2 rep1 |  | 4541696 | 439416 | 3710950 | 350098 | 13842245 | 9338455 | 15011034 | 2850886
199 | STD1 rep1 |  | 9586001 | 440184 | 7190498 | 380143 | 27996349 | 9464889 | 28493233 | 2746527
200 | STD6 rep2 |  | 106708 | 446221 | 35865 | 355396 | 137834 | 10204947 | 148188 | 3032690
201 | STD5 rep2 |  | 112972 | 447382 | 44023 | 343693 | 178041 | 9960270 | 159631 | 2905798
202 | STD4 rep2 |  | 148930 | 427395 | 80096 | 352074 | 345272 | 10151367 | 338477 | 2925397
203 | STD3 rep2 |  | 553048 | 440539 | 377525 | 351469 | 1566274 | 10616561 | 1749875 | 3131775
204 | STD2 rep2 |  | 4500370 | 413715 | 3572211 | 326080 | 14513356 | 9439707 | 14861746 | 2836059
205 | STD1 rep2 |  | 9697098 | 445014 | 6840343 | 339044 | 27412441 | 9166301 | 29152699 | 2693014
206 | STD6 rep3 |  | 116484 | 515494 | 38501 | 367372 | 144872 | 10542342 | 159262 | 3440082
207 | STD5 rep3 |  | 105905 | 444220 | 48439 | 371383 | 179916 | 10369821 | 186763 | 3064654
208 | STD4 rep3 |  | 158632 | 463465 | 77220 | 349471 | 346677 | 10080590 | 373774 | 2954301
209 | STD3 rep3 |  | 582048 | 454690 | 359197 | 351848 | 1630274 | 10467091 | 1682758 | 3071388
210 | STD2 rep3 |  | 4477437 | 442670 | 3529956 | 358874 | 13542426 | 8829206 | 14537508 | 2801136
211 | STD1 rep3 |  | 8836843 | 434222 | 6782790 | 325012 | 26186426 | 9078214 | 27723440 | 2699043
212 | Matrix blank |  | 453531 | 0 | 156425 | 0 | 1954704 | 0 | 774636 | 0
213 | Total Blank 3 |  | 53098 | 0 | 5066 | 0 | 3640 | 0 | 2015 | 0