## Participant Information
Participant Information
Participant Information |  | Group Name
|  | Lab code | 9
| Institution | Institution name
Street
City
Postal code
Country
| Contact person | Name
Address
Email
Phone

## Summary
Summary
Summary | This protocol describes the extraction and LC-MRM procedure for absolute quantification of Cer d18:1/16:0, Cer d18:1/18:0, Cer d18:1/24:0 and Cer d18:1/24:1 in human plasma according to Kauhanen et al, Anal Bioanal Chem (2016) 408:3475–3483.
A calibration curve is built by injecting several dilutions of the non-labelled standards mixed with a fixed concentration of labelled standards. Each calibration point is prepared and analysed in triplicates (n=3).
To establish the endogenous concentration of the four selected ceramides in the study samples (NIST SRM1950, NIST high TAG, NIST T1D, and NIST young A-A) each sample is extracted and analyzed in six replicates (n =6).
Prepare the pooled QC by mixing 100 mL each of NIST SRM1950, NIST high TAG, NIST T1D, and NIST young AA (total 400 mL).
We suggest proceeding first to measure the ceramides concentration in the four study samples and then continue with the intra-assay validation to avoid wasting material.

## Extraction - Overview
Extraction Overview
Extraction Overview | Table 1 - Extraction - Materials
Chemicals, solvents and materials
Ethyl acetate (EtAC)(LC-MS garde)
2-propanol (IPA)(LC-MS grade)
5% Bovine Serum Albumin (BSA) (Roth T844.2, molecular biology)
Water (LC-MS grade)
Dichloromethane (DCM)(HPLC, >99.8%, amylene as stabilizer)
Methanol (MeOH)(LC-MS grade)
Ammonium acetate (LC-MS grade)(Sigma, 73594-25G-F)
2.0 mL Safe-Lock Eppendorf PPT tubes, pcr clean
Table 2 - Extraction - Sample Types
| Sample type | Components
| Pooled Quality Control samples (QC) | 10 µL pooled QC |  |  | 20 µL Labelled IS |  | 570 µL EtAc:IPA 2:8 (v/v)
| Matrix Blank (MB) | 10 µL Matrix (pooled QC) only |  |  | - |  | 590 µL EtAc:IPA 2:8
| Study Samples (S) | 10 µL Human plasma (NIST SRM1950, high TAG, T1D, young AA) |  |  | 20 µL Labelled IS |  | 570 µL EtAc:IPA 2:8
| Total Blank (TB) | 10 µL 5% BSA |  |  | - |  | 590 µL EtAc:IPA 2:8
| Calibration line standards (STD) | 10 µL 5% BSA + 20 µL non labelled STD mixture (Table 5) |  |  | 20 µL Labelled IS |  | 550 µL EtAc:IPA 2:8
Table 3 - Extraction - Labelled Internal Standard Mixture Solution (provided)
| Labelled IS mixture in EtAc:IPA 2:8 (v/v) | Concentration |  |  |  |  | Unit | Vendor
| D7-Cer d18:1/16:0 | 0.125 |  |  |  |  | pmol/µL [µM] | Avanti Polar Lipids
| D7-Cer d18:1/18:0 | 0.050
| D7-Cer d18:1/24:0 | 1500
| D7-Cer d18:1/24:1 | 0.500
Table 4 - Extraction - Non-Labelled Standards - for STD1 Mixture Solution (provided)
| Non-labelled standards in EtAc:IPA 2:8 (v/v) | Concentration |  |  |  |  | Unit | Vendor
| Cer d18:1/16:0 | 2.0 |  |  |  |  | pmol/µL [µM] | Avanti Polar Lipids
| Cer d18:1/18:0 | 2.0
| Cer d18:1/24:0 | 20.0
| Cer d18:1/24:1 | 20.0
Table 5 - Serial dilutions of Non-Labelled Standards from STD1 Mixture Solution for calibration line
| Add 20 µL of | STD1 Conc. | STD2 Conc. | STD3 Conc. | STD4 Conc. | STD5 Conc. | STD6 Conc. | Unit
| Cer d18:1/16:0 | 2 | 1 | 0.1 | 0.02 | 0.01 | 0.008 | pmol/µL [µM]
| Cer d18:1/18:0 | 2 | 1 | 0.1 | 0.02 | 0.01 | 0.008
| Cer d18:1/24:0 | 20 | 10 | 1 | 0.2 | 0.1 | 0.08
| Cer d18:1/24:1 | 20 | 10 | 1 | 0.2 | 0.1 | 0.08

## Extraction - Procedure
Extraction Procedure
Extraction Procedure | Extraction Procedure
| Step | Description | Reference Table
| 1 | Prepare different dilutions of STD1 by diluting the provided mixture (Extraction - Overview, Table 4) with EtAc:IPA 2:8 as suggested in table | Extraction - Overview'!B34
| 2 | Thaw plasma on ice (+4°C) and bring to room temperature prior to extraction. Mix by vortexing before use.
| 3 | Prepare samples in 2 mL Eppendorf tubes manually according to specifications in column C of table | Extraction - Overview'!C14
| 4 | Add 20 µL of labelled internal standard mixture to samples according to column F of table | Extraction - Overview'!F14
| 5 | For the calibration line, add 20 µL of different dilutions of the non-labelled mixture (STD1-6).
| 6 | Add 590 µL of ethyl acetate:isopropanol (2:8, vol/vol) to Total Blank (TB) and Matrix Blank (MB); 550 µL to calibration line (STD); 570 µL to study samples and QCs according to column H of table | Extraction - Overview'!H14
| 7 | Mix samples 10 min by pipetting or vortexing
| 8 | Centrifuge samples for 10 min at 3000 x gat room temperature
| 9 | Transfer 25 µL of the clear supernatant to MS-vials with 250 µl inserts
| 10 | Dry extracts under nitrogen
| 11 | Resuspend with 250 µl DCM/MeOH (1:1, v/v), 10 mM ammonium acetate

## LC-MRM - Overview
Flow Injection Analysis-MRM Overview
LC-MRM Overview | Table 6 Chemicals and solvents
Dichloromethane, HPLC, >99.8%, amylene as stabilizer
Methanol LC-MS grade (MeOH)
Ammonium acetate ()
Table 7 FIA-MRM FIA-MS Details
| Property | Specifications
| Mass spectrometry | QTRAP (5500 Sciex)
| LC | UHPLC (Shimadzu Nexera) without any column
| Injection volume (indicative) | 50 µl
| Weak wash | DCM/MeOH (1:1, v/v)
Strong wash
| Pre-column | None
| Column | None
| Column temperature | None
| Mobile phase A | DCM/MeOH (1:1, v/v), ammonium acetate 10 mM
Table 8 FIA-MRM gradient
| Time (minutes) | Flow (µL/min)
| 0.01 | 7
| 2.4 | 7
| 2.5 | 30
| 7.5 | 30
| 7.6 | 7
Table 9 FIA-MRM MS conditions
| Polarity mode | Positive
| Ion spray (volt) | 4100
| Curtain gas | 17
| Source temperature (°C) | 200
| Gas 1 | 17
| Gas 2 | 25
| DP (volt) | 80
| EP (volt) | 10
| Collision cell exit potential (volt) | 16
| Collision energy (CE) (volt) | 43
| Acquisition time (min) | 0.2
Table 10 FIA-MRM Mass transitions used for absolute quantification
| Analyte | Mass transitions | IS | Mass transitions | RT (min) | CE
| Cer d18:1/16:0 | 538.6 – 264.4 | D7- Cer d18:1/16:0 | 545.6 - 271.3 | None | 43
C34H68NO3 – C18H34N
| Cer d18:1/18:0 | 566.7 – 264.4 | D7- Cer d18:1/18:0 | 573.6 – 271.3 | None | 43
C36H72NO3 – C18H34N
| Cer d18:1/22:0 | 622.7 – 264.4 |  |  | None | 43
C40H80NO3 – C18H34N
| Cer d18:1/24:0 | 650.8 – 264.4 | D7- Cer d18:1/24:0 | 657.7 – 271.3 | None | 43
C42H84NO3 – C18H34N
| Cer d18:1/24:1 | 648.8 – 264.4 | D7- Cer d18:1/24:1 | 655.7 – 271.3 | None | 43
C42H82NO3 – C18H34N
Table 11 Data processing
| Processing software | Multiquant 3.0.3 (Summation integration algorithm, noise 0%)

## FIA-MRM - Sample Sequence
Sample Sequence
Sample Sequence (Example) | Sample Sequence 1 - Sample Assay
| Number | Sample Name | Remarks
| 1 | TotalBlank-01
| 2 | MatrixBlank-01
| 3 | Line1_STD1_01
| 4 | Line1_STD1_02
| 5 | Line1_STD1_03
| 6 | Line1_STD2_01
| 7 | Line1_STD2_02
| 8 | Line1_STD2_03
| 9 | Line1_STD3_01
| 10 | Line1_STD3_02
| 11 | Line1_STD3_03
| 12 | Line1_STD4_01
| 13 | Line1_STD4_02
| 14 | Line1_STD4_03
| 15 | Line1_STD5_01
| 16 | Line1_STD5_02
| 17 | Line1_STD5_03
| 18 | Line1_STD6_01
| 19 | Line1_STD6_02
| 20 | Line1_STD6_03
| 21 | PooledQC-01
| 22 | PooledQC-02
| 23 | SRM1950-01_01
| 24 | SRM1950-01_02
| 25 | SRM1950-01_03
| 26 | SRM1950-02_01
| 27 | SRM1950-02_02
| 28 | SRM1950-02_03
| 29 | SRM1950-03_01
| 30 | SRM1950-03_02
| 31 | SRM1950-03_03
| 32 | SRM1950-04_01
| 33 | SRM1950-04_02
| 34 | SRM1950-04_03
| 35 | SRM1950-05_01
| 36 | SRM1950-05_02
| 37 | SRM1950-05_03
| 38 | SRM1950-06_01
| 39 | SRM1950-06_02
| 40 | SRM1950-06_03
| 41 | PooledQC-03
| 42 | PooledQC-04
| 43 | T1D-01_01
| 44 | T1D-01_02
| 45 | T1D-01_03
| 46 | T1D-02_01
| 47 | T1D-02_02
| 48 | T1D-02_03
| 49 | T1D-03_01
| 50 | T1D-03_02
| 51 | T1D-03_03
| 52 | T1D-04_01
| 53 | T1D-04_02
| 54 | T1D-04_03
| 55 | T1D-05_01
| 56 | T1D-05_02
| 57 | T1D-05_03
| 58 | T1D-06_01
| 59 | T1D-06_02
| 60 | T1D-06_03
| 61 | TotalBlank-02
| 62 | PooledQC-05
| 63 | PooledQC-06
| 64 | YoungAA-01_01
| 65 | YoungAA-01_02
| 66 | YoungAA-01_03
| 67 | YoungAA-02_01
| 68 | YoungAA-02_02
| 69 | YoungAA-02_03
| 70 | YoungAA-03_01
| 71 | YoungAA-03_02
| 72 | YoungAA-03_03
| 73 | YoungAA-04_01
| 74 | YoungAA-04_02
| 75 | YoungAA-04_03
| 76 | YoungAA-05_01
| 77 | YoungAA-05_02
| 78 | YoungAA-05_03
| 79 | YoungAA-06_01
| 80 | YoungAA-06_02
| 81 | YoungAA-06_03
| 82 | Pooled_QC-07
| 83 | Pooled_QC-08
| 84 | hTAG-01_01
| 85 | hTAG-01_02
| 86 | hTAG-01_03
| 87 | hTAG-02_01
| 88 | hTAG-02_02
| 89 | hTAG-02_03
| 90 | hTAG-03_01
| 91 | hTAG-03_02
| 92 | hTAG-03_03
| 93 | hTAG-04_01
| 94 | hTAG-04_02
| 95 | hTAG-04_03
| 96 | hTAG-05_01
| 97 | hTAG-05_02
| 98 | hTAG-05_03
| 99 | hTAG-06_01
| 100 | hTAG-06_02
| 101 | hTAG-06_03
| 102 | PooledQC-09 | Correspond to second injection of PooledQC-01 (prepared only 8 pooledQC)
| 103 | PooledQC-010 | Correspond to second injection of PooledQC-02 (prepared only 8 poolerQC)
| 104 | Line2_STD6_01
| 105 | Line2_STD6_02
| 106 | Line2_STD6_03
| 107 | Line2_STD5_01
| 108 | Line2_STD5_02
| 109 | Line2_STD5_03
| 110 | Line2_STD4_01
| 111 | Line2_STD4_02
| 112 | Line2_STD4_03
| 113 | Line2_STD3_01
| 114 | Line2_STD3_02
| 115 | Line2_STD3_03
| 116 | Line2_STD2_01
| 117 | Line2_STD2_02
| 118 | Line2_STD2_03
| 119 | Line2_STD1_01
| 120 | Line2_STD1_02
| 121 | Line2_STD1_03
| 122 | MatrixBlank_02
| 123 | TotalBlank_03

## FIA-MRM - Intra Assay QC Proc
Intra Assay Procedure
Intra Assay Procedure | Quality Controls for Intra-Assay Variation
| Step | Description | Reference Table
| 1 | Prepare the Pooled NIST QC sample containing equal volumes of the four NIST reference plasma you received.
| 2 | The pooled QC representing middle quality control (MQC), low QC (LQC), and lower limit of quantification QC (LLOQ) is prepared by diluting MQC (e.g. two- and three-fold) with water.
| 3 | High QC (HQC) and upper limit of quantification (ULOQ) QC are prepared by spiking in 10 µL and 20 µL respectively of STD3 endogenous standards from table | Extraction - Overview'!B34
| 4 | Internal standards are added to all QC samples.
Intra-Assay Variation
| 6 | The precision and accuracy of the assay is determined for each of the four ceramides measured in LLOQ, LQC, MQC, HQC, and HLOQ.
| 7 | Intra-assay variation, precision, and accuracy is calculated for each of the ceramide in replicates of six (n = 6) at each QC concentration independently on three separate occasions.
| 8 | The intra-assay precision (percentage coefficient variance, %CV) and accuracy (percentage accuracy, %Accuracy) are calculated from the nominal concentrations according to the formulas reported in corresponding sheets and tables of this document.
| 9 | Recovery estimation is not necessary for the published method but only for different ones. Stability estimation is not necessary for the published method but only for different ones. Calibration line is prepared in 5% BSA.

## IntraAssay - Sample Sequence
Sample Sequence
Sample Sequence (Example) | Sample Sequence 2 - Intra Assay
| Number | Sample Name | Remarks
| 1 | TotalBlank-01
| 2 | MatrixBlank-01
| 3 | LLQC-01_001
| 4 | LLQC-01_002
| 5 | LLQC-01_003
| 6 | LLQC-02_001
| 7 | LLQC-02_002
| 8 | LLQC-02_003
| 9 | LLQC-03_001
| 10 | LLQC-03_002
| 11 | LLQC-03_003
| 12 | LLQC-04_001
| 13 | LLQC-04_002
| 14 | LLQC-04_003
| 15 | LLQC-05_001
| 16 | LLQC-05_002
| 17 | LLQC-05_003
| 18 | LLQC-06_001
| 19 | LLQC-06_002
| 20 | LLQC-06_003
| 21 | LQC-01_001
| 22 | LQC-01_002
| 23 | LQC-01_003
| 24 | LQC-02_001
| 25 | LQC-02_002
| 26 | LQC-02_003
| 27 | LQC-03_001
| 28 | LQC-03_002
| 29 | LQC-03_003
| 30 | LQC-04_001
| 31 | LQC-04_002
| 32 | LQC-04_003
| 33 | LQC-05_001
| 34 | LQC-05_002
| 35 | LQC-05_003
| 36 | LQC-06_001
| 37 | LQC-06_002
| 38 | LQC-06_003
| 39 | MQC-01_001
| 40 | MQC-01_002
| 41 | MQC-01_003
| 42 | MQC-02_001
| 43 | MQC-02_002
| 44 | MQC-02_003
| 45 | MQC-03_001
| 46 | MQC-03_002
| 47 | MQC-03_003
| 48 | MQC-04_001
| 49 | MQC-04_002
| 50 | MQC-04_003
| 51 | MQC-05_001
| 52 | MQC-05_002
| 53 | MQC-05_003
| 54 | MQC-06_001
| 55 | MQC-06_002
| 56 | MQC-06_003
| 57 | TotalBlank-02
| 58 | HQC-01_001
| 59 | HQC-01_002
| 60 | HQC-01_003
| 61 | HQC-02_001
| 62 | HQC-02_002
| 63 | HQC-02_003
| 64 | HQC-03_001
| 65 | HQC-03_002
| 66 | HQC-03_003
| 67 | HQC-04_001
| 68 | HQC-04_002
| 69 | HQC-04_003
| 70 | HQC-05_001
| 71 | HQC-05_002
| 72 | HQC-05_003
| 73 | HQC-06_001
| 74 | HQC-06_002
| 75 | HQC-06_003
| 76 | HLQC-01_001
| 77 | HLQC-01_002
| 78 | HLQC-01_003
| 79 | HLQC-02_001
| 80 | HLQC-02_002
| 81 | HLQC-02_003
| 82 | HLQC-03_001
| 83 | HLQC-03_002
| 84 | HLQC-03_003
| 85 | HLQC-04_001
| 86 | HLQC-04_002
| 87 | HLQC-04_003
| 88 | HLQC-05_001
| 89 | HLQC-05_002
| 90 | HLQC-05_003
| 91 | HLQC-06_001
| 92 | HLQC-06_002
| 93 | HLQC-06_003
| 94 | TotalBlank-03
| 95 | MatrixBlank-03

## LC-MRM - Intra Assay QC Res
INTRA ASSAY - QC RESULTS
INTRA ASSAY - QC RESULTS | Table 12 LC-MRM Calibration Line 1 Data (Beginning). Please report the values of each of the three replicates separately for each compound and standard.
| Calibration Line 1 | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| STD 1 | 49915.65 | 50570.3 | 65938.425 | 43152.1 | 44283.775 | 60417.025 | 556399.2 | 595789.85 | 815486.1 | 497075.95 | 547468.775 | 739441.75 | 2278.025 | 3077.55 | 3926.575 | 2008.775 | 2165.2 | 3377.7 | 47217.65 | 49290.55 | 71659.15 | 15777.6 | 16733.975 | 24546.6
| STD 2 | 23234.6 | 23496.2 | 26418.875 | 22521.8 | 23441.825 | 25075.825 | 282234.95 | 301585.35 | 338087.85 | 261176.325 | 274881.45 | 304703.975 | 3414.675 | 2124.1 | 2560.275 | 2313.9 | 2013.775 | 2645.575 | 49801.65 | 56339.025 | 59030.525 | 18062.525 | 16186.325 | 19365.325
| STD 3 | 2822.2 | 2403.85 | 3145.825 | 2017.625 | 2367.825 | 2521.925 | 29210.85 | 31166.825 | 32924.05 | 27367.05 | 27493.675 | 29887.875 | 2581.575 | 2279.05 | 2171.45 | 2028.9 | 2176.525 | 2094.15 | 48729.775 | 49848.2 | 51326.95 | 15581.925 | 17436.575 | 17032.025
| STD 4 | 432.85 | 540.525 | 749.25 | 251.5 | 720.5 | 120.1 | 6902.85 | 5900.475 | 6652.525 | 5836.2 | 5152.925 | 5212.8 | 2436.575 | 2604.15 | 2485.2 | 2381.775 | 2135.3 | 2174.075 | 44605.15 | 44610.15 | 43477.65 | 15039.65 | 15150.55 | 15132.325
| STD 5 | 304.025 | 480.4 | 300.225 | 362.8 | 465.25 | 220.15 | 2813.65 | 2979.85 | 3090.2 | 2823.075 | 2307.575 | 2914.45 | 2006.225 | 2101.6 | 2629.475 | 1897.675 | 2879.75 | 1893.875 | 42709.775 | 43222.5 | 45566.55 | 13552.5 | 13453.275 | 14800.825
| STD 6 | 80.05 | 260.1 | 254.025 | 222.65 | 140.175 | 254.025 | 2830.825 | 2223.025 | 2733.525 | 2020.25 | 2269.35 | 2207.525 | 2047.575 | 2723.375 | 2256.55 | 1535.875 | 1856.325 | 2135.1 | 40856.025 | 40543.8 | 40224.825 | 13373.975 | 14127.775 | 14319.35
Table 13 LC-MRM Calibration Line 2 Data (End). Please report the values of each of the three replicates separately for each compound and standard.
| Calibration Line 2 | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| STD 1 | 43825.4 | 48311.5 | 55717.25 | 42488.475 | 43733.425 | 50988.55 | 622880.875 | 620589.9 | 711007.675 | 546000.45 | 567257.325 | 632902.1 | 2650.625 | 2849.525 | 2752.05 | 1841.35 | 2737.9 | 2594.15 | 52155.95 | 50666.05 | 62372.725 | 16976.1 | 17111.35 | 20238.7
| STD 2 | 23787.375 | 23994.075 | 24565.95 | 21273.4 | 22169.1 | 20697.75 | 297065.225 | 292102.925 | 306141.55 | 276183.2 | 261785.3 | 282214.575 | 1992.475 | 2185.35 | 2735.675 | 2584 | 2397.8 | 2593.65 | 53154.25 | 52755.3 | 53843.3 | 17336.275 | 16667.175 | 17802.45
| STD 3 | 2565.35 | 2716.7 | 2517.725 | 2082.65 | 2296.4 | 2898.175 | 32358.325 | 30317.6 | 34947.8 | 31356 | 29176.45 | 29954.4 | 2663.45 | 3127.475 | 3245.95 | 2634.725 | 2037.45 | 1989.05 | 57317.425 | 58235.15 | 55794.225 | 19149.025 | 18444.5 | 19189.35
| STD 4 | 640.35 | 858.075 | 900.675 | 663.05 | 402.875 | 462.775 | 6353.4 | 4411.8 | 6360.5 | 5283.95 | 4230.65 | 6609.05 | 2797.925 | 1723.925 | 2579.6 | 2080.125 | 1872.45 | 2635.425 | 48523.225 | 38381.8 | 54984.925 | 15662.525 | 12464.925 | 17065.725
| STD 5 | 290.2 | 511.65 | 282.65 | 460.45 | 280.3 | 231.425 | 2872.95 | 3134.85 | 2708.4 | 2938.525 | 2734.425 | 2481.875 | 2060.175 | 1660.85 | 2270.325 | 2040.125 | 1504.9 | 1912.4 | 46729.1 | 43190.625 | 42376.4 | 15922.325 | 14889.025 | 14966.25
| STD 6 | 82.55 | 525.325 | 470.5 | 100.075 | 200.2 | 302.775 | 2480.425 | 2527.9 | 1913.725 | 2311.575 | 2403.95 | 1977.6 | 1983.75 | 2209.225 | 2410.475 | 1583.825 | 1651.2 | 1272.175 | 37552.5 | 43160.375 | 42369.55 | 13825.95 | 13943.975 | 13349.9
Table 14 FIA-MRM Sample Data. Please report the values of each of the three replicates separately for each compound and standard.
| NIST SRM | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| SRM 1 | 3213.825 | 3632.5 | 3212.275 | 998.3 | 1036.9 | 836.875 | 28865.475 | 26949.9 | 27356.7 | 9770.925 | 11167.025 | 10140.85 | 2633 | 3109.925 | 2370.4 | 3004.725 | 2570.55 | 1957.475 | 52858.475 | 54372.075 | 46812.275 | 16865.5 | 16581.1 | 15558.9
| SRM 2 | 2945.75 | 2227.925 | 2434.25 | 1065.7 | 901.9 | 900.575 | 27751.75 | 28110.075 | 29754.325 | 9397.05 | 9598.125 | 10686.05 | 2061.5 | 2219.45 | 2136.25 | 1331 | 2168.85 | 1766.2 | 36677.2 | 39662.05 | 38205.95 | 11956 | 13002.175 | 14179.525
| SRM 3 | 2264.275 | 2768.25 | 3137.4 | 1096.95 | 631.75 | 1080.55 | 30579.7 | 30212.875 | 31201.55 | 11100.6 | 10090.1 | 9925.325 | 2605.475 | 2565.425 | 1629.975 | 1965.25 | 1837.7 | 1880.575 | 39474.7 | 40208.975 | 38987.625 | 13829.675 | 13268.175 | 13374.9
| SRM 4 | 2265.15 | 2554.15 | 2401.8 | 1043.325 | 865.425 | 711.775 | 28413.6 | 26767.425 | 30062.925 | 8631.675 | 9546.175 | 9624.975 | 1902.7 | 2188.825 | 2243.85 | 2037.8 | 1612.7 | 1775.075 | 39581.95 | 39818.825 | 40330.825 | 13867.275 | 13293.575 | 13268.575
| SRM 5 | 2881.575 | 2578.5 | 2504.15 | 844.5 | 770.475 | 671.55 | 31473.625 | 27135.725 | 31826.025 | 10838.9 | 9774.075 | 11060.2 | 2394.225 | 2625.575 | 2320.3 | 1877.45 | 1353.65 | 1723.975 | 40334.9 | 38986.9 | 40396.725 | 14761.1 | 13249.85 | 13809.025
| SRM 6 | 2695.65 | 2770.5 | 2693.075 | 863.3 | 795.475 | 981.925 | 27222.225 | 28865.425 | 27613.9 | 9145.45 | 11599.375 | 9716.175 | 1523.4 | 1670.925 | 2059.025 | 1833.875 | 1835.2 | 1683.6 | 36838.025 | 40323.175 | 36921.4 | 12670.95 | 14606.225 | 13183.975
| NIST hTAG | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| hTAG 1 | 2711.975 | 2316.775 | 2481.975 | 741.775 | 910.4 | 634.25 | 34109.45 | 36400.175 | 33950.8 | 11560.8 | 11121.025 | 10760.175 | 2179 | 1605.075 | 1926.5 | 1827.325 | 1293.45 | 1621.025 | 32095.325 | 32169.9 | 34240.4 | 11767.225 | 10592.4 | 10316.95
| hTAG 2 | 2101.575 | 2623.1 | 2057.725 | 728 | 720.475 | 525.25 | 32158.725 | 35706.475 | 34206.525 | 9674.675 | 11700.05 | 11803.05 | 1963.9 | 1507.15 | 1469.825 | 1127.975 | 2282.975 | 1635.225 | 31449.1 | 31559.525 | 30302.6 | 11109.2 | 11227.85 | 9939.15
| hTAG 3 | 2061.425 | 2087.725 | 2863.25 | 590.55 | 1205.75 | 834.525 | 34865.575 | 33255.925 | 35773.025 | 11639.575 | 11899.15 | 11640.85 | 1887.175 | 1173.35 | 1512.05 | 1347.075 | 1384.625 | 1617.525 | 31738.775 | 30789.875 | 29751.175 | 10148.275 | 10101.875 | 10653.6
| hTAG 4 | 3016.1 | 2176.575 | 2616.75 | 923.2 | 624.1 | 662.95 | 35396.2 | 37223.7 | 37882.575 | 10680.175 | 12586.25 | 9950.45 | 1280.95 | 1654.775 | 2130.05 | 1271.075 | 1308.675 | 1309.625 | 31812.075 | 31590.075 | 31544.05 | 11344.375 | 10384.4 | 10296.025
| hTAG 5 | 2501.775 | 2551.825 | 2480.3 | 734.275 | 580.4 | 865.875 | 36139.875 | 37306.725 | 38554.075 | 13452.525 | 10664.075 | 12134.375 | 2036.5 | 1718.375 | 2180.3 | 1102.025 | 1825.25 | 1513.7 | 34515.825 | 33591.475 | 35478.175 | 12244.025 | 10516.1 | 12515.25
| hTAG 6 | 3152.175 | 2180.1 | 2007.825 | 1174.3 | 555.55 | 740.625 | 42000.425 | 36089.625 | 41538 | 13295.1 | 10426 | 13724.825 | 2412.925 | 1514.95 | 1608.675 | 1618.7 | 1932.875 | 1899.9 | 38491.45 | 32198.75 | 35348.9 | 12087.875 | 11126.475 | 12057.675
| NIST T1D | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| T1D 1 | 2056.45 | 2032.575 | 2484.025 | 1105.65 | 1055.825 | 852 | 28655 | 29017.225 | 31473.475 | 10844.45 | 9627 | 10243.475 | 2111.25 | 2574.15 | 2379.15 | 2231.7 | 1645.05 | 1995.175 | 40621.825 | 37306.125 | 42906.925 | 13636.1 | 13384.95 | 13710.825
| T1D 2 | 2790.775 | 2609.825 | 2332.375 | 1012.15 | 1360.85 | 670.5 | 27547.85 | 29982.625 | 32735.4 | 10082.15 | 10571.625 | 11337.025 | 2050.4 | 2464.2 | 2307.7 | 1801.55 | 1658.85 | 1976.425 | 39331.875 | 39706.525 | 43829.825 | 13162.7 | 13974.525 | 14180.925
| T1D 3 | 2327 | 3067.05 | 2574.075 | 1190.625 | 941.95 | 1274.4 | 35299.975 | 31477.9 | 33083.175 | 12459.425 | 11327.775 | 12707.125 | 2704.3 | 2397.875 | 2526.85 | 2328.9 | 2028.675 | 1973.8 | 49070.6 | 45420.975 | 47562.025 | 16549.875 | 16860.325 | 16254.65
| T1D 4 | 2598.175 | 3253.575 | 3003.125 | 667.975 | 953.075 | 856.8 | 33649.225 | 34557.325 | 35737.25 | 11080.225 | 11698.375 | 12497.6 | 2571.675 | 2420.4 | 2094.3 | 2520.4 | 1741.575 | 2495.4 | 42310.9 | 46041.525 | 46868.35 | 14345.925 | 15076.05 | 13553.375
| T1D 5 | 3472.225 | 2534.45 | 2174.2 | 1339.6 | 1221.95 | 1258.6 | 37806.45 | 31799.075 | 32459.225 | 12391.325 | 11543.5 | 12064.625 | 2954 | 2272.825 | 2390.65 | 2046.25 | 1846.35 | 2315.325 | 48345.625 | 45685.775 | 43625.725 | 16509.925 | 16110.15 | 16435.775
| T1D 6 | 2713.125 | 2273.75 | 2674.425 | 900.375 | 1010.8 | 1091.875 | 31298.65 | 33468.15 | 33460.625 | 11400.25 | 10943.725 | 11390.525 | 2544.15 | 2086.55 | 2369.15 | 1567.25 | 1938.775 | 2438.075 | 42908.15 | 45594.825 | 45897.525 | 15043.675 | 14091.425 | 14865.475
| NIST Young AA | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| yA-A 1 | 2016.625 | 2392.975 | 2359.35 | 501.5 | 700.6 | 491.425 | 24499.475 | 23839.475 | 24969.65 | 8222.75 | 8386.825 | 8474.9 | 3121.875 | 2739.65 | 2839.6 | 2025.15 | 2332.675 | 2339.6 | 49255.075 | 52509.425 | 55496.825 | 18113.225 | 18354.1 | 16331.725
| yA-A 2 | 2475.325 | 2435.775 | 2077.575 | 603.05 | 440.5 | 641.9 | 25731.4 | 25123.65 | 24586.4 | 7180.3 | 6537.275 | 8025.95 | 3034.525 | 2124.975 | 2428.225 | 2474.25 | 2440.4 | 1872.35 | 53771.2 | 50728.375 | 48105.525 | 18994.25 | 16566.05 | 17136.125
| yA-A 3 | 1940.175 | 1609.65 | 2208.95 | 918.3 | 860.775 | 600.35 | 25760.9 | 25970.725 | 23269.4 | 8698.65 | 6986.15 | 7838.4 | 2723.35 | 2301.35 | 1815.025 | 1867.425 | 1574.75 | 2522.05 | 49133.45 | 47873.275 | 48080.625 | 16076.4 | 15299.125 | 15586.5
| yA-A 4 | 2062.575 | 1872.425 | 1902.6 | 578.05 | 391.5 | 580.425 | 21455.325 | 24169.4 | 21671 | 7168.2 | 7940.125 | 7273.9 | 2406.6 | 2106.625 | 2140.325 | 1974.875 | 2205.45 | 2524.6 | 42639.05 | 47400.425 | 45811.225 | 14857 | 14714 | 15340.7
| yA-A 5 | 1820.05 | 1837.925 | 1460.85 | 492.925 | 544.275 | 525.375 | 23087.775 | 24034.15 | 21842.9 | 7259.025 | 6902.325 | 7189.275 | 2566.9 | 2228.95 | 2252.55 | 2017.625 | 1723.95 | 2436.725 | 44240.975 | 45469.9 | 45302.4 | 16146.7 | 14725.775 | 15772.725
| yA-A 6 | 2212.025 | 2006.5 | 2032.275 | 1014.525 | 525.3 | 532.875 | 21955.075 | 21001.425 | 23867.75 | 7847.875 | 7450.1 | 7197.625 | 2421.675 | 1906.075 | 2315.375 | 2512.6 | 2257.5 | 2329.025 | 42727.75 | 46270.55 | 45877.425 | 18032.4 | 16156.525 | 15992.55
Table 15 FIA-MRM Intra Assay QC Data. Please report the values of each of the three replicates separately for each compound and standard.
| Lowest Level QC | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| LLQC 1 | 1133.075 | 1554.775 | 1515.925 | 640.4 | 600.5 | 444.15 | 15815.225 | 17147.675 | 18687.2 | 5314.1 | 6261.75 | 6409.275 | 2990.625 | 2012.625 | 2470.575 | 1957.7 | 1829.75 | 1997.575 | 41931.475 | 47127.7 | 44104.6 | 15478.025 | 14306.725 | 17187.85
| LLQC 2 | 1463.725 | 1130.975 | 1616.475 | 390.3 | 812.825 | 482.875 | 16191.225 | 15865.8 | 19379.875 | 6540.8 | 5741.925 | 5721.35 | 2327.925 | 1951.325 | 2297.7 | 1916.1 | 1945.225 | 1797.375 | 40668.225 | 38833.45 | 43397.425 | 14170.3 | 14036.75 | 15444.725
| LLQC 3 | 1538.875 | 1660.175 | 1435.975 | 640.375 | 482.8 | 577.95 | 16902.625 | 17813.825 | 15789.25 | 5179.025 | 5861.825 | 6194.625 | 2220.3 | 2457.775 | 2650.775 | 2016.625 | 1677.425 | 1867.425 | 44196.7 | 47277.3 | 41880.1 | 13910.875 | 17224.35 | 15854.225
| LLQC 4 | 1391 | 1555.775 | 1578.6 | 480.375 | 520.35 | 463.05 | 16719.225 | 14422.775 | 16754.825 | 7489.125 | 5052.9 | 6512.625 | 2298.975 | 2528.175 | 2449.4 | 2215.4 | 1916.3 | 1763.625 | 44522.45 | 42855.475 | 44375.875 | 14774.075 | 13486.525 | 16272.45
| LLQC 5 | 1752.775 | 1627.225 | 1313.175 | 725.5 | 420.3 | 580.325 | 18401.95 | 18120.85 | 16851.55 | 6805.125 | 5636.325 | 7486.05 | 3444.4 | 2858.075 | 2511.6 | 1972.55 | 2372.925 | 2034.125 | 47889.45 | 47743.775 | 47605.375 | 15537.45 | 14938.575 | 17910.5
| LLQC 6 | 1163.525 | 1406.125 | 1819.95 | 320.125 | 474 | 485.225 | 17249.15 | 17558.025 | 18935.175 | 5871.475 | 6024.325 | 5480.65 | 2656.75 | 2719.65 | 2666.075 | 2040.075 | 1948.8 | 2040.3 | 46341.975 | 45411.725 | 49618.1 | 15643.575 | 16485.925 | 18746.975
| Low QC | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| LQC 1 | 1416.175 | 1336.15 | 1160.725 | 630.675 | 373.975 | 589.05 | 15358.75 | 12647.875 | 12491.8 | 4967.925 | 4380.55 | 4319.375 | 1976.325 | 1715.075 | 2041.6 | 1903.8 | 1453.7 | 1379.95 | 38270.75 | 32655.95 | 31353.425 | 14163.05 | 11691.625 | 11526.425
| LQC 2 | 1043.125 | 988.125 | 785.55 | 484.075 | 565.425 | 531.575 | 13145.3 | 11965.3 | 12885.05 | 4879.775 | 3792.825 | 4673.3 | 1656.05 | 1859.825 | 2122.625 | 1577.225 | 1339.875 | 1909 | 36358.9 | 33887.25 | 33340.75 | 12254 | 11464.15 | 10308.025
| LQC 3 | 1045.75 | 1338.15 | 1274.775 | 420.4 | 471.5 | 440.3 | 14870.25 | 15349.45 | 13879.225 | 5270.575 | 4140.325 | 3827.525 | 1534.8 | 2066.625 | 1394.775 | 1787.5 | 2057.625 | 2133.825 | 40069.2 | 38128.6 | 36281.275 | 12942.8 | 11763.575 | 11939.7
| LQC 4 | 1515 | 750.675 | 1187.175 | 584.3 | 472.7 | 620.275 | 13620.375 | 13034.825 | 12354.65 | 4573.275 | 4813.6 | 4767.05 | 1743.8 | 2065.225 | 1791.025 | 1666.1 | 1452.275 | 1513.575 | 34378.9 | 34183.225 | 33883.125 | 12545.025 | 10633.725 | 9771.275
| LQC 5 | 1365.85 | 1506.125 | 1013.25 | 220.05 | 390.275 | 555.375 | 13692.85 | 13554.275 | 14788.55 | 4874.225 | 4728.05 | 4930.475 | 2080.375 | 1326.85 | 1897.425 | 1380.9 | 1467.15 | 1674.925 | 34882.575 | 34830.55 | 36894.675 | 11187.25 | 12147.85 | 12549.6
| LQC 6 | 1449.8 | 1549.725 | 1364.8 | 60.05 | 474.075 | 430.4 | 15643.075 | 15584.625 | 16618.45 | 5021.2 | 5619.05 | 5446.1 | 1915.375 | 1644.95 | 2098.825 | 2041.425 | 1130.9 | 2070.325 | 39755.6 | 40312.325 | 42869.525 | 14693.275 | 12809.8 | 15236.8
| Middle QC | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| MQC 1 | 2918.375 | 2987.25 | 2911.925 | 1058.225 | 980.7 | 1290.575 | 26285.7 | 31575.5 | 29287.85 | 9768.025 | 10389.925 | 10530.575 | 2451.425 | 2023.9 | 2316.175 | 2172.85 | 1865.175 | 1853.85 | 36733.175 | 39988.275 | 38587.7 | 12837.175 | 13866.7 | 13051.6
| MQC 2 | 2689.25 | 1735.1 | 2100.175 | 734.3 | 815.5 | 856.975 | 26769.575 | 27442.05 | 26062.575 | 8582.925 | 8387.525 | 8409.2 | 1450.85 | 1469.975 | 2037.7 | 1898.625 | 1651.075 | 1535.975 | 34367.25 | 31208.3 | 31283.1 | 12359.6 | 11138.5 | 12376.7
| MQC 3 | 1892.775 | 1972.55 | 1752.4 | 540.35 | 921.825 | 1046.9 | 22455.375 | 23885.825 | 22711.725 | 7114.925 | 7911.125 | 7963 | 1225.75 | 1949.05 | 1300.85 | 1695.8 | 1697.5 | 1388.325 | 28921.375 | 32339.325 | 33021.95 | 10577.55 | 10815.725 | 11806.775
| MQC 4 | 1647.45 | 2015.225 | 1922.675 | 500.375 | 800.475 | 831.775 | 21297.175 | 28107.775 | 23199.3 | 8509.275 | 9174.35 | 8626.175 | 1314.375 | 1444.525 | 1876.15 | 1558.7 | 1757.45 | 1307.15 | 28381.675 | 31615.875 | 28015.45 | 10606.875 | 11488.625 | 10528.15
| MQC 5 | 2341.475 | 2339.075 | 1985.125 | 903.325 | 723.075 | 662.975 | 25149.675 | 27385.9 | 23638.025 | 8791.375 | 8415.75 | 7945.475 | 1836.15 | 1842.65 | 1718.6 | 1372.25 | 1033.175 | 2016.4 | 34286.925 | 37188.475 | 28888.3 | 12414.525 | 10641.7 | 9899.75
| MQC 6 | 2004.95 | 1464.825 | 1786.45 | 638 | 630.55 | 570.35 | 22757.25 | 20045.15 | 19773.3 | 7202.375 | 7566.825 | 7246.45 | 1356.1 | 1416.15 | 2324.175 | 1601.1 | 1462.35 | 1592.725 | 30967.975 | 25393.975 | 27249.525 | 10889.225 | 9365.8 | 9219.975
| High QC | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| HQC 1 | 2429.35 | 2400.375 | 2379.425 | 1254.575 | 1257.075 | 1374.9 | 31793.075 | 31040.075 | 27437.35 | 14960.15 | 15128.35 | 15484.95 | 1733.875 | 1494.7 | 1495.975 | 1412 | 1549.775 | 1639.675 | 33735.425 | 33198.85 | 34007.55 | 11545.75 | 12673.65 | 12258.25
HQC 2
| HQC 3 | 3094.4 | 2479.325 | 3061.15 | 1576.4 | 1663.5 | 1694.6 | 34342.475 | 35692.525 | 35866.5 | 16504.2 | 18462.4 | 17474.725 | 1939.9 | 1956.45 | 1765.15 | 1497.55 | 1717.425 | 1573.625 | 37516.05 | 40519.775 | 40575.125 | 12029.375 | 12952.275 | 13688.25
| HQC 4 | 2050.275 | 2843.125 | 2872.45 | 1472.125 | 1715.075 | 1684.85 | 33830.45 | 33651.725 | 37384.425 | 18439.55 | 17803.075 | 17947.65 | 1932.325 | 2300.6 | 1976.525 | 1428.375 | 1619.9 | 1508.55 | 32139.3 | 33602.2 | 34625.175 | 11266.475 | 12745.5 | 12125.975
| HQC 5 | 2291.625 | 3421.2 | 2942.125 | 1414.575 | 1897.575 | 1273.425 | 31739.675 | 33308.925 | 32034.825 | 15696.2 | 18081.225 | 15572.45 | 1497.15 | 2315.925 | 1792.45 | 1695.05 | 1724.975 | 1341.975 | 35212.75 | 36793.7 | 33503.15 | 12507.175 | 13354.5 | 10147.425
| HQC 6 | 2785.6 | 3577.25 | 3008.5 | 1102.1 | 2347.675 | 1582.2 | 33588.5 | 44229.6 | 38205.725 | 17162 | 22968.9 | 19616.55 | 1353.6 | 1993.85 | 1977.65 | 1745.125 | 2766.85 | 1353.45 | 33056.45 | 43632.95 | 37580.425 | 10603.775 | 13888.975 | 13463.125
| Highest Level QC | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| HLQC 1 | 4346.7 | 4261.75 | 4778.225 | 1991.375 | 2398.075 | 2511.8 | 55282.525 | 49561.125 | 53985.675 | 30871.7 | 28768.475 | 32416.4 | 2643 | 2450.95 | 1822.3 | 2295.675 | 1445.95 | 1936.425 | 43240.35 | 37897.15 | 40892.45 | 12980.625 | 12203 | 13930.6
| HLQC 2 | 3789.3 | 3976.425 | 3846.375 | 2353.125 | 2141.175 | 2898.55 | 53454.2 | 50473.6 | 56262.625 | 28590.45 | 29905.175 | 33475.325 | 2204.225 | 2044 | 1879.125 | 1851.025 | 1641.1 | 2205.35 | 36322.05 | 38654.75 | 43213.175 | 12232.25 | 14098.6 | 14430.15
| HLQC 3 | 6429.675 | 4918.225 | 4637.125 | 2746.975 | 3042.15 | 2779.525 | 57177.05 | 57862.925 | 55614.1 | 34236.425 | 33679.625 | 33178.65 | 2558 | 2434.25 | 1903.725 | 2336.375 | 1652.525 | 1668.875 | 46418.625 | 43211.625 | 42433.025 | 15822.275 | 16903.125 | 13751.7
| HLQC 4 | 5621.025 | 4734.725 | 4528.175 | 3155.15 | 2297.925 | 2708.275 | 61248.7 | 56382.65 | 62751.4 | 35173.025 | 33327.475 | 37516.025 | 2267.775 | 2273.075 | 2481.725 | 2240.125 | 1750 | 2278 | 43967.075 | 42704.75 | 49475.85 | 15827.85 | 15331.85 | 16789.6
| HLQC 5 | 3789.3 | 4672.575 | 3927.675 | 2560.475 | 3383.725 | 2883.05 | 46125.95 | 55775.725 | 54794.175 | 29141.45 | 35046.1 | 31547.4 | 1225.725 | 1532.55 | 1786.2 | 1124.5 | 1399.8 | 860.625 | 23406.95 | 29634.625 | 29322.2 | 6728.15 | 8943.125 | 9990.05
HLQC 6
| Blank QC | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| MatrixBlank 1 | 280.075 |  |  | 120.1 |  |  | 322.8 |  |  | 240.075 |  |  | 160.075 |  |  | 0 |  |  | 302.675 |  |  | 140.1
| MatrixBlank 2 | 100 |  |  | 92.6 |  |  | 340.275 |  |  | 140.125 |  |  | 20 |  |  | 0 |  |  | 280.25 |  |  | 100.075
| TotalBlank 1 | 330.275 |  |  | 60.05 |  |  | 271.375 |  |  | 471.475 |  |  | 100.075 |  |  | 20.025 |  |  | 420.225 |  |  | 242.625
| TotalBlank 2 | 20 |  |  | 0 |  |  | 302.7 |  |  | 200.075 |  |  | 21.275 |  |  | 100.05 |  |  | 140.075 |  |  | 101.4
| TotalBlank 3 | 220 |  |  | 120.15 |  |  | 320.225 |  |  | 276.4 |  |  | 20 |  |  | 100.05 |  |  | 192.55 |  |  | 80.05
Table 13-1 FIA-MRM PooledQC. Please report the values of each of the three replicates separately for each compound and standard.
| pooled QC | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 |  |  | Repl 1 |  |  | Repl 1 |  |  | Repl 1 |  |  | Repl 1 |  |  | Repl 1 |  |  | Repl 1 |  |  | Repl 1
| pooledQC1 | 2364 |  |  | 1172 |  |  | 28648.375 |  |  | 10249.5 |  |  | 1503.475 |  |  | 1822.7 |  |  | 38366.075 |  |  | 13236.775
| pooledQC2 | 3053.725 |  |  | 850.575 |  |  | 31505.6 |  |  | 11094.85 |  |  | 2111.275 |  |  | 1790.225 |  |  | 41974.375 |  |  | 13069.075
| pooledQC3 | 2543.975 |  |  | 874.3 |  |  | 31446.5 |  |  | 10108.025 |  |  | 2046.325 |  |  | 1912.55 |  |  | 38391.975 |  |  | 12121.8
| pooledQC4 | 2817.075 |  |  | 800.4 |  |  | 32092.85 |  |  | 11735.5 |  |  | 1508.6 |  |  | 1704.95 |  |  | 39395.175 |  |  | 13239.4
| pooledQC5 | 2665.45 |  |  | 655.5 |  |  | 37125 |  |  | 12787.05 |  |  | 2007.525 |  |  | 1876.45 |  |  | 44422.375 |  |  | 16177.8
| pooledQC6 | 2928.65 |  |  | 1143.2 |  |  | 35489.175 |  |  | 12805.425 |  |  | 2455.275 |  |  | 1895.1 |  |  | 45592.925 |  |  | 14394.6
| pooledQC7 | 2150.35 |  |  | 930.675 |  |  | 27352.225 |  |  | 9572.05 |  |  | 1539.825 |  |  | 1329.5 |  |  | 36202.95 |  |  | 12609.6
| pooledQC8 | 3102.225 |  |  | 828.225 |  |  | 30900.525 |  |  | 10706.2 |  |  | 1781.275 |  |  | 1733.5 |  |  | 38729.375 |  |  | 14147.725
| pooledQC9 | 2224.125 |  |  | 940.7 |  |  | 26909.725 |  |  | 10012.725 |  |  | 2016.475 |  |  | 1578.65 |  |  | 35597.375 |  |  | 11590.2
| pooledQC10 | 2301.9 |  |  | 908.325 |  |  | 29543.975 |  |  | 8899.825 |  |  | 1937.575 |  |  | 1629.675 |  |  | 34489.975 |  |  | 10724.175
| Blank QC | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 |  |  | Repl 1 |  |  | Repl 1 |  |  | Repl 1 |  |  | Repl 1 |  |  | Repl 1 |  |  | Repl 1 |  |  | Repl 1
| MatrixBlank1 | 3252.275 |  |  | 889.325 |  |  | 36221.875 |  |  | 12928.375 |  |  | 130.1 |  |  | 80 |  |  | 442.975 |  |  | 40.05
| MatrixBlank2 | 2511.675 |  |  | 970.625 |  |  | 29341.325 |  |  | 11310.7 |  |  | 40.05 |  |  | 20.025 |  |  | 360.25 |  |  | 163.8
| TotalBlank1 | 260.175 |  |  | 40.025 |  |  | 274.025 |  |  | 422.8 |  |  | 20.025 |  |  | 40.05 |  |  | 373.975 |  |  | 200.2
| TotalBlank2 | 100.075 |  |  | 20.025 |  |  | 370.2 |  |  | 236.3 |  |  | 0 |  |  | 60 |  |  | 633.025 |  |  | 160.125
| TotalBlank3 | 101.35 |  |  | 20 |  |  | 400.25 |  |  | 281.35 |  |  | 100.075 |  |  | 100 |  |  | 400.35 |  |  | 120.075

## Raw Peak Areas
Sample_Name | Cer d18:1/16:0 [M>264] | Cer d18:1/16:0-d7 [M>264] | Cer d18:1/18:0 [M>264] | Cer d18:1/18:0-d7 [M>264] | Cer d18:1/24:0 [M>264] | Cer d18:1/24:0-d7 [M>264] | Cer d18:1/24:1 [M>264] | Cer d18:1/24:1-d7 [M>264] | Cer d18:1/22:0 [M>264]
TotalBlank-01 | 260.175 | 20.025 | 40.025 | 40.05 | 274.025 | 373.975 | 422.8 | 200.2 | 220.2
MatrixBlank-01 | 3252.275 | 130.1 | 889.325 | 80 | 36221.875 | 442.975 | 12928.375 | 40.05 | 8708.975
Line1-STD1_01 | 49915.65 | 2278.025 | 43152.1 | 2008.775 | 556399.2 | 47217.65 | 497075.95 | 15777.6 | 3867.875
Line1-STD1_02 | 50570.3 | 3077.55 | 44283.775 | 2165.2 | 595789.85 | 49290.55 | 547468.775 | 16733.975 | 3898.125
Line1-STD1_03 | 65938.425 | 3926.575 | 60417.025 | 3377.7 | 815486.1 | 71659.15 | 739441.75 | 24546.6 | 4850.775
Line1-STD2_01 | 23234.6 | 3414.675 | 22521.8 | 2313.9 | 282234.95 | 49801.65 | 261176.325 | 18062.525 | 1892.5
Line1-STD2_02 | 23496.2 | 2124.1 | 23441.825 | 2013.775 | 301585.35 | 56339.025 | 274881.45 | 16186.325 | 1577.35
Line1-STD2_03 | 26418.875 | 2560.275 | 25075.825 | 2645.575 | 338087.85 | 59030.525 | 304703.975 | 19365.325 | 2291.525
Line1-STD3_01 | 2822.2 | 2581.575 | 2017.625 | 2028.9 | 29210.85 | 48729.775 | 27367.05 | 15581.925 | 662.975
Line1-STD3_02 | 2403.85 | 2279.05 | 2367.825 | 2176.525 | 31166.825 | 49848.2 | 27493.675 | 17436.575 | 526.55
Line1-STD3_03 | 3145.825 | 2171.45 | 2521.925 | 2094.15 | 32924.05 | 51326.95 | 29887.875 | 17032.025 | 534.225
Line1-STD4_01 | 432.85 | 2436.575 | 251.5 | 2381.775 | 6902.85 | 44605.15 | 5836.2 | 15039.65 | 260.225
Line1-STD4_02 | 540.525 | 2604.15 | 720.5 | 2135.3 | 5900.475 | 44610.15 | 5152.925 | 15150.55 | 200.1
Line1-STD4_03 | 749.25 | 2485.2 | 120.1 | 2174.075 | 6652.525 | 43477.65 | 5212.8 | 15132.325 | 80
Line1-STD5_01 | 304.025 | 2006.225 | 362.8 | 1897.675 | 2813.65 | 42709.775 | 2823.075 | 13552.5 | 400.275
Line1-STD5_02 | 480.4 | 2101.6 | 465.25 | 2879.75 | 2979.85 | 43222.5 | 2307.575 | 13453.275 | 160
Line1-STD5_03 | 300.225 | 2629.475 | 220.15 | 1893.875 | 3090.2 | 45566.55 | 2914.45 | 14800.825 | 222.75
Line1-STD6_01 | 80.05 | 2047.575 | 222.65 | 1535.875 | 2830.825 | 40856.025 | 2020.25 | 13373.975 | 100.1
Line1-STD6_02 | 260.1 | 2723.375 | 140.175 | 1856.325 | 2223.025 | 40543.8 | 2269.35 | 14127.775 | 242.7
Line1-STD6_03 | 254.025 | 2256.55 | 254.025 | 2135.1 | 2733.525 | 40224.825 | 2207.525 | 14319.35 | 180.225
PooledQC-01 | 2364 | 1503.475 | 1172 | 1822.7 | 28648.375 | 38366.075 | 10249.5 | 13236.775 | 7199.025
PooledQC-02 | 3053.725 | 2111.275 | 850.575 | 1790.225 | 31505.6 | 41974.375 | 11094.85 | 13069.075 | 8277.875
SRM1950-01_01 | 3213.825 | 2633 | 998.3 | 3004.725 | 28865.475 | 52858.475 | 9770.925 | 16865.5 | 8449.4
SRM1950-01_02 | 3632.5 | 3109.925 | 1036.9 | 2570.55 | 26949.9 | 54372.075 | 11167.025 | 16581.1 | 8121.7
SRM1950-01_03 | 3212.275 | 2370.4 | 836.875 | 1957.475 | 27356.7 | 46812.275 | 10140.85 | 15558.9 | 7437.6
SRM1950-02_01 | 2945.75 | 2061.5 | 1065.7 | 1331 | 27751.75 | 36677.2 | 9397.05 | 11956 | 8121.7
SRM1950-02_02 | 2227.925 | 2219.45 | 901.9 | 2168.85 | 28110.075 | 39662.05 | 9598.125 | 13002.175 | 7434.55
SRM1950-02_03 | 2434.25 | 2136.25 | 900.575 | 1766.2 | 29754.325 | 38205.95 | 10686.05 | 14179.525 | 8032.05
SRM1950-03_01 | 2264.275 | 2605.475 | 1096.95 | 1965.25 | 30579.7 | 39474.7 | 11100.6 | 13829.675 | 8345.725
SRM1950-03_02 | 2768.25 | 2565.425 | 631.75 | 1837.7 | 30212.875 | 40208.975 | 10090.1 | 13268.175 | 7478.3
SRM1950-03_03 | 3137.4 | 1629.975 | 1080.55 | 1880.575 | 31201.55 | 38987.625 | 9925.325 | 13374.9 | 8166.825
SRM1950-04_01 | 2265.15 | 1902.7 | 1043.325 | 2037.8 | 28413.6 | 39581.95 | 8631.675 | 13867.275 | 6986.15
SRM1950-04_02 | 2554.15 | 2188.825 | 865.425 | 1612.7 | 26767.425 | 39818.825 | 9546.175 | 13293.575 | 8574.8
SRM1950-04_03 | 2401.8 | 2243.85 | 711.775 | 1775.075 | 30062.925 | 40330.825 | 9624.975 | 13268.575 | 8115.8
SRM1950-05_01 | 2881.575 | 2394.225 | 844.5 | 1877.45 | 31473.625 | 40334.9 | 10838.9 | 14761.1 | 9434.15
SRM1950-05_02 | 2578.5 | 2625.575 | 770.475 | 1353.65 | 27135.725 | 38986.9 | 9774.075 | 13249.85 | 7175.45
SRM1950-05_03 | 2504.15 | 2320.3 | 671.55 | 1723.975 | 31826.025 | 40396.725 | 11060.2 | 13809.025 | 8064.175
SRM1950-06_01 | 2695.65 | 1523.4 | 863.3 | 1833.875 | 27222.225 | 36838.025 | 9145.45 | 12670.95 | 7488.975
SRM1950-06_02 | 2770.5 | 1670.925 | 795.475 | 1835.2 | 28865.425 | 40323.175 | 11599.375 | 14606.225 | 8031.975
SRM1950-06_03 | 2693.075 | 2059.025 | 981.925 | 1683.6 | 27613.9 | 36921.4 | 9716.175 | 13183.975 | 7137.175
PooledQC-03 | 2543.975 | 2046.325 | 874.3 | 1912.55 | 31446.5 | 38391.975 | 10108.025 | 12121.8 | 8129.5
PooledQC-04 | 2817.075 | 1508.6 | 800.4 | 1704.95 | 32092.85 | 39395.175 | 11735.5 | 13239.4 | 7614.5
T1D-01_01 | 2056.45 | 2111.25 | 1105.65 | 2231.7 | 28655 | 40621.825 | 10844.45 | 13636.1 | 8073.625
T1D-01_02 | 2032.575 | 2574.15 | 1055.825 | 1645.05 | 29017.225 | 37306.125 | 9627 | 13384.95 | 8379.375
T1D-01_03 | 2484.025 | 2379.15 | 852 | 1995.175 | 31473.475 | 42906.925 | 10243.475 | 13710.825 | 8857.275
T1D-02_01 | 2790.775 | 2050.4 | 1012.15 | 1801.55 | 27547.85 | 39331.875 | 10082.15 | 13162.7 | 7942.95
T1D-02_02 | 2609.825 | 2464.2 | 1360.85 | 1658.85 | 29982.625 | 39706.525 | 10571.625 | 13974.525 | 7082.425
T1D-02_03 | 2332.375 | 2307.7 | 670.5 | 1976.425 | 32735.4 | 43829.825 | 11337.025 | 14180.925 | 8568.675
T1D-03_01 | 2327 | 2704.3 | 1190.625 | 2328.9 | 35299.975 | 49070.6 | 12459.425 | 16549.875 | 8908.95
T1D-03_02 | 3067.05 | 2397.875 | 941.95 | 2028.675 | 31477.9 | 45420.975 | 11327.775 | 16860.325 | 8876.9
T1D-03_03 | 2574.075 | 2526.85 | 1274.4 | 1973.8 | 33083.175 | 47562.025 | 12707.125 | 16254.65 | 8384.75
T1D-04_01 | 2598.175 | 2571.675 | 667.975 | 2520.4 | 33649.225 | 42310.9 | 11080.225 | 14345.925 | 8313.175
T1D-04_02 | 3253.575 | 2420.4 | 953.075 | 1741.575 | 34557.325 | 46041.525 | 11698.375 | 15076.05 | 8426.025
T1D-04_03 | 3003.125 | 2094.3 | 856.8 | 2495.4 | 35737.25 | 46868.35 | 12497.6 | 13553.375 | 8652.975
T1D-05_01 | 3472.225 | 2954 | 1339.6 | 2046.25 | 37806.45 | 48345.625 | 12391.325 | 16509.925 | 9422.35
T1D-05_02 | 2534.45 | 2272.825 | 1221.95 | 1846.35 | 31799.075 | 45685.775 | 11543.5 | 16110.15 | 8861.95
T1D-05_03 | 2174.2 | 2390.65 | 1258.6 | 2315.325 | 32459.225 | 43625.725 | 12064.625 | 16435.775 | 8414.725
T1D-06_01 | 2713.125 | 2544.15 | 900.375 | 1567.25 | 31298.65 | 42908.15 | 11400.25 | 15043.675 | 7353.975
T1D-06_02 | 2273.75 | 2086.55 | 1010.8 | 1938.775 | 33468.15 | 45594.825 | 10943.725 | 14091.425 | 8384.65
T1D-06_03 | 2674.425 | 2369.15 | 1091.875 | 2438.075 | 33460.625 | 45897.525 | 11390.525 | 14865.475 | 8623.775
TotalBlank-02 | 100.075 | 0 | 20.025 | 60 | 370.2 | 633.025 | 236.3 | 160.125 | 20.025
PooledQC-05 | 2665.45 | 2007.525 | 655.5 | 1876.45 | 37125 | 44422.375 | 12787.05 | 16177.8 | 10455.875
PooledQC-06 | 2928.65 | 2455.275 | 1143.2 | 1895.1 | 35489.175 | 45592.925 | 12805.425 | 14394.6 | 8105.6
YoungAA-01_01 | 2016.625 | 3121.875 | 501.5 | 2025.15 | 24499.475 | 49255.075 | 8222.75 | 18113.225 | 5933.7
YoungAA-01_02 | 2392.975 | 2739.65 | 700.6 | 2332.675 | 23839.475 | 52509.425 | 8386.825 | 18354.1 | 7280
YoungAA-01_03 | 2359.35 | 2839.6 | 491.425 | 2339.6 | 24969.65 | 55496.825 | 8474.9 | 16331.725 | 6414.125
YoungAA-02_01 | 2475.325 | 3034.525 | 603.05 | 2474.25 | 25731.4 | 53771.2 | 7180.3 | 18994.25 | 6934.975
YoungAA-02_02 | 2435.775 | 2124.975 | 440.5 | 2440.4 | 25123.65 | 50728.375 | 6537.275 | 16566.05 | 6986.9
YoungAA-02_03 | 2077.575 | 2428.225 | 641.9 | 1872.35 | 24586.4 | 48105.525 | 8025.95 | 17136.125 | 6551.7
YoungAA-03_01 | 1940.175 | 2723.35 | 918.3 | 1867.425 | 25760.9 | 49133.45 | 8698.65 | 16076.4 | 6573.325
YoungAA-03_02 | 1609.65 | 2301.35 | 860.775 | 1574.75 | 25970.725 | 47873.275 | 6986.15 | 15299.125 | 6598.175
YoungAA-03_03 | 2208.95 | 1815.025 | 600.35 | 2522.05 | 23269.4 | 48080.625 | 7838.4 | 15586.5 | 5947.725
YoungAA-04_01 | 2062.575 | 2406.6 | 578.05 | 1974.875 | 21455.325 | 42639.05 | 7168.2 | 14857 | 5937.35
YoungAA-04_02 | 1872.425 | 2106.625 | 391.5 | 2205.45 | 24169.4 | 47400.425 | 7940.125 | 14714 | 6784.825
YoungAA-04_03 | 1902.6 | 2140.325 | 580.425 | 2524.6 | 21671 | 45811.225 | 7273.9 | 15340.7 | 5859.95
YoungAA-05_01 | 1820.05 | 2566.9 | 492.925 | 2017.625 | 23087.775 | 44240.975 | 7259.025 | 16146.7 | 6134.65
YoungAA-05_02 | 1837.925 | 2228.95 | 544.275 | 1723.95 | 24034.15 | 45469.9 | 6902.325 | 14725.775 | 5481.6
YoungAA-05_03 | 1460.85 | 2252.55 | 525.375 | 2436.725 | 21842.9 | 45302.4 | 7189.275 | 15772.725 | 6366.7
YoungAA-06_01 | 2212.025 | 2421.675 | 1014.525 | 2512.6 | 21955.075 | 42727.75 | 7847.875 | 18032.4 | 5985.725
YoungAA-06_02 | 2006.5 | 1906.075 | 525.3 | 2257.5 | 21001.425 | 46270.55 | 7450.1 | 16156.525 | 6199.575
YoungAA-06_03 | 2032.275 | 2315.375 | 532.875 | 2329.025 | 23867.75 | 45877.425 | 7197.625 | 15992.55 | 6306.8
PooledQC-07 | 2150.35 | 1539.825 | 930.675 | 1329.5 | 27352.225 | 36202.95 | 9572.05 | 12609.6 | 7354.925
PooledQC-08 | 3102.225 | 1781.275 | 828.225 | 1733.5 | 30900.525 | 38729.375 | 10706.2 | 14147.725 | 8431.275
hTAG-01_01 | 2711.975 | 2179 | 741.775 | 1827.325 | 34109.45 | 32095.325 | 11560.8 | 11767.225 | 8898.775
hTAG-01_02 | 2316.775 | 1605.075 | 910.4 | 1293.45 | 36400.175 | 32169.9 | 11121.025 | 10592.4 | 10334.85
hTAG-01_03 | 2481.975 | 1926.5 | 634.25 | 1621.025 | 33950.8 | 34240.4 | 10760.175 | 10316.95 | 9162.8
hTAG-02_01 | 2101.575 | 1963.9 | 728 | 1127.975 | 32158.725 | 31449.1 | 9674.675 | 11109.2 | 9912.55
hTAG-02_02 | 2623.1 | 1507.15 | 720.475 | 2282.975 | 35706.475 | 31559.525 | 11700.05 | 11227.85 | 8810.9
hTAG-02_03 | 2057.725 | 1469.825 | 525.25 | 1635.225 | 34206.525 | 30302.6 | 11803.05 | 9939.15 | 9097.525
hTAG-03_01 | 2061.425 | 1887.175 | 590.55 | 1347.075 | 34865.575 | 31738.775 | 11639.575 | 10148.275 | 9675.325
hTAG-03_02 | 2087.725 | 1173.35 | 1205.75 | 1384.625 | 33255.925 | 30789.875 | 11899.15 | 10101.875 | 8343.125
hTAG-03_03 | 2863.25 | 1512.05 | 834.525 | 1617.525 | 35773.025 | 29751.175 | 11640.85 | 10653.6 | 9472.625
hTAG-04_01 | 3016.1 | 1280.95 | 923.2 | 1271.075 | 35396.2 | 31812.075 | 10680.175 | 11344.375 | 7457.775
hTAG-04_02 | 2176.575 | 1654.775 | 624.1 | 1308.675 | 37223.7 | 31590.075 | 12586.25 | 10384.4 | 9481.65
hTAG-04_03 | 2616.75 | 2130.05 | 662.95 | 1309.625 | 37882.575 | 31544.05 | 9950.45 | 10296.025 | 9518.15
hTAG-05_01 | 2501.775 | 2036.5 | 734.275 | 1102.025 | 36139.875 | 34515.825 | 13452.525 | 12244.025 | 8715.1
hTAG-05_02 | 2551.825 | 1718.375 | 580.4 | 1825.25 | 37306.725 | 33591.475 | 10664.075 | 10516.1 | 10247.025
hTAG-05_03 | 2480.3 | 2180.3 | 865.875 | 1513.7 | 38554.075 | 35478.175 | 12134.375 | 12515.25 | 9927.825
hTAG-06_01 | 3152.175 | 2412.925 | 1174.3 | 1618.7 | 42000.425 | 38491.45 | 13295.1 | 12087.875 | 9828.075
hTAG-06_02 | 2180.1 | 1514.95 | 555.55 | 1932.875 | 36089.625 | 32198.75 | 10426 | 11126.475 | 8987.475
hTAG-06_03 | 2007.825 | 1608.675 | 740.625 | 1899.9 | 41538 | 35348.9 | 13724.825 | 12057.675 | 9961.025
PooledQC-09 | 2224.125 | 2016.475 | 940.7 | 1578.65 | 26909.725 | 35597.375 | 10012.725 | 11590.2 | 7782.6
PooledQC-010 | 2301.9 | 1937.575 | 908.325 | 1629.675 | 29543.975 | 34489.975 | 8899.825 | 10724.175 | 8426.1
Line2-STD6_01 | 43825.4 | 2650.625 | 42488.475 | 1841.35 | 622880.875 | 52155.95 | 546000.45 | 16976.1 | 3972.675
Line2-STD6_02 | 48311.5 | 2849.525 | 43733.425 | 2737.9 | 620589.9 | 50666.05 | 567257.325 | 17111.35 | 3386.425
Line2-STD6_03 | 55717.25 | 2752.05 | 50988.55 | 2594.15 | 711007.675 | 62372.725 | 632902.1 | 20238.7 | 4609.75
Line2-STD5_01 | 23787.375 | 1992.475 | 21273.4 | 2584 | 297065.225 | 53154.25 | 276183.2 | 17336.275 | 1878.5
Line2-STD5_02 | 23994.075 | 2185.35 | 22169.1 | 2397.8 | 292102.925 | 52755.3 | 261785.3 | 16667.175 | 1893.625
Line2-STD5_03 | 24565.95 | 2735.675 | 20697.75 | 2593.65 | 306141.55 | 53843.3 | 282214.575 | 17802.45 | 2195.25
Line2-STD4_01 | 2565.35 | 2663.45 | 2082.65 | 2634.725 | 32358.325 | 57317.425 | 31356 | 19149.025 | 271.3
Line2-STD4_02 | 2716.7 | 3127.475 | 2296.4 | 2037.45 | 30317.6 | 58235.15 | 29176.45 | 18444.5 | 291.475
Line2-STD4_03 | 2517.725 | 3245.95 | 2898.175 | 1989.05 | 34947.8 | 55794.225 | 29954.4 | 19189.35 | 380.25
Line2-STD3_01 | 640.35 | 2797.925 | 663.05 | 2080.125 | 6353.4 | 48523.225 | 5283.95 | 15662.525 | 431.5
Line2-STD3_02 | 858.075 | 1723.925 | 402.875 | 1872.45 | 4411.8 | 38381.8 | 4230.65 | 12464.925 | 123.9
Line2-STD3_03 | 900.675 | 2579.6 | 462.775 | 2635.425 | 6360.5 | 54984.925 | 6609.05 | 17065.725 | 220.15
Line2-STD2_01 | 290.2 | 2060.175 | 460.45 | 2040.125 | 2872.95 | 46729.1 | 2938.525 | 15922.325 | 200.075
Line2-STD2_02 | 511.65 | 1660.85 | 280.3 | 1504.9 | 3134.85 | 43190.625 | 2734.425 | 14889.025 | 60
Line2-STD2_03 | 282.65 | 2270.325 | 231.425 | 1912.4 | 2708.4 | 42376.4 | 2481.875 | 14966.25 | 260.175
Line2-STD1_01 | 82.55 | 1983.75 | 100.075 | 1583.825 | 2480.425 | 37552.5 | 2311.575 | 13825.95 | 280.225
Line2-STD1_02 | 525.325 | 2209.225 | 200.2 | 1651.2 | 2527.9 | 43160.375 | 2403.95 | 13943.975 | 221.375
Line2-STD1_03 | 470.5 | 2410.475 | 302.775 | 1272.175 | 1913.725 | 42369.55 | 1977.6 | 13349.9 | 120.075
MatrixBlank-02 | 2511.675 | 40.05 | 970.625 | 20.025 | 29341.325 | 360.25 | 11310.7 | 163.8 | 7299
TotalBlank-03 | 101.35 | 100.075 | 20 | 100 | 400.25 | 400.35 | 281.35 | 120.075 | 120.05

## Raw Peak Areas - IntraAssay
Sample_Name | Cer d18:1/16:0 [M>264] | Cer d18:1/16:0-d7 [M>264] | Cer d18:1/18:0 [M>264] | Cer d18:1/18:0-d7 [M>264] | Cer d18:1/24:0 [M>264] | Cer d18:1/24:0-d7 [M>264] | Cer d18:1/24:1 [M>264] | Cer d18:1/24:1-d7 [M>264] | Cer d18:1/22:0 [M>264]
TotalBlank-01 | 330.275 | 100.075 | 60.05 | 20.025 | 271.375 | 420.225 | 471.475 | 242.625 | 301.525
MatrixBlank-01 | 280.075 | 160.075 | 120.1 | 0 | 322.8 | 302.675 | 240.075 | 140.1 | 140.075
LLQC-01_001 | 1133.075 | 2990.625 | 640.4 | 1957.7 | 15815.225 | 41931.475 | 5314.1 | 15478.025 | 4061.4
LLQC-01_002 | 1554.775 | 2012.625 | 600.5 | 1829.75 | 17147.675 | 47127.7 | 6261.75 | 14306.725 | 4896.3
LLQC-01_003 | 1515.925 | 2470.575 | 444.15 | 1997.575 | 18687.2 | 44104.6 | 6409.275 | 17187.85 | 4426.775
LLQC-02_001 | 1463.725 | 2327.925 | 390.3 | 1916.1 | 16191.225 | 40668.225 | 6540.8 | 14170.3 | 4260.05
LLQC-02_002 | 1130.975 | 1951.325 | 812.825 | 1945.225 | 15865.8 | 38833.45 | 5741.925 | 14036.75 | 4806.3
LLQC-02_003 | 1616.475 | 2297.7 | 482.875 | 1797.375 | 19379.875 | 43397.425 | 5721.35 | 15444.725 | 3973.025
LLQC-03_001 | 1538.875 | 2220.3 | 640.375 | 2016.625 | 16902.625 | 44196.7 | 5179.025 | 13910.875 | 5274.975
LLQC-03_002 | 1660.175 | 2457.775 | 482.8 | 1677.425 | 17813.825 | 47277.3 | 5861.825 | 17224.35 | 5217
LLQC-03_003 | 1435.975 | 2650.775 | 577.95 | 1867.425 | 15789.25 | 41880.1 | 6194.625 | 15854.225 | 4483.375
LLQC-04_001 | 1391 | 2298.975 | 480.375 | 2215.4 | 16719.225 | 44522.45 | 7489.125 | 14774.075 | 4798.55
LLQC-04_002 | 1555.775 | 2528.175 | 520.35 | 1916.3 | 14422.775 | 42855.475 | 5052.9 | 13486.525 | 5096.35
LLQC-04_003 | 1578.6 | 2449.4 | 463.05 | 1763.625 | 16754.825 | 44375.875 | 6512.625 | 16272.45 | 4535.55
LLQC-05_001 | 1752.775 | 3444.4 | 725.5 | 1972.55 | 18401.95 | 47889.45 | 6805.125 | 15537.45 | 4819.8
LLQC-05_002 | 1627.225 | 2858.075 | 420.3 | 2372.925 | 18120.85 | 47743.775 | 5636.325 | 14938.575 | 4621.475
LLQC-05_003 | 1313.175 | 2511.6 | 580.325 | 2034.125 | 16851.55 | 47605.375 | 7486.05 | 17910.5 | 4488.425
LLQC-06_001 | 1163.525 | 2656.75 | 320.125 | 2040.075 | 17249.15 | 46341.975 | 5871.475 | 15643.575 | 4541.425
LLQC-06_002 | 1406.125 | 2719.65 | 474 | 1948.8 | 17558.025 | 45411.725 | 6024.325 | 16485.925 | 5359.9
LLQC-06_003 | 1819.95 | 2666.075 | 485.225 | 2040.3 | 18935.175 | 49618.1 | 5480.65 | 18746.975 | 4994.5
LQC-01_001 | 1416.175 | 1976.325 | 630.675 | 1903.8 | 15358.75 | 38270.75 | 4967.925 | 14163.05 | 4072.75
LQC-01_002 | 1336.15 | 1715.075 | 373.975 | 1453.7 | 12647.875 | 32655.95 | 4380.55 | 11691.625 | 3650.925
LQC-01_003 | 1160.725 | 2041.6 | 589.05 | 1379.95 | 12491.8 | 31353.425 | 4319.375 | 11526.425 | 3629.95
LQC-02_001 | 1043.125 | 1656.05 | 484.075 | 1577.225 | 13145.3 | 36358.9 | 4879.775 | 12254 | 3807.525
LQC-02_002 | 988.125 | 1859.825 | 565.425 | 1339.875 | 11965.3 | 33887.25 | 3792.825 | 11464.15 | 3215.925
LQC-02_003 | 785.55 | 2122.625 | 531.575 | 1909 | 12885.05 | 33340.75 | 4673.3 | 10308.025 | 2850.6
LQC-03_001 | 1045.75 | 1534.8 | 420.4 | 1787.5 | 14870.25 | 40069.2 | 5270.575 | 12942.8 | 4566.275
LQC-03_002 | 1338.15 | 2066.625 | 471.5 | 2057.625 | 15349.45 | 38128.6 | 4140.325 | 11763.575 | 3520.15
LQC-03_003 | 1274.775 | 1394.775 | 440.3 | 2133.825 | 13879.225 | 36281.275 | 3827.525 | 11939.7 | 3712.025
LQC-04_001 | 1515 | 1743.8 | 584.3 | 1666.1 | 13620.375 | 34378.9 | 4573.275 | 12545.025 | 3322.15
LQC-04_002 | 750.675 | 2065.225 | 472.7 | 1452.275 | 13034.825 | 34183.225 | 4813.6 | 10633.725 | 3674.825
LQC-04_003 | 1187.175 | 1791.025 | 620.275 | 1513.575 | 12354.65 | 33883.125 | 4767.05 | 9771.275 | 3823.85
LQC-05_001 | 1365.85 | 2080.375 | 220.05 | 1380.9 | 13692.85 | 34882.575 | 4874.225 | 11187.25 | 3580.9
LQC-05_002 | 1506.125 | 1326.85 | 390.275 | 1467.15 | 13554.275 | 34830.55 | 4728.05 | 12147.85 | 3659.75
LQC-05_003 | 1013.25 | 1897.425 | 555.375 | 1674.925 | 14788.55 | 36894.675 | 4930.475 | 12549.6 | 4641.7
LQC-06_001 | 1449.8 | 1915.375 | 60.05 | 2041.425 | 15643.075 | 39755.6 | 5021.2 | 14693.275 | 4166.825
LQC-06_002 | 1549.725 | 1644.95 | 474.075 | 1130.9 | 15584.625 | 40312.325 | 5619.05 | 12809.8 | 3812.525
LQC-06_003 | 1364.8 | 2098.825 | 430.4 | 2070.325 | 16618.45 | 42869.525 | 5446.1 | 15236.8 | 4078.8
MQC-01_001 | 2918.375 | 2451.425 | 1058.225 | 2172.85 | 26285.7 | 36733.175 | 9768.025 | 12837.175 | 7429.175
MQC-01_002 | 2987.25 | 2023.9 | 980.7 | 1865.175 | 31575.5 | 39988.275 | 10389.925 | 13866.7 | 7834.1
MQC-01_003 | 2911.925 | 2316.175 | 1290.575 | 1853.85 | 29287.85 | 38587.7 | 10530.575 | 13051.6 | 7778.15
MQC-02_001 | 2689.25 | 1450.85 | 734.3 | 1898.625 | 26769.575 | 34367.25 | 8582.925 | 12359.6 | 6642.025
MQC-02_002 | 1735.1 | 1469.975 | 815.5 | 1651.075 | 27442.05 | 31208.3 | 8387.525 | 11138.5 | 8696.575
MQC-02_003 | 2100.175 | 2037.7 | 856.975 | 1535.975 | 26062.575 | 31283.1 | 8409.2 | 12376.7 | 6779.325
MQC-03_001 | 1892.775 | 1225.75 | 540.35 | 1695.8 | 22455.375 | 28921.375 | 7114.925 | 10577.55 | 6043.825
MQC-03_002 | 1972.55 | 1949.05 | 921.825 | 1697.5 | 23885.825 | 32339.325 | 7911.125 | 10815.725 | 8324.675
MQC-03_003 | 1752.4 | 1300.85 | 1046.9 | 1388.325 | 22711.725 | 33021.95 | 7963 | 11806.775 | 6415.9
MQC-04_001 | 1647.45 | 1314.375 | 500.375 | 1558.7 | 21297.175 | 28381.675 | 8509.275 | 10606.875 | 6346.65
MQC-04_002 | 2015.225 | 1444.525 | 800.475 | 1757.45 | 28107.775 | 31615.875 | 9174.35 | 11488.625 | 7311.425
MQC-04_003 | 1922.675 | 1876.15 | 831.775 | 1307.15 | 23199.3 | 28015.45 | 8626.175 | 10528.15 | 6492.125
MQC-05_001 | 2341.475 | 1836.15 | 903.325 | 1372.25 | 25149.675 | 34286.925 | 8791.375 | 12414.525 | 6806.3
MQC-05_002 | 2339.075 | 1842.65 | 723.075 | 1033.175 | 27385.9 | 37188.475 | 8415.75 | 10641.7 | 6105.5
MQC-05_003 | 1985.125 | 1718.6 | 662.975 | 2016.4 | 23638.025 | 28888.3 | 7945.475 | 9899.75 | 5275.325
MQC-06_001 | 2004.95 | 1356.1 | 638 | 1601.1 | 22757.25 | 30967.975 | 7202.375 | 10889.225 | 7140.175
MQC-06_002 | 1464.825 | 1416.15 | 630.55 | 1462.35 | 20045.15 | 25393.975 | 7566.825 | 9365.8 | 5619.475
MQC-06_003 | 1786.45 | 2324.175 | 570.35 | 1592.725 | 19773.3 | 27249.525 | 7246.45 | 9219.975 | 6909.375
TotalBlank-02 | 20 | 21.275 | 0 | 100.05 | 302.7 | 140.075 | 200.075 | 101.4 | 273.875
HQC-01_001 | 2429.35 | 1733.875 | 1254.575 | 1412 | 31793.075 | 33735.425 | 14960.15 | 11545.75 | 5660.6
HQC-01_002 | 2400.375 | 1494.7 | 1257.075 | 1549.775 | 31040.075 | 33198.85 | 15128.35 | 12673.65 | 6169.35
HQC-01_003 | 2379.425 | 1495.975 | 1374.9 | 1639.675 | 27437.35 | 34007.55 | 15484.95 | 12258.25 | 5562.65
HQC-02_001 | 850.675 | 742.975 | 340.3 | 382.825 | 9155.625 | 8850.1 | 4683.025 | 2850.575 | 2299.15
HQC-02_002 | 739.125 | 683.1 | 462.775 | 570.475 | 9616.075 | 9242.35 | 4995.375 | 2937.975 | 2724.125
HQC-02_003 | 600.475 | 380.325 | 470.475 | 434.025 | 9620.325 | 8833.675 | 5536.325 | 2969.225 | 1363.65
HQC-03_001 | 3094.4 | 1939.9 | 1576.4 | 1497.55 | 34342.475 | 37516.05 | 16504.2 | 12029.375 | 6661.95
HQC-03_002 | 2479.325 | 1956.45 | 1663.5 | 1717.425 | 35692.525 | 40519.775 | 18462.4 | 12952.275 | 6028.425
HQC-03_003 | 3061.15 | 1765.15 | 1694.6 | 1573.625 | 35866.5 | 40575.125 | 17474.725 | 13688.25 | 7765.35
HQC-04_001 | 2050.275 | 1932.325 | 1472.125 | 1428.375 | 33830.45 | 32139.3 | 18439.55 | 11266.475 | 6629.475
HQC-04_002 | 2843.125 | 2300.6 | 1715.075 | 1619.9 | 33651.725 | 33602.2 | 17803.075 | 12745.5 | 7060.275
HQC-04_003 | 2872.45 | 1976.525 | 1684.85 | 1508.55 | 37384.425 | 34625.175 | 17947.65 | 12125.975 | 7850.35
HQC-05_001 | 2291.625 | 1497.15 | 1414.575 | 1695.05 | 31739.675 | 35212.75 | 15696.2 | 12507.175 | 6947.3
HQC-05_002 | 3421.2 | 2315.925 | 1897.575 | 1724.975 | 33308.925 | 36793.7 | 18081.225 | 13354.5 | 6242.05
HQC-05_003 | 2942.125 | 1792.45 | 1273.425 | 1341.975 | 32034.825 | 33503.15 | 15572.45 | 10147.425 | 5943.775
HQC-06_001 | 2785.6 | 1353.6 | 1102.1 | 1745.125 | 33588.5 | 33056.45 | 17162 | 10603.775 | 6055.325
HQC-06_002 | 3577.25 | 1993.85 | 2347.675 | 2766.85 | 44229.6 | 43632.95 | 22968.9 | 13888.975 | 9360.625
HQC-06_003 | 3008.5 | 1977.65 | 1582.2 | 1353.45 | 38205.725 | 37580.425 | 19616.55 | 13463.125 | 7338.525
HLQC-01_001 | 4346.7 | 2643 | 1991.375 | 2295.675 | 55282.525 | 43240.35 | 30871.7 | 12980.625 | 8893.925
HLQC-01_002 | 4261.75 | 2450.95 | 2398.075 | 1445.95 | 49561.125 | 37897.15 | 28768.475 | 12203 | 7541.275
HLQC-01_003 | 4778.225 | 1822.3 | 2511.8 | 1936.425 | 53985.675 | 40892.45 | 32416.4 | 13930.6 | 7757.8
HLQC-02_001 | 3789.3 | 2204.225 | 2353.125 | 1851.025 | 53454.2 | 36322.05 | 28590.45 | 12232.25 | 8392.825
HLQC-02_002 | 3976.425 | 2044 | 2141.175 | 1641.1 | 50473.6 | 38654.75 | 29905.175 | 14098.6 | 7866.4
HLQC-02_003 | 3846.375 | 1879.125 | 2898.55 | 2205.35 | 56262.625 | 43213.175 | 33475.325 | 14430.15 | 8150.425
HLQC-03_001 | 6429.675 | 2558 | 2746.975 | 2336.375 | 57177.05 | 46418.625 | 34236.425 | 15822.275 | 10268.05
HLQC-03_002 | 4918.225 | 2434.25 | 3042.15 | 1652.525 | 57862.925 | 43211.625 | 33679.625 | 16903.125 | 8736.225
HLQC-03_003 | 4637.125 | 1903.725 | 2779.525 | 1668.875 | 55614.1 | 42433.025 | 33178.65 | 13751.7 | 8085.025
HLQC-04_001 | 5621.025 | 2267.775 | 3155.15 | 2240.125 | 61248.7 | 43967.075 | 35173.025 | 15827.85 | 9547.7
HLQC-04_002 | 4734.725 | 2273.075 | 2297.925 | 1750 | 56382.65 | 42704.75 | 33327.475 | 15331.85 | 7518.325
HLQC-04_003 | 4528.175 | 2481.725 | 2708.275 | 2278 | 62751.4 | 49475.85 | 37516.025 | 16789.6 | 8414.575
HLQC-05_001 | 3789.3 | 1225.725 | 2560.475 | 1124.5 | 46125.95 | 23406.95 | 29141.45 | 6728.15 | 6639.875
HLQC-05_002 | 4672.575 | 1532.55 | 3383.725 | 1399.8 | 55775.725 | 29634.625 | 35046.1 | 8943.125 | 9765.8
HLQC-05_003 | 3927.675 | 1786.2 | 2883.05 | 860.625 | 54794.175 | 29322.2 | 31547.4 | 9990.05 | 9718.125
HLQC-06_001 | 4708.55 | 90.1 | 2485.475 | 60.025 | 50373.825 | 160.025 | 28293.3 | 83.825 | 7691.9
HLQC-06_002 | 4255.225 | 20 | 2394.675 | 80.075 | 53091.05 | 494.175 | 32145.4 | 120.1 | 8985.175
HLQC-06_003 | 4338.1 | 80.025 | 2987.025 | 40 | 48248.725 | 320.225 | 27657.35 | 80.075 | 7549.45
TotalBlank-03 | 220 | 20 | 120.15 | 100.05 | 320.225 | 192.55 | 276.4 | 80.05 | 342.725
MatrixBlank-02 | 100 | 20 | 92.6 | 0 | 340.275 | 280.25 | 140.125 | 100.075 | 240.2