## Participant Information
Participant Information
Participant Information |  | Group Name
Lab code
| Institution | Institution name
Street
City
Postal code
Country
| Contact person | Name
Address
Email
Phone

## Summary
Summary
Summary | This protocol describes the extraction and LC-MRM procedure for absolute quantification of Cer d18:1/16:0, Cer d18:1/18:0, Cer d18:1/24:0 and Cer d18:1/24:1 in human plasma according to a preferred LC-MS method
A calibration curve is built by injecting several dilutions of the non-labelled standards mixed with a fixed concentration of labelled standards. Each calibration point is prepared and analysed in triplicates (n=3).
To establish the endogenous concentration of the four selected ceramides in the study samples (NIST SRM1950, NIST high TAG, NIST T1D, and NIST young A-A) each sample is extracted and analyzed in six replicates (n =6).
Prepare the pooled QC by mixing 100 mL each of NIST SRM1950, NIST high TAG, NIST T1D, and NIST young AA (total 400 mL).
We suggest proceeding first to measure the ceramides concentration in the four study samples and then continue with the intra-assay validation to avoid wasting material.

## Extraction - Overview
Extraction Overview
Extraction Overview | Table 1 - Extraction - Materials
Chemicals, solvents and materials
Ethyl acetate (EtAc)
2-propanol (IPA)
5% Bovine Serum Albumin (BSA)
96-well plates/tubes
Sealing foil
2 mL DeepWell plates/tubes/LC vials
Table 2 - Extraction - Sample Types
| Sample type | Components
| Pooled Quality Control samples (QC) | 10 µL pooled QC |  |  | 20 µL Labelled IS |  | 570 µL EtAc:IPA 2:8 (v/v)
| Matrix Blank (MB) | 10 µL Matrix (pooled QC) only |  |  | - |  | 590 µL EtAc:IPA 2:8
| Study Samples (S) | 10 µL Human plasma (NIST SRM1950, high TAG, T1D, young AA) |  |  | 20 µL Labelled IS |  | 570 µL EtAc:IPA 2:8
| Total Blank (TB) | 10 µL 5% BSA |  |  | - |  | 590 µL EtAc:IPA 2:8
| Calibration line standards (STD) | 10 µL 5% BSA + 20 µL non labelled STD mixture (Table 5) |  |  | 20 µL Labelled IS |  | 550 µL EtAc:IPA 2:8
Table 3 - Extraction - Labelled Internal Standard Mixture Solution (provided)
| Labelled IS mixture in EtAc:IPA 2:8 (v/v) | Concentration |  |  |  |  | Unit | Vendor
| D7-Cer d18:1/16:0 | 0.125 |  |  |  |  | pmol/µL [µM] | Avanti Polar Lipids
| D7-Cer d18:1/18:0 | 0.050
| D7-Cer d18:1/24:0 | 1500
| D7-Cer d18:1/24:1 | 0.500
Table 4 - Extraction - Non-Labelled Standards - for STD1 Mixture Solution (provided)
| Non-labelled standards in EtAc:IPA 2:8 (v/v) | Concentration |  |  |  |  | Unit | Vendor
| Cer d18:1/16:0 | 2.0 |  |  |  |  | pmol/µL [µM] | Avanti Polar Lipids
| Cer d18:1/18:0 | 2.0
| Cer d18:1/24:0 | 20.0
| Cer d18:1/24:1 | 20.0
Table 5 - Serial dilutions of Non-Labelled Standards from STD1 Mixture Solution for calibration line
| Add 20 µL of | STD1 Conc. | STD2 Conc. | STD3 Conc. | STD4 Conc. | STD5 Conc. | STD6 Conc. | Unit
| Cer d18:1/16:0 | 2.0 | 1.0 | 0.10 | 0.02 | 0.01 | 0.008 | pmol/µL [µM]
| Cer d18:1/18:0 | 2.0 | 1.0 | 0.10 | 0.02 | 0.01 | 0.008
| Cer d18:1/24:0 | 20.0 | 10.0 | 1.0 | 0.2 | 0.1 | 0.08
| Cer d18:1/24:1 | 20.0 | 10.0 | 1.0 | 0.2 | 0.1 | 0.08

## Extraction - Procedure
Extraction Procedure
Extraction Procedure | Extraction Procedure
| Step | Description | Reference Table
| 1 | Prepare different dilutions of STD1 by diluting the provided mixture (Extraction - Overview, Table 4) with EtAc:IPA 2:8 as suggested in table | Extraction - Overview'!B34
| 2 | Thaw plasma on ice (+4°C) and bring to room temperature prior to extraction. Mix by pipetting before use.
| 3 | Prepare samples in plates/tubes manually or with a robot, according to specifications in column C of table | Extraction - Overview'!C14
| 4 | Add 20 µL of labelled internal standard mixture to samples according to column F of table | Extraction - Overview'!F14
| 5 | For the calibration line, add 20 µL of different dilutions of the non-labelled mixture (STD1-6).
| 6 | Add 590 µL of ethyl acetate:isopropanol (2:8, vol/vol) to Total Blank (TB) and Matrix Blank (MB); 550 µL to calibration line (STD); 570 µL to study samples and QCs according to column H of table | Extraction - Overview'!H14
| 7 | Mix samples 10 min by pipetting or vortexing
| 8 | Centrifuge samples for 10 min at 3000 x gat room temperature
| 9 | Transfer 50 µL of the clear supernatant to MS-vials or MS plates, close/seal, store samples at -20 °C prior to LC-MRM analysis.

## LC-MRM - Overview
LC-MRM Overview
LC-MRM Overview | Table 5 Chemicals and solvents
Ammonium acetate
Formic acid
Ultra-pure water
Isopropanol
Methanol
Table 7 LC-MRM LC-MS Details
| Property | Specifications
| Mass spectrometry | QQQ Agilent 6495
| LC | UPLC
| Injection volume (indicative) | 2 µL
| Weak wash | Water
| Strong wash | Acetonitrile
Pre-column
| Column | Agilent EclipsePlus C18, 2.1 × 100 mm id. 1.8 µm
| Column temperature | 40 °C
| Mobile phase A | 60% Methanol, 40% water, 0.2% formic acid, 10mM Ammonium acetate.
| Mobile phase B | 60% Methanol, 40% Isopropanol, 0.2% formic acid, 10mM Ammonium acetate
Table 8 LC-MRM LC gradient
| Time (minutes) | A% | B% |  | Flow (µL/min)
| 0.00 | 100 | 0 |  | 400
| 3.00 | 90 | 10 |  | 400
| 5.00 | 60 | 40 |  | 400
| 5.30 | 45 | 55 |  | 400
| 8.00 | 40 | 60 |  | 400
| 8.50 | 20 |  | 80 |  | 400
| 10.50 | 20 |  | 80 |  | 400
| 16.00 | 10 |  | 90 |  | 400
| 19.00 | 10 |  | 90 |  | 400
| 22.00 | 0 |  | 100 |  | 400
| 22.10 | 100 | 0 |  | 400
| 25 | 100 | 0 |  | 400
Table 9 LC-MRM MS conditions (EXAMPLE: this only applies to Sciex QTRAP, please adapt)
| Polarity mode | Pos
Ion spray
Curtain gas
Source temperature
Gas 1
Gas 2
DP
EP
Collision cell exit potential
Collision energy (CE)
Table 10 LC-MRM Mass transitions used for absolute quantification. On different systems, RT and CE might be different!
| Analyte | Mass transitions | IS | Mass transitions | RT (min) | CE
| Cer d18:1/16:0 | 538.5 – 264.3 | D7- Cer d18:1/16:0 | 545.6 - 271.3 | 12.1-12.2 | 25
C34H68NO3 – C18H34N
| Cer d18:1/18:0 | 566.6 – 264.3 | D7- Cer d18:1/18:0 | 573.6 – 271.3 | 13.3-13.5 | 25
C36H72NO3 – C18H34N
| Cer d18:1/24:0 | 650.6 – 264.3 | D7- Cer d18:1/24:0 | 657.7 – 271.3 | 17.2-17.4 | 25
C42H84NO3 – C18H34N
| Cer d18:1/24:1 | 648.6 – 264.3 | D7- Cer d18:1/24:1 | 655.7 – 271.3 | 16.0-16.2 | 25
C42H82NO3 – C18H34N
Table 11 Data processing
| Processing software | Agilent MassHunter

## LC-MRM - Sample Sequence
Sample Sequence (Example)
Sample Sequence (Example) | Sample Sequence 1
| Number | Sample Name | Remarks
| 1 | Total Blank 1
| 2 | Matrix blank
| 3 | Calibration line 1 from STD6 to 1 (n=3)
| 4 | […]
| 5 | Pooled QC
| 6 | Pooled QC
| 7 | SRM 1950-1
| 8 | SRM 1950-2
| 9 | SRM 1950-3
| 10 | SRM 1950-4
| 11 | SRM 1950-5
| 12 | SRM 1950-6
| 13 | Pooled QC
| 14 | Pooled QC
| 15 | T1D-1
| 16 | T1D-2
| 17 | T1D-3
| 18 | T1D-4
| 19 | T1D-5
| 20 | T1D-6
| 21 | Total Blank 2
| 22 | Pooled QC
| 23 | Pooled QC
| 24 | Young AA-1
| 25 | Young AA-2
| 26 | Young AA-3
| 27 | Young AA-4
| 28 | Young AA-5
| 29 | Young AA-6
| 30 | Pooled QC
| 31 | Pooled QC
| 32 | hTAG-1
| 33 | hTAG-2
| 34 | hTAG-3
| 35 | hTAG-4
| 36 | hTAG-5
| 37 | hTAG-6
| 38 | Pooled QC
| 39 | Pooled QC
| 40 | Calibration line 2 from STD1 to 6 (n=3)
| 41 | Intra-assay QCs: LLQC, LQC, MQC, HQC
| 42 | Matrix blank
| 43 | Total Blank 3

## LC-MRM - Intra Assay QC Proc
Intra Assay Procedure
Intra Assay Procedure | Quality Controls for Intra-Assay Variation
| Step | Description | Reference Table
| 1 | Prepare the Pooled NIST QC sample containing equal volumes of the four NIST reference plasma you received.
| 2 | The pooled QC representing middle quality control (MQC), low QC (LQC), and lower limit of quantification QC (LLOQ) is prepared by diluting MQC (e.g. two- and three-fold) with water.
| 3 | High QC (HQC) and upper limit of quantification (ULOQ) QC are prepared by spiking in 10 µL and 20 µL respectively of STD3 endogenous standards from table | Extraction - Overview'!B34
| 4 | Internal standards are added to all QC samples.
Intra-Assay Variation
| 6 | The precision and accuracy of the assay is determined for each of the four ceramides measured in LLOQ, LQC, MQC, HQC, and HLOQ.
| 7 | Intra-assay variation, precision, and accuracy is calculated for each of the ceramide in replicates of six (n = 6) at each QC concentration.
| 8 | The intra-assay precision (percentage coefficient variance, %CV) and accuracy (percentage accuracy, %Accuracy) are calculated from the nominal concentrations according to the formulas reported in corresponding sheets and tables of this document.
| 9 | Recovery estimation is not necessary for the published method but only for different ones. Stability estimation is not necessary for the published method but only for different ones. Calibration line is prepared in 5% BSA.

## LC-MRM - Intra Assay QC Res
INTRA ASSAY - QC RESULTS
INTRA ASSAY - QC RESULTS | Table 12 LC-MRM Calibration Line 1 Data (Beginning). Please report the values of each of the three replicates separately for each compound and standard. Each STD prepared in triplicate, injected once from each vial
| Calibration Line 1 | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| STD 1 | 724406 | 728437 | 754521 | 822340 | 829620 | 870966 | 31388293 | 31084632 | 33053390 | 21769283 | 22531890 | 22931371 | 30597 | 30326 | 32711 | 26864 | 30349 | 29993 | 1844375 | 1892690 | 1971422 | 445172 | 444661 | 470656
| STD 2 | 390646 | 389209 | 391892 | 440824 | 449607 | 453107 | 15413643 | 15563676 | 15954809 | 10818415 | 11083859 | 11133527 | 30951 | 29823 | 30976 | 28152 | 30081 | 28591 | 1784665 | 1824027 | 1819687 | 425327 | 435013 | 442515
| STD 3 | 48178 | 44670 | 48167 | 42079 | 42178 | 46069 | 1546845 | 1599073 | 1607324 | 1088706 | 1081727 | 1167605 | 27805 | 29337 | 29038 | 24335 | 26080 | 27766 | 1622089 | 1658669 | 1737367 | 422642 | 421416 | 425688
| STD 4 | 17251 | 18222 | 17381 | 12264 | 11750 | 12946 | 382173 | 396568 | 405108 | 271520 | 290164 | 290426 | 29405 | 27749 | 31556 | 29536 | 26837 | 27936 | 1652182 | 1722675 | 1683894 | 409161 | 422027 | 427899
| STD 5 | 12794 | 11825 | 13794 | 7603 | 7054 | 7760 | 216567 | 220208 | 228492 | 158971 | 153444 | 154928 | 28820 | 29465 | 27727 | 27196 | 27967 | 26705 | 1718484 | 1668250 | 1699348 | 448197 | 428299 | 419862
| STD 6 | 11036 | 13144 | 12380 | 5110 | 6449 | 6119 | 211296 | 202801 | 209534 | 126186 | 131513 | 133244 | 28173 | 30093 | 29619 | 29887 | 28147 | 29378 | 2133311 | 1739839 | 1736049 | 444907 | 437046 | 445876
Table 13 LC-MRM Calibration Line 2 Data (End). Please report the values of each of the three replicates separately for each compound and standard. Each STD prepared in triplicate, injected once from each vial
| Calibration Line 2 | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| STD 1 | 697701 | 649165 | 677983 | 756788 | 753029 | 757124 | 33965913 | 28696806 | 30323264 | 22471750 | 20103282 | 20364385 | 30838 | 27987 | 29334 | 25899 | 24647 | 26678 | 2193865 | 1809433 | 1887443 | 511335 | 445800 | 453160
| STD 2 | 384727 | 365829 | 363987 | 414150 | 409424 | 411886 | 15241256 | 15413567 | 15144521 | 10697622 | 10494584 | 10498060 | 30329 | 27912 | 25946 | 24611 | 26637 | 26282 | 1838899 | 1812081 | 1818270 | 452155 | 401333 | 432137
| STD 3 | 46275 | 51533 | 45502 | 46513 | 48775 | 49031 | 1628421 | 1725695 | 1704835 | 1149647 | 1147044 | 1186478 | 27522 | 29079 | 30978 | 27199 | 28530 | 27233 | 1776421 | 1784162 | 1799487 | 420442 | 428852 | 439920
| STD 4 | 12558 | 12035 | 12705 | 6838 | 8405 | 8159 | 220876 | 228846 | 229881 | 149437 | 151401 | 156017 | 25009 | 26997 | 27566 | 27040 | 26566 | 26827 | 1740185 | 1799683 | 1807762 | 413349 | 421962 | 420152
| STD 5 | 18383 | 17117 | 17246 | 11600 | 12039 | 12261 | 403433 | 415183 | 428355 | 280683 | 284485 | 281563 | 29657 | 29334 | 28886 | 28591 | 26984 | 27501 | 1731026 | 1765414 | 1791029 | 430054 | 434424 | 429107
| STD 6 | 11001 | 11898 | 11398 | 4438 | 5674 | 5454 | 178636 | 185452 | 187931 | 121390 | 123272 | 133926 | 30042 | 31229 | 30560 | 27307 | 28048 | 26435 | 1784794 | 1832784 | 1903322 | 430917 | 444288 | 454850
Table 14 LC-MRM Sample Data. Please report the values of each of the three replicates separately for each compound and standard. Each vial injected 3 times
| NIST SRM | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| SRM 1 | 44245 |  |  | 17212 |  |  | 1921981 |  |  | 440438 |  |  | 36299 |  |  | 35171 |  |  | 2472061 |  |  | 564374
| SRM 2 | 44637 |  |  | 16696 |  |  | 1784821 |  |  | 405819 |  |  | 29554 |  |  | 31246 |  |  | 2201377 |  |  | 517169
| SRM 3 | 43442 |  |  | 13025 |  |  | 1759030 |  |  | 406130 |  |  | 29507 |  |  | 34734 |  |  | 2206640 |  |  | 532053
| SRM 4 | 48168 |  |  | 18561 |  |  | 1837212 |  |  | 424716 |  |  | 32304 |  |  | 35840 |  |  | 2280309 |  |  | 542414
| SRM 5 | 42300 |  |  | 15355 |  |  | 1769185 |  |  | 416176 |  |  | 33293 |  |  | 29067 |  |  | 2238350 |  |  | 524312
| SRM 6 | 38519 |  |  | 15311 |  |  | 1732838 |  |  | 393455 |  |  | 34493 |  |  | 30469 |  |  | 2229946 |  |  | 534017
| NIST hTAG | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| hTAG 1 | 46711 |  |  | 17256 |  |  | 2416874 |  |  | 509264 |  |  | 28797 |  |  | 30748 |  |  | 2141412 |  |  | 504071
| hTAG 2 | 50354 |  |  | 15625 |  |  | 2788730 |  |  | 550362 |  |  | 32136 |  |  | 33516 |  |  | 2339327 |  |  | 540830
| hTAG 3 | 49426 |  |  | 17887 |  |  | 2977561 |  |  | 650927 |  |  | 30379 |  |  | 28767 |  |  | 2395142 |  |  | 563199
| hTAG 4 | 50028 |  |  | 18219 |  |  | 2854402 |  |  | 611641 |  |  | 36932 |  |  | 30941 |  |  | 2421923 |  |  | 572645
| hTAG 5 | 52070 |  |  | 17945 |  |  | 2843677 |  |  | 618002 |  |  | 36339 |  |  | 35858 |  |  | 2410607 |  |  | 595542
| hTAG 6 | 42522 |  |  | 14692 |  |  | 2487251 |  |  | 501197 |  |  | 32178 |  |  | 30740 |  |  | 2267440 |  |  | 530670
| NIST T1D | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| T1D 1 | 37517 |  |  | 18333 |  |  | 1809812 |  |  | 369843 |  |  | 35003 |  |  | 35115 |  |  | 2346411 |  |  | 564123
| T1D 2 | 43569 |  |  | 18092 |  |  | 2012488 |  |  | 409198 |  |  | 37966 |  |  | 38082 |  |  | 2587083 |  |  | 617575
| T1D 3 | 33087 |  |  | 17436 |  |  | 1590200 |  |  | 316825 |  |  | 30860 |  |  | 28361 |  |  | 2128985 |  |  | 480111
| T1D 4 | 36522 |  |  | 14270 |  |  | 1500750 |  |  | 367445 |  |  | 28902 |  |  | 31684 |  |  | 2090703 |  |  | 486952
| T1D 5 | 33894 |  |  | 21238 |  |  | 1548276 |  |  | 332979 |  |  | 29464 |  |  | 28304 |  |  | 2002024 |  |  | 488012
| T1D 6 | 40723 |  |  | 14266 |  |  | 1767685 |  |  | 359392 |  |  | 31653 |  |  | 29190 |  |  | 2332066 |  |  | 538678
| NIST Young AA | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| yA-A 1 | 28041 |  |  | 10574 |  |  | 1328918 |  |  | 276930 |  |  | 31219 |  |  | 33920 |  |  | 2291397 |  |  | 523630
| yA-A 2 | 26709 |  |  | 8824 |  |  | 1137060 |  |  | 244292 |  |  | 30320 |  |  | 32788 |  |  | 2143086 |  |  | 514448
| yA-A 3 | 28079 |  |  | 10420 |  |  | 1323930 |  |  | 244035 |  |  | 36744 |  |  | 39098 |  |  | 2447512 |  |  | 578695
| yA-A 4 | 25742 |  |  | 7293 |  |  | 1159612 |  |  | 251011 |  |  | 29971 |  |  | 30824 |  |  | 2140221 |  |  | 498888
| yA-A 5 | 29147 |  |  | 10973 |  |  | 1301467 |  |  | 270824 |  |  | 33540 |  |  | 32638 |  |  | 2343470 |  |  | 529795
| yA-A 6 | 19524 |  |  | 5998 |  |  | 915012 |  |  | 180648 |  |  | 30981 |  |  | 29008 |  |  | 2093059 |  |  | 492264
Table 15 LC-MRM Intra Assay QC Data. Please report the values of each of the three replicates separately for each compound and standard. QCs have to be measured in one occasion only. The independent 6 replicates injected 3 times each
| Lowest Level QC | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| LLQC 1 | 12527 |  |  | 5074 |  |  | 474661 |  |  | 110493 |  |  | 33335 |  |  | 28322 |  |  | 1889067 |  |  | 495868
| LLQC 2 | 10620 |  |  | 3845 |  |  | 492500 |  |  | 110918 |  |  | 30023 |  |  | 31920 |  |  | 2047727 |  |  | 485242
| LLQC 3 | 10021 |  |  | 3465 |  |  | 494696 |  |  | 110960 |  |  | 26804 |  |  | 30244 |  |  | 2007112 |  |  | 487601
| LLQC 4 | 13795 |  |  | 4470 |  |  | 547224 |  |  | 120040 |  |  | 37241 |  |  | 33809 |  |  | 2226191 |  |  | 538937
| LLQC 5 | 13451 |  |  | 3982 |  |  | 495316 |  |  | 101115 |  |  | 30005 |  |  | 30955 |  |  | 1979932 |  |  | 474180
| LLQC 6 | 9759 |  |  | 6041 |  |  | 487103 |  |  | 113936 |  |  | 32039 |  |  | 31677 |  |  | 2048275 |  |  | 501056
| Low QC | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| LQC 1 | 18272 |  |  | 6091 |  |  | 789935 |  |  | 177129 |  |  | 27332 |  |  | 29409 |  |  | 1986585 |  |  | 460224
| LQC 2 | 20058 |  |  | 7167 |  |  | 889230 |  |  | 189149 |  |  | 32748 |  |  | 32520 |  |  | 2184799 |  |  | 503673
| LQC 3 | 18306 |  |  | 6861 |  |  | 842562 |  |  | 189465 |  |  | 31600 |  |  | 30403 |  |  | 2019117 |  |  | 480567
| LQC 4 | 21939 |  |  | 6689 |  |  | 937846 |  |  | 218153 |  |  | 32415 |  |  | 36459 |  |  | 2266727 |  |  | 533229
| LQC 5 | 20926 |  |  | 7223 |  |  | 964829 |  |  | 203445 |  |  | 28780 |  |  | 30415 |  |  | 2047230 |  |  | 483373
| LQC 6 | 17062 |  |  | 5830 |  |  | 840714 |  |  | 199064 |  |  | 27805 |  |  | 27161 |  |  | 2064870 |  |  | 480324
| Middle QC | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| MQC 1 | 37844 |  |  | 14301 |  |  | 1766023 |  |  | 401018 |  |  | 30463 |  |  | 31604 |  |  | 2060810 |  |  | 508233
| MQC 2 | 35856 |  |  | 13456 |  |  | 1856622 |  |  | 401716 |  |  | 29396 |  |  | 31190 |  |  | 2225090 |  |  | 491845
| MQC 3 | 43214 |  |  | 14017 |  |  | 2028104 |  |  | 432891 |  |  | 32712 |  |  | 28145 |  |  | 2254268 |  |  | 532769
| MQC 4 | 44276 |  |  | 18527 |  |  | 2120466 |  |  | 493959 |  |  | 33650 |  |  | 33273 |  |  | 2368978 |  |  | 549695
| MQC 5 | 43059 |  |  | 15711 |  |  | 1985871 |  |  | 458028 |  |  | 33334 |  |  | 33182 |  |  | 2288202 |  |  | 530820
| MQC 6 | 38007 |  |  | 13836 |  |  | 1937919 |  |  | 401511 |  |  | 31905 |  |  | 32457 |  |  | 2231867 |  |  | 516020
| High QC | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| HQC 1 | 64813 |  |  | 45167 |  |  | 3017807 |  |  | 1180672 |  |  | 31594 |  |  | 29708 |  |  | 2236298 |  |  | 534189
| HQC 2 | 68368 |  |  | 41662 |  |  | 3064744 |  |  | 1163501 |  |  | 30916 |  |  | 30366 |  |  | 2225457 |  |  | 521876
| HQC 3 | 66779 |  |  | 37226 |  |  | 3106376 |  |  | 1172142 |  |  | 33354 |  |  | 31552 |  |  | 2254914 |  |  | 528732
| HQC 4 | 62964 |  |  | 38990 |  |  | 3025734 |  |  | 1112781 |  |  | 33496 |  |  | 25891 |  |  | 2202660 |  |  | 520407
| HQC 5 | 55134 |  |  | 39604 |  |  | 2744457 |  |  | 1081394 |  |  | 30676 |  |  | 30882 |  |  | 2224324 |  |  | 523952
| HQC 6 | 69623 |  |  | 40801 |  |  | 3110702 |  |  | 1215737 |  |  | 35094 |  |  | 32451 |  |  | 2281761 |  |  | 543511
| Highest Level QC | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| HLQC 1 | 84913 |  |  | 63244 |  |  | 3991513 |  |  | 1792740 |  |  | 30939 |  |  | 27479 |  |  | 2239406 |  |  | 517921
| HLQC 2 | 89391 |  |  | 64641 |  |  | 3962891 |  |  | 1800145 |  |  | 31681 |  |  | 32850 |  |  | 2196319 |  |  | 521081
| HLQC 3 | 84947 |  |  | 63863 |  |  | 4082627 |  |  | 1771321 |  |  | 31054 |  |  | 29058 |  |  | 2217698 |  |  | 521412
| HLQC 4 | 87260 |  |  | 68537 |  |  | 3908791 |  |  | 1899794 |  |  | 27790 |  |  | 30128 |  |  | 2098913 |  |  | 498650
| HLQC 5 | 83373 |  |  | 59966 |  |  | 3746561 |  |  | 1688520 |  |  | 29768 |  |  | 33326 |  |  | 2219981 |  |  | 525140
| HLQC 6 | 91198 |  |  | 63490 |  |  | 4100450 |  |  | 1942843 |  |  | 31703 |  |  | 35525 |  |  | 2254036 |  |  | 563541
| Blank QC | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| Matrix Blank | 36191 | 39659 | 42708 | 14575 | 16139 | 16288 | 1363055 | 1658024 | 2141597 | 334884 | 391080 | 464273 | 128 | 116 | 162 | 520 | 100 | 97 | 360 | 7037 | 7216 | 43 | 1317 | 1813
| Total Blank | 7860 | 7077 | 6190 | 1349 | 2023 | 1165 | 7599 | 11222 | 10432 | 2725 | 2970 | 2444 | 165 | 92 | 207 | 26 | 41 | 49 | 1145 | 2074 | 2205 | 237 | 123 | 55

## Raw Peak Areas
Sample_Name | Cer d18:1/16:0(SphB) | D7-Cer d18:1/16:0(SphB) | Cer d18:1/18:0(SphB) | D7-Cer d18:1/18:0(SphB) | Cer d18:1/24:0(SphB) | D7-Cer d18:1/24:0(SphB) | Cer d18:1/24:1(SphB) | D7-Cer d18:1/24:1(SphB)
highTAG-1 | 46711 | 28797 | 17256 | 30748 | 2416874 | 2141412 | 509264 | 504071
highTAG-2 | 50354 | 32136 | 15625 | 33516 | 2788730 | 2339327 | 550362 | 540830
highTAG-3 | 49426 | 30379 | 17887 | 28767 | 2977561 | 2395142 | 650927 | 563199
highTAG-4 | 50028 | 36932 | 18219 | 30941 | 2854402 | 2421923 | 611641 | 572645
highTAG-5 | 52070 | 36339 | 17945 | 35858 | 2843677 | 2410607 | 618002 | 595542
highTAG-6 | 42522 | 32178 | 14692 | 30740 | 2487251 | 2267440 | 501197 | 530670
HLQC 1 | 84913 | 30939 | 63244 | 27479 | 3991513 | 2239406 | 1792740 | 517921
HLQC 2 | 89391 | 31681 | 64641 | 32850 | 3962891 | 2196319 | 1800145 | 521081
HLQC 3 | 84947 | 31054 | 63863 | 29058 | 4082627 | 2217698 | 1771321 | 521412
HLQC 4 | 87260 | 27790 | 68537 | 30128 | 3908791 | 2098913 | 1899794 | 498650
HLQC 5 | 83373 | 29768 | 59966 | 33326 | 3746561 | 2219981 | 1688520 | 525140
HLQC 6 | 91198 | 31703 | 63490 | 35525 | 4100450 | 2254036 | 1942843 | 563541
HQC 1 | 64813 | 31594 | 45167 | 29708 | 3017807 | 2236298 | 1180672 | 534189
HQC 2 | 68368 | 30916 | 41662 | 30366 | 3064744 | 2225457 | 1163501 | 521876
HQC 3 | 66779 | 33354 | 37226 | 31552 | 3106376 | 2254914 | 1172142 | 528732
HQC 4 | 62964 | 33496 | 38990 | 25891 | 3025734 | 2202660 | 1112781 | 520407
HQC 5 | 55134 | 30676 | 39604 | 30882 | 2744457 | 2224324 | 1081394 | 523952
HQC 6 | 69623 | 35094 | 40801 | 32451 | 3110702 | 2281761 | 1215737 | 543511
LLQC 1 | 12527 | 33335 | 5074 | 28322 | 474661 | 1889067 | 110493 | 495868
LLQC 2 | 10620 | 30023 | 3845 | 31920 | 492500 | 2047727 | 110918 | 485242
LLQC 3 | 10021 | 26804 | 3465 | 30244 | 494696 | 2007112 | 110960 | 487601
LLQC 4 | 13795 | 37241 | 4470 | 33809 | 547224 | 2226191 | 120040 | 538937
LLQC 5 | 13451 | 30005 | 3982 | 30955 | 495316 | 1979932 | 101115 | 474180
LLQC 6 | 9759 | 32039 | 6041 | 31677 | 487103 | 2048275 | 113936 | 501056
LQC 1 | 18272 | 27332 | 6091 | 29409 | 789935 | 1986585 | 177129 | 460224
LQC 2 | 20058 | 32748 | 7167 | 32520 | 889230 | 2184799 | 189149 | 503673
LQC 3 | 18306 | 31600 | 6861 | 30403 | 842562 | 2019117 | 189465 | 480567
LQC 4 | 21939 | 32415 | 6689 | 36459 | 937846 | 2266727 | 218153 | 533229
LQC 5 | 20926 | 28780 | 7223 | 30415 | 964829 | 2047230 | 203445 | 483373
LQC 6 | 17062 | 27805 | 5830 | 27161 | 840714 | 2064870 | 199064 | 480324
Matrix Blank1 | 36191 | 128 | 14575 | 520 | 1363055 | 360 | 334884 | 43
Matrix Blank2 | 39659 | 116 | 16139 | 100 | 1658024 | 7037 | 391080 | 1317
Matrix Blank3 | 42708 | 162 | 16288 | 97 | 2141597 | 7216 | 464273 | 1813
MQC 1 | 37844 | 30463 | 14301 | 31604 | 1766023 | 2060810 | 401018 | 508233
MQC 2 | 35856 | 29396 | 13456 | 31190 | 1856622 | 2225090 | 401716 | 491845
MQC 3 | 43214 | 32712 | 14017 | 28145 | 2028104 | 2254268 | 432891 | 532769
MQC 4 | 44276 | 33650 | 18527 | 33273 | 2120466 | 2368978 | 493959 | 549695
MQC 5 | 43059 | 33334 | 15711 | 33182 | 1985871 | 2288202 | 458028 | 530820
MQC 6 | 38007 | 31905 | 13836 | 32457 | 1937919 | 2231867 | 401511 | 516020
Pooled QC1 | 38246 | 29959 | 13443 | 27000 | 1551424 | 1793495 | 396239 | 474226
Pooled QC2 | 40509 | 34507 | 14002 | 29988 | 1949598 | 2183773 | 412481 | 484174
Pooled QC3 | 40820 | 34278 | 13357 | 33370 | 1810844 | 2256232 | 403888 | 560721
Pooled QC4 | 46313 | 33539 | 18487 | 31037 | 2287324 | 2422867 | 479790 | 578795
Pooled QC5 | 38115 | 30704 | 15951 | 31820 | 1605426 | 1866998 | 402816 | 496860
Pooled QC6 | 38544 | 35900 | 16604 | 33342 | 2041234 | 2328007 | 443853 | 523750
Pooled QC7 | 35408 | 29706 | 13664 | 26692 | 1553239 | 1807241 | 355228 | 444551
Pooled QC8 | 33934 | 29615 | 13895 | 28955 | 1677049 | 2174219 | 361580 | 489495
Pooled QC9 | 48003 | 33705 | 18784 | 31207 | 2358656 | 2311464 | 526378 | 547397
Pooled QC10 | 40582 | 32791 | 14045 | 29429 | 1942911 | 2270304 | 434861 | 517328
SRM 1950-1 | 44245 | 36299 | 17212 | 35171 | 1921981 | 2472061 | 440438 | 564374
SRM 1950-2 | 44637 | 29554 | 16696 | 31246 | 1784821 | 2201377 | 405819 | 517169
SRM 1950-3 | 43442 | 29507 | 13025 | 34734 | 1759030 | 2206640 | 406130 | 532053
SRM 1950-4 | 48168 | 32304 | 18561 | 35840 | 1837212 | 2280309 | 424716 | 542414
SRM 1950-5 | 42300 | 33293 | 15355 | 29067 | 1769185 | 2238350 | 416176 | 524312
SRM 1950-6 | 38519 | 34493 | 15311 | 30469 | 1732838 | 2229946 | 393455 | 534017
STD-A1-1 | 724406 | 30597 | 822340 | 26864 | 31388293 | 1844375 | 21769283 | 445172
STD-A1-2 | 390646 | 30951 | 440824 | 28152 | 15413643 | 1784665 | 10818415 | 425327
STD-A1-3 | 48178 | 27805 | 42079 | 24335 | 1546845 | 1622089 | 1088706 | 422642
STD-A1-4 | 17251 | 29405 | 12264 | 29536 | 382173 | 1652182 | 271520 | 409161
STD-A1-5 | 12794 | 28820 | 7603 | 27196 | 216567 | 1718484 | 158971 | 448197
STD-A1-6 | 11036 | 28173 | 5110 | 29887 | 211296 | 2133311 | 126186 | 444907
STD-A2-1 | 728437 | 30326 | 829620 | 30349 | 31084632 | 1892690 | 22531890 | 444661
STD-A2-2 | 389209 | 29823 | 449607 | 30081 | 15563676 | 1824027 | 11083859 | 435013
STD-A2-3 | 44670 | 29337 | 42178 | 26080 | 1599073 | 1658669 | 1081727 | 421416
STD-A2-4 | 18222 | 27749 | 11750 | 26837 | 396568 | 1722675 | 290164 | 422027
STD-A2-5 | 11825 | 29465 | 7054 | 27967 | 220208 | 1668250 | 153444 | 428299
STD-A2-6 | 13144 | 30093 | 6449 | 28147 | 202801 | 1739839 | 131513 | 437046
STD-A3-1 | 754521 | 32711 | 870966 | 29993 | 33053390 | 1971422 | 22931371 | 470656
STD-A3-2 | 391892 | 30976 | 453107 | 28591 | 15954809 | 1819687 | 11133527 | 442515
STD-A3-3 | 48167 | 29038 | 46069 | 27766 | 1607324 | 1737367 | 1167605 | 425688
STD-A3-4 | 17381 | 31556 | 12946 | 27936 | 405108 | 1683894 | 290426 | 427899
STD-A3-5 | 13794 | 27727 | 7760 | 26705 | 228492 | 1699348 | 154928 | 419862
STD-A3-6 | 12380 | 29619 | 6119 | 29378 | 209534 | 1736049 | 133244 | 445876
STD-B1-1 | 697701 | 30838 | 756788 | 25899 | 33965913 | 2193865 | 22471750 | 511335
STD-B1-2 | 384727 | 30329 | 414150 | 24611 | 15241256 | 1838899 | 10697622 | 452155
STD-B1-3 | 46275 | 27522 | 46513 | 27199 | 1628421 | 1776421 | 1149647 | 420442
STD-B1-4 | 12558 | 25009 | 6838 | 27040 | 220876 | 1740185 | 149437 | 413349
STD-B1-5 | 18383 | 29657 | 11600 | 28591 | 403433 | 1731026 | 280683 | 430054
STD-B1-6 | 11001 | 30042 | 4438 | 27307 | 178636 | 1784794 | 121390 | 430917
STD-B2-1 | 649165 | 27987 | 753029 | 24647 | 28696806 | 1809433 | 20103282 | 445800
STD-B2-2 | 365829 | 27912 | 409424 | 26637 | 15413567 | 1812081 | 10494584 | 401333
STD-B2-3 | 51533 | 29079 | 48775 | 28530 | 1725695 | 1784162 | 1147044 | 428852
STD-B2-4 | 12035 | 26997 | 8405 | 26566 | 228846 | 1799683 | 151401 | 421962
STD-B2-5 | 17117 | 29334 | 12039 | 26984 | 415183 | 1765414 | 284485 | 434424
STD-B2-6 | 11898 | 31229 | 5674 | 28048 | 185452 | 1832784 | 123272 | 444288
STD-B3-1 | 677983 | 29334 | 757124 | 26678 | 30323264 | 1887443 | 20364385 | 453160
STD-B3-2 | 363987 | 25946 | 411886 | 26282 | 15144521 | 1818270 | 10498060 | 432137
STD-B3-3 | 45502 | 30978 | 49031 | 27233 | 1704835 | 1799487 | 1186478 | 439920
STD-B3-4 | 12705 | 27566 | 8159 | 26827 | 229881 | 1807762 | 156017 | 420152
STD-B3-5 | 17246 | 28886 | 12261 | 27501 | 428355 | 1791029 | 281563 | 429107
STD-B3-6 | 11398 | 30560 | 5454 | 26435 | 187931 | 1903322 | 133926 | 454850
T1D-1 | 37517 | 35003 | 18333 | 35115 | 1809812 | 2346411 | 369843 | 564123
T1D-2 | 43569 | 37966 | 18092 | 38082 | 2012488 | 2587083 | 409198 | 617575
T1D-3 | 33087 | 30860 | 17436 | 28361 | 1590200 | 2128985 | 316825 | 480111
T1D-4 | 36522 | 28902 | 14270 | 31684 | 1500750 | 2090703 | 367445 | 486952
T1D-5 | 33894 | 29464 | 21238 | 28304 | 1548276 | 2002024 | 332979 | 488012
T1D-6 | 40723 | 31653 | 14266 | 29190 | 1767685 | 2332066 | 359392 | 538678
Total Blank1 | 7860 | 165 | 1349 | 26 | 7599 | 1145 | 2725 | 237
Total Blank2 | 7077 | 92 | 2023 | 41 | 11222 | 2074 | 2970 | 123
Total Blank3 | 6190 | 207 | 1165 | 49 | 10432 | 2205 | 2444 | 55
Total Blank4 | 6066 | 274 | 1265 | 72 | 10182 | 1620 | 2886 | 83
YOUNG AA-1 | 28041 | 31219 | 10574 | 33920 | 1328918 | 2291397 | 276930 | 523630
YOUNG AA-2 | 26709 | 30320 | 8824 | 32788 | 1137060 | 2143086 | 244292 | 514448
YOUNG AA-3 | 28079 | 36744 | 10420 | 39098 | 1323930 | 2447512 | 244035 | 578695
YOUNG AA-4 | 25742 | 29971 | 7293 | 30824 | 1159612 | 2140221 | 251011 | 498888
YOUNG AA-5 | 29147 | 33540 | 10973 | 32638 | 1301467 | 2343470 | 270824 | 529795
YOUNG AA-6 | 19524 | 30981 | 5998 | 29008 | 915012 | 2093059 | 180648 | 492264