## Participant Information
Participant Information
Participant Information |  | Group Name
|  | Lab code | 2b
| Institution | Institution name
Street
City
Postal code
Country
| Contact person | Name
Address
Email
Phone

## Summary
Summary
Summary | This protocol describes the extraction and LC-MRM procedure for absolute quantification of Cer d18:1/16:0, Cer d18:1/18:0, Cer d18:1/24:0 and Cer d18:1/24:1 in human plasma according to Kauhanen et al, Anal Bioanal Chem (2016) 408:3475–3483.
A calibration curve is built by injecting several dilutions of the non-labelled standards mixed with a fixed concentration of labelled standards. Each calibration point is prepared and analysed in triplicates (n=3).
To establish the endogenous concentration of the four selected ceramides in the study samples (NIST SRM1950, NIST high TAG, NIST T1D, and NIST young A-A) each sample is extracted and analyzed in six replicates (n =6).
Prepare the pooled QC by mixing 100 mL each of NIST SRM1950, NIST high TAG, NIST T1D, and NIST young AA (total 400 mL).
We suggest proceeding first to measure the ceramides concentration in the four study samples and then continue with the intra-assay validation to avoid wasting material.

## Extraction - Overview
Extraction Overview
Extraction Overview | Table 1 - Extraction - Materials
Chemicals, solvents and materials
Ethyl acetate (EtAc)
2-propanol (IPA)
5% Bovine Serum Albumin (BSA)
96-well plates/tubes
Sealing foil
2 mL DeepWell plates/tubes/LC vials
Table 2 - Extraction - Sample Types
| Sample type | Components
| Pooled Quality Control samples (QC) | 10 µL pooled QC |  |  | 20 µL Labelled IS |  | 570 µL EtAc:IPA 2:8 (v/v)
| Matrix Blank (MB) | 10 µL Matrix (pooled QC) only |  |  | - |  | 590 µL EtAc:IPA 2:8
| Study Samples (S) | 10 µL Human plasma (NIST SRM1950, high TAG, T1D, young AA) |  |  | 20 µL Labelled IS |  | 570 µL EtAc:IPA 2:8
| Total Blank (TB) | 10 µL 5% BSA |  |  | - |  | 590 µL EtAc:IPA 2:8
| Calibration line standards (STD) | 10 µL 5% BSA + 20 µL non labelled STD mixture (Table 5) |  |  | 20 µL Labelled IS |  | 550 µL EtAc:IPA 2:8
Table 3 - Extraction - Labelled Internal Standard Mixture Solution (provided)
| Labelled IS mixture in EtAc:IPA 2:8 (v/v) | Concentration |  |  |  |  | Unit | Vendor
| D7-Cer d18:1/16:0 | 0.125 |  |  |  |  | pmol/µL [µM] | Avanti Polar Lipids
| D7-Cer d18:1/18:0 | 0.050
| D7-Cer d18:1/24:0 | 1500
| D7-Cer d18:1/24:1 | 0.500
Table 4 - Extraction - Non-Labelled Standards - for STD1 Mixture Solution (provided)
| Non-labelled standards in EtAc:IPA 2:8 (v/v) | Concentration |  |  |  |  | Unit | Vendor
| Cer d18:1/16:0 | 2.0 |  |  |  |  | pmol/µL [µM] | Avanti Polar Lipids
| Cer d18:1/18:0 | 2.0
| Cer d18:1/24:0 | 20.0
| Cer d18:1/24:1 | 20.0
Table 5 - Serial dilutions of Non-Labelled Standards from STD1 Mixture Solution for calibration line
| Add 20 µL of | STD1 Conc. | STD2 Conc. | STD3 Conc. | STD4 Conc. | STD5 Conc. | STD6 Conc. | Unit
| Cer d18:1/16:0 | 2.0 | 1.0 | 0.10 | 0.02 | 0.01 | 0.008 | pmol/µL [µM]
| Cer d18:1/18:0 | 2.0 | 1.0 | 0.10 | 0.02 | 0.01 | 0.008
| Cer d18:1/24:0 | 20.0 | 10.0 | 1.0 | 0.2 | 0.1 | 0.08
| Cer d18:1/24:1 | 20.0 | 10.0 | 1.0 | 0.2 | 0.1 | 0.08

## Extraction - Procedure
Extraction Procedure
Extraction Procedure | Extraction Procedure
| Step | Description | Reference Table
| 1 | Prepare different dilutions of STD1 by diluting the provided mixture (Extraction - Overview, Table 4) with EtAc:IPA 2:8 as suggested in table | Extraction - Overview'!B34
| 2 | Thaw plasma on ice (+4°C) and bring to room temperature prior to extraction. Mix by pipetting before use.
| 3 | Prepare samples in plates/tubes manually or with a robot, according to specifications in column C of table | Extraction - Overview'!C14
| 4 | Add 20 µL of labelled internal standard mixture to samples according to column F of table | Extraction - Overview'!F14
| 5 | For the calibration line, add 20 µL of different dilutions of the non-labelled mixture (STD1-6).
| 6 | Add 590 µL of ethyl acetate:isopropanol (2:8, vol/vol) to Total Blank (TB) and Matrix Blank (MB); 550 µL to calibration line (STD); 570 µL to study samples and QCs according to column H of table | Extraction - Overview'!H14
| 7 | Mix samples 10 min by pipetting or vortexing
| 8 | Centrifuge samples for 10 min at 3000 x gat room temperature
| 9 | Transfer 50 µL of the clear supernatant to MS-vials or MS plates, close/seal, store samples at -20 °C prior to LC-MRM analysis.

## LC-MRM - Overview
LC-MRM Overview
LC-MRM Overview | Table 6 Chemicals and solvents
Ammonium formate
Formic acid
Ultra-pure water
Acetonitrile
Table 7 LC-MRM LC-MS Details
| Property | Specifications
| Mass spectrometry | QQQ
| LC | UPLC
| Injection volume (indicative) | 3 µL
| Weak wash | Water
| Strong wash | Acetonitrile
| Pre-column | Phenomenex SecurityGuard™ ULTRA Cartridge UHPLC C18 2.1mm ID Column
| Column | Phenomenex Kinetex C18, 2.1 × 50 mm id., 2.6 µm particle size
| Column temperature | 60 °C
| Mobile phase A | 10 mM ammonium acetate in water with 0.1% formic acid
| Mobile phase B | 10 mM ammonium acetate in acetonitrile:2-propanol:water (4:3:0.2, vol/vol/vol) with 0.1% formic acid
Table 8 LC-MRM LC gradient
| Time (minutes) | A% | B% |  | Flow (µL/min)
| 0.00 | 15 | 85 |  | 500
| 0.50 | 15 | 85 |  | 500
| 1.50 | 0 | 100 |  | 500
| 4.00 | 0 | 100 |  | 500
| 4.10 | 15 | 85 |  | 500
| 5.00 | 15 | 85 |  | 500
Table 9 LC-MRM MS conditions (EXAMPLE: this only applies to Sciex QTRAP, please adapt)
| Polarity mode | Positive
| Ion spray | 5000 V
| Curtain gas | 30 psi
| Source temperature | 300 °C
| Gas 1 | 50 psi
| Gas 2 | 35 psi
| DP | 30 V
| EP | 10 V
| Collision cell exit potential | 20 V
| Collision energy (CE) | 40 eV
Table 10 LC-MRM Mass transitions used for absolute quantification. On different systems, RT and CE might be different!
| Analyte | Mass transitions | IS | Mass transitions | RT (min) | CE
| Cer d18:1/16:0 | 538.5 – 264.3 | D7- Cer d18:1/16:0 | 545.6 - 271.3 | 1.9-2.0 | 40
C34H68NO3 – C18H34N
| Cer d18:1/18:0 | 566.6 – 264.3 | D7- Cer d18:1/18:0 | 573.6 – 271.3 | 2.1-2.2 | 40
C36H72NO3 – C18H34N
| Cer d18:1/24:0 | 650.6 – 264.3 | D7- Cer d18:1/24:0 | 657.7 – 271.3 | 2.4-2.5 | 40
C42H84NO3 – C18H34N
| Cer d18:1/24:1 | 648.6 – 264.3 | D7- Cer d18:1/24:1 | 655.7 – 271.3 | 2.3-2.4 | 40
C42H82NO3 – C18H34N
Table 11 Data processing
| Processing software | Analyst 1.6.3

## LC-MRM - Sample Sequence
Sample Sequence
Sample Sequence | Sample Sequence 1
| Number | Sample Name | Remarks
| 1 | Total Blank 1
| 2 | Matrix blank
| 3 | Std 6a
| 4 | Std 6b
| 5 | Std 6c
| 6 | Std 5a
| 7 | Std 5b
| 8 | Std 5c
| 9 | Std 4a
| 10 | Std 4b
| 11 | Std 4c
| 12 | Std 3a
| 13 | Std 3b
| 14 | Std 3c
| 15 | Std 2a
| 16 | Std 2b
| 17 | Std 2c
| 18 | Std 1a
| 19 | Std 1b
| 20 | Std 1c
| 21 | pooled QC
| 22 | pooled QC
| 23 | SRM 1950-1
| 24 | SRM 1950-2
| 25 | SRM 1950-3
| 26 | SRM 1950-4
| 27 | SRM 1950-5
| 28 | SRM 1950-6
| 29 | pooled QC
| 30 | pooled QC
| 31 | T1D-1
| 32 | T1D-2
| 33 | T1D-3
| 34 | T1D-4
| 35 | T1D-5
| 36 | T1D-6
| 37 | total blank 2
| 38 | pooled QC
| 39 | pooled QC
| 40 | Young AA-1
| 41 | Young AA-2
| 42 | Young AA-3
| 43 | Young AA-4
| 44 | Young AA-5
| 45 | Young AA-6
| 46 | Pooled QC
| 47 | Pooled QC
| 48 | hTAG-1
| 49 | hTAG-2
| 50 | hTAG-3
| 51 | hTAG-4
| 52 | hTAG-5
| 53 | hTAG-6
| 54 | Pooled QC
| 55 | Pooled QC
| 56 | Std 6a
| 57 | Std 6b
| 58 | Std 6c
| 59 | Std 5a
| 60 | Std 5b
| 61 | Std 5c
| 62 | Std 4a
| 63 | Std 4b
| 64 | Std 4c
| 65 | Std 3a
| 66 | Std 3b
| 67 | Std 3c
| 68 | Std 2a
| 69 | Std 2b
| 70 | Std 2c
| 71 | Std 1a
| 72 | Std 1b
| 73 | Std 1c
| 74 | Total Blank 1
| 75 | Matrix blank

## LC-MRM - Intra Assay QC Proc
Intra Assay Procedure
Intra Assay Procedure | Quality Controls for Intra-Assay Variation
| Step | Description | Reference Table
| 1 | Prepare the Pooled NIST QC sample containing equal volumes of the four NIST reference plasma you received.
| 2 | The pooled QC representing middle quality control (MQC), low QC (LQC), and lower limit of quantification QC (LLOQ) is prepared by diluting MQC (e.g. two- and three-fold) with water.
| 3 | High QC (HQC) and upper limit of quantification (ULOQ) QC are prepared by spiking in 10 µL and 20 µL respectively of STD3 endogenous standards from table | Extraction - Overview'!B34
| 4 | Internal standards are added to all QC samples.
Intra-Assay Variation
| 6 | The precision and accuracy of the assay is determined for each of the four ceramides measured in LLOQ, LQC, MQC, HQC, and HLOQ.
| 7 | Intra-assay variation, precision, and accuracy is calculated for each of the ceramide in replicates of six (n = 6) at each QC concentration independently on three separate occasions.
| 8 | The intra-assay precision (percentage coefficient variance, %CV) and accuracy (percentage accuracy, %Accuracy) are calculated from the nominal concentrations according to the formulas reported in corresponding sheets and tables of this document.
| 9 | Recovery estimation is not necessary for the published method but only for different ones. Stability estimation is not necessary for the published method but only for different ones. Calibration line is prepared in 5% BSA.

## LC-MRM - Intra Assay QC Res
INTRA ASSAY - QC RESULTS
INTRA ASSAY - QC RESULTS | Table 12 LC-MRM Calibration Line 1 Data (Beginning). Please report the values of each of the three replicates separately for each compound and standard.
| Calibration Line 1 | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| STD 1 | 10600000 | 12700000 | 11400000 | 21500000 | 18800000 | 18200000 | 209000000 | 193000000 | 197000000 | 190000000 | 175000000 | 174000000 | 507000 | 624000 | 580000 | 890000 | 754000 | 686000 | 18700000 | 17000000 | 17200000 | 6140000 | 5600000 | 5370000
| STD 2 | 5130000 | 5280000 | 5500000 | 9690000 | 9680000 | 9600000 | 119000000 | 114000000 | 116000000 | 108000000 | 105000000 | 106000000 | 501000 | 514000 | 516000 | 767000 | 747000 | 748000 | 18900000 | 19200000 | 18800000 | 6380000 | 6150000 | 5930000
| STD 3 | 563000 | 563000 | 577000 | 1080000 | 1030000 | 1080000 | 14900000 | 14000000 | 13800000 | 13100000 | 12400000 | 13000000 | 494000 | 534000 | 508000 | 774000 | 787000 | 810000 | 21500000 | 20400000 | 20000000 | 6760000 | 6750000 | 6600000
| STD 4 | 139000 | 120000 | 109000 | 218000 | 213000 | 209000 | 3170000 | 2920000 | 3050000 | 2670000 | 2480000 | 2660000 | 617000 | 582000 | 484000 | 770000 | 790000 | 782000 | 21100000 | 20800000 | 20800000 | 6760000 | 6490000 | 6680000
| STD 5 | 67200 | 58500 | 58600 | 98700 | 102000 | 103000 | 1610000 | 1570000 | 1610000 | 1260000 | 1280000 | 1360000 | 577000 | 479000 | 473000 | 775000 | 727000 | 715000 | 20700000 | 20600000 | 20600000 | 6410000 | 6460000 | 6360000
| STD 6 | 65000 | 57100 | 48100 | 98600 | 88100 | 87300 | 1410000 | 1230000 | 1340000 | 1110000 | 1020000 | 1050000 | 687000 | 506000 | 485000 | 832000 | 804000 | 793000 | 21800000 | 20800000 | 20800000 | 7000000 | 6810000 | 6750000
Table 13 LC-MRM Calibration Line 2 Data (End). Please report the values of each of the three replicates separately for each compound and standard.
| Calibration Line 2 | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| STD 1 | 7280000 | 9230000 | 8690000 | 8210000 | 8280000 | 8150000 | 107000000 | 113000000 | 108000000 | 97200000 | 96200000 | 96600000 | 318000 | 422000 | 391000 | 324000 | 334000 | 319000 | 9020000 | 8720000 | 8480000 | 2740000 | 2670000 | 2590000
| STD 2 | 3800000 | 3610000 | 3850000 | 4320000 | 4000000 | 4180000 | 58300000 | 55000000 | 56400000 | 55000000 | 48700000 | 54000000 | 329000 | 311000 | 323000 | 352000 | 322000 | 310000 | 9010000 | 8590000 | 8510000 | 2900000 | 2780000 | 2790000
| STD 3 | 392000 | 418000 | 406000 | 472000 | 449000 | 452000 | 6470000 | 6210000 | 6200000 | 6210000 | 5740000 | 5680000 | 307000 | 352000 | 323000 | 352000 | 332000 | 322000 | 9200000 | 9090000 | 8810000 | 3100000 | 2920000 | 2850000
| STD 4 | 109000 | 89400 | 78400 | 92200 | 84300 | 91200 | 1360000 | 1260000 | 1270000 | 1190000 | 1110000 | 1130000 | 393000 | 360000 | 298000 | 344000 | 342000 | 341000 | 8940000 | 9030000 | 8910000 | 2890000 | 2950000 | 2940000
| STD 5 | 39000 | 43300 | 40000 | 39600 | 50400 | 41500 | 697000 | 698000 | 689000 | 571000 | 557000 | 596000 | 365000 | 306000 | 304000 | 318000 | 332000 | 326000 | 9090000 | 9090000 | 8800000 | 2910000 | 2930000 | 2770000
| STD 6 | 40200 | 30500 | 31800 | 39400 | 32400 | 39100 | 592000 | 552000 | 573000 | 484000 | 431000 | 478000 | 386000 | 301000 | 304000 | 321000 | 331000 | 333000 | 9180000 | 9030000 | 9150000 | 2880000 | 2940000 | 2910000
Table 14 LC-MRM Sample Data. Please report the values of each of the three replicates separately for each compound and standard.
| NIST SRM | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| SRM 1 | 517000 | 496000 | 493000 | 425000 | 421000 | 403000 | 16000000 | 16000000 | 15800000 | 5020000 | 4880000 | 4870000 | 338000 | 330000 | 332000 | 714000 | 713000 | 682000 | 19100000 | 18200000 | 18600000 | 5500000 | 5510000 | 5580000
| SRM 2 | 464000 | 507000 | 505000 | 410000 | 399000 | 383000 | 15000000 | 15300000 | 14900000 | 4690000 | 4720000 | 4600000 | 341000 | 351000 | 341000 | 702000 | 715000 | 723000 | 17500000 | 18200000 | 17500000 | 5370000 | 5390000 | 5330000
| SRM 3 | 476000 | 499000 | 467000 | 400000 | 410000 | 425000 | 14600000 | 15300000 | 15100000 | 4560000 | 4630000 | 4470000 | 328000 | 338000 | 333000 | 708000 | 737000 | 720000 | 17500000 | 18500000 | 18000000 | 5240000 | 5410000 | 5470000
| SRM 4 | 471000 | 464000 | 474000 | 399000 | 401000 | 392000 | 14700000 | 14200000 | 15000000 | 4350000 | 4340000 | 4410000 | 306000 | 317000 | 319000 | 694000 | 679000 | 689000 | 17000000 | 16900000 | 16700000 | 5170000 | 5090000 | 4970000
| SRM 5 | 454000 | 464000 | 467000 | 386000 | 390000 | 397000 | 14000000 | 13900000 | 14700000 | 4310000 | 4330000 | 4290000 | 306000 | 306000 | 300000 | 670000 | 691000 | 683000 | 17300000 | 16800000 | 17200000 | 5080000 | 4880000 | 5080000
| SRM 6 | 450000 | 484000 | 465000 | 372000 | 362000 | 409000 | 14000000 | 14500000 | 14900000 | 4300000 | 4200000 | 4240000 | 300000 | 308000 | 320000 | 658000 | 661000 | 689000 | 16300000 | 16600000 | 17300000 | 4760000 | 4980000 | 5000000
| NIST hTAG | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| hTAG 1 | 476000 | 464000 | 472000 | 320000 | 326000 | 333000 | 16000000 | 14900000 | 15300000 | 4250000 | 4300000 | 4050000 | 268000 | 259000 | 256000 | 483000 | 464000 | 471000 | 11900000 | 11600000 | 11300000 | 3250000 | 3330000 | 3200000
| hTAG 2 | 453000 | 440000 | 444000 | 335000 | 315000 | 307000 | 15000000 | 14400000 | 14600000 | 3990000 | 4080000 | 4080000 | 262000 | 248000 | 263000 | 471000 | 444000 | 461000 | 11400000 | 11100000 | 10900000 | 3140000 | 3190000 | 3190000
| hTAG 3 | 458000 | 451000 | 451000 | 322000 | 298000 | 295000 | 15700000 | 14500000 | 14500000 | 3990000 | 4030000 | 3930000 | 257000 | 254000 | 258000 | 475000 | 421000 | 432000 | 11700000 | 11000000 | 10800000 | 3140000 | 3120000 | 3220000
| hTAG 4 | 436000 | 446000 | 445000 | 298000 | 294000 | 320000 | 14200000 | 14900000 | 14600000 | 3920000 | 4060000 | 3970000 | 268000 | 260000 | 267000 | 447000 | 451000 | 445000 | 10700000 | 11200000 | 10900000 | 3130000 | 3240000 | 3210000
| hTAG 5 | 446000 | 452000 | 462000 | 304000 | 296000 | 292000 | 14200000 | 14700000 | 13800000 | 3900000 | 3740000 | 3950000 | 249000 | 241000 | 254000 | 424000 | 449000 | 440000 | 10700000 | 11000000 | 10600000 | 3100000 | 2990000 | 3150000
| hTAG 6 | 471000 | 466000 | 498000 | 310000 | 284000 | 284000 | 14600000 | 14200000 | 14500000 | 3920000 | 3960000 | 3840000 | 246000 | 241000 | 268000 | 437000 | 404000 | 417000 | 10900000 | 10600000 | 10700000 | 3080000 | 3080000 | 3070000
| NIST T1D | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| T1D 1 | 430000 | 448000 | 441000 | 437000 | 423000 | 423000 | 14400000 | 13900000 | 13700000 | 4380000 | 4360000 | 4320000 | 322000 | 331000 | 317000 | 661000 | 629000 | 658000 | 16900000 | 16000000 | 15800000 | 4980000 | 4890000 | 4730000
| T1D 2 | 426000 | 434000 | 428000 | 407000 | 389000 | 389000 | 13000000 | 13200000 | 12900000 | 4030000 | 4010000 | 4090000 | 313000 | 311000 | 311000 | 628000 | 604000 | 581000 | 14400000 | 15000000 | 14600000 | 4480000 | 4550000 | 4530000
| T1D 3 | 410000 | 411000 | 435000 | 400000 | 395000 | 380000 | 13000000 | 13000000 | 12800000 | 3920000 | 3960000 | 3980000 | 284000 | 308000 | 297000 | 592000 | 575000 | 580000 | 15000000 | 14800000 | 14300000 | 4470000 | 4450000 | 4350000
| T1D 4 | 409000 | 411000 | 419000 | 396000 | 398000 | 369000 | 12900000 | 12700000 | 12600000 | 3800000 | 3930000 | 3830000 | 303000 | 297000 | 289000 | 593000 | 581000 | 576000 | 15200000 | 15100000 | 14900000 | 4450000 | 4370000 | 4390000
| T1D 5 | 411000 | 421000 | 408000 | 367000 | 368000 | 374000 | 12800000 | 12000000 | 12100000 | 3950000 | 3840000 | 3740000 | 293000 | 292000 | 307000 | 554000 | 545000 | 540000 | 14000000 | 14000000 | 13900000 | 4250000 | 4080000 | 4150000
| T1D 6 | 410000 | 426000 | 466000 | 386000 | 352000 | 367000 | 12300000 | 12000000 | 12100000 | 3620000 | 3640000 | 3690000 | 296000 | 297000 | 343000 | 552000 | 528000 | 559000 | 13900000 | 13700000 | 13800000 | 4040000 | 4050000 | 4100000
| NIST Young AA | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| yA-A 1 | 311000 | 316000 | 338000 | 188000 | 176000 | 175000 | 7810000 | 7590000 | 7680000 | 1980000 | 2030000 | 1990000 | 324000 | 333000 | 352000 | 521000 | 496000 | 496000 | 12100000 | 11700000 | 11800000 | 3710000 | 3590000 | 3680000
| yA-A 2 | 311000 | 327000 | 312000 | 183000 | 183000 | 184000 | 7920000 | 7910000 | 7920000 | 2030000 | 2030000 | 2010000 | 328000 | 322000 | 328000 | 480000 | 498000 | 486000 | 11400000 | 11600000 | 11500000 | 3570000 | 3560000 | 3580000
| yA-A 3 | 307000 | 310000 | 323000 | 176000 | 173000 | 184000 | 7660000 | 8090000 | 7820000 | 1990000 | 2000000 | 2050000 | 312000 | 315000 | 321000 | 463000 | 480000 | 547000 | 10900000 | 12100000 | 11800000 | 3480000 | 3540000 | 3650000
| yA-A 4 | 331000 | 325000 | 333000 | 181000 | 183000 | 190000 | 7780000 | 7860000 | 8070000 | 2090000 | 1980000 | 2070000 | 323000 | 323000 | 313000 | 486000 | 475000 | 533000 | 11400000 | 11200000 | 12400000 | 3630000 | 3490000 | 3640000
| yA-A 5 | 332000 | 318000 | 321000 | 183000 | 198000 | 169000 | 7620000 | 7810000 | 7160000 | 1980000 | 2120000 | 2010000 | 321000 | 324000 | 330000 | 500000 | 522000 | 461000 | 11200000 | 11700000 | 11000000 | 3550000 | 3670000 | 3610000
| yA-A 6 | 316000 | 316000 | 327000 | 174000 | 174000 | 175000 | 7440000 | 7460000 | 7370000 | 1940000 | 1970000 | 1890000 | 319000 | 314000 | 312000 | 430000 | 469000 | 455000 | 10600000 | 10400000 | 10600000 | 3320000 | 3450000 | 3330000
Table 15 LC-MRM Intra Assay QC Data. Please report the values of each of the three replicates separately for each compound and standard.
| Lowest Level QC | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| LLQC 1 | 225000 | 237000 | 227000 | 102000 | 110000 | 108000 | 4220000 | 4280000 | 4230000 | 1270000 | 1210000 | 1260000 | 368000 | 373000 | 368000 | 415000 | 410000 | 424000 | 10900000 | 10700000 | 10600000 | 3500000 | 3590000 | 3650000
| LLQC 2 | 196000 | 206000 | 199000 | 103000 | 91300 | 92200 | 3780000 | 3810000 | 3720000 | 1170000 | 1130000 | 1120000 | 361000 | 377000 | 387000 | 440000 | 419000 | 409000 | 10700000 | 10600000 | 10600000 | 3730000 | 3530000 | 3590000
| LLQC 3 | 275000 | 291000 | 275000 | 144000 | 139000 | 139000 | 6080000 | 6110000 | 5600000 | 1780000 | 1770000 | 1660000 | 344000 | 370000 | 362000 | 450000 | 474000 | 441000 | 11300000 | 10700000 | 10600000 | 3710000 | 3680000 | 3630000
| LLQC 4 | 263000 | 268000 | 273000 | 131000 | 133000 | 138000 | 4930000 | 4900000 | 4970000 | 1460000 | 1460000 | 1400000 | 366000 | 380000 | 381000 | 433000 | 431000 | 443000 | 10700000 | 10800000 | 10700000 | 3530000 | 3610000 | 3660000
| LLQC 5 | 199000 | 214000 | 209000 | 96400 | 96900 | 104000 | 4200000 | 3910000 | 4160000 | 1180000 | 1160000 | 1240000 | 381000 | 394000 | 383000 | 451000 | 441000 | 439000 | 11400000 | 10500000 | 11100000 | 3710000 | 3610000 | 3660000
| LLQC 6 | 201000 | 195000 | 200000 | 95100 | 92500 | 103000 | 3930000 | 3780000 | 3960000 | 1070000 | 1110000 | 1200000 | 372000 | 359000 | 385000 | 421000 | 420000 | 433000 | 10800000 | 11100000 | 10800000 | 3560000 | 3560000 | 3690000
| Low QC | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| LQC 1 | 303000 | 293000 | 311000 | 165000 | 168000 | 158000 | 6530000 | 6350000 | 6550000 | 1960000 | 2010000 | 2000000 | 356000 | 347000 | 355000 | 470000 | 459000 | 458000 | 11900000 | 11800000 | 11600000 | 3720000 | 3630000 | 3710000
| LQC 2 | 283000 | 279000 | 292000 | 141000 | 147000 | 146000 | 5690000 | 5390000 | 5550000 | 1730000 | 1710000 | 1700000 | 363000 | 372000 | 390000 | 457000 | 476000 | 464000 | 11400000 | 11800000 | 11800000 | 3710000 | 3680000 | 3740000
| LQC 3 | 331000 | 343000 | 330000 | 174000 | 177000 | 170000 | 6420000 | 6530000 | 6360000 | 1870000 | 1880000 | 1930000 | 359000 | 382000 | 357000 | 461000 | 510000 | 473000 | 11600000 | 12400000 | 11900000 | 3600000 | 3730000 | 3590000
| LQC 4 | 303000 | 308000 | 308000 | 156000 | 156000 | 155000 | 6330000 | 6620000 | 6160000 | 1870000 | 1840000 | 1860000 | 356000 | 355000 | 364000 | 440000 | 464000 | 463000 | 11700000 | 11700000 | 11300000 | 3580000 | 3550000 | 3510000
| LQC 5 | 312000 | 332000 | 345000 | 157000 | 164000 | 176000 | 6300000 | 6290000 | 6420000 | 1840000 | 1820000 | 2000000 | 353000 | 354000 | 365000 | 458000 | 434000 | 448000 | 11400000 | 11300000 | 11400000 | 3450000 | 3460000 | 3680000
| LQC 6 | 305000 | 300000 | 307000 | 153000 | 151000 | 155000 | 6210000 | 5910000 | 6080000 | 1830000 | 1810000 | 1800000 | 359000 | 360000 | 353000 | 438000 | 437000 | 427000 | 11300000 | 11200000 | 11500000 | 3550000 | 3560000 | 3540000
| Middle QC | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| MQC 1 | 390000 | 409000 | 403000 | 222000 | 223000 | 230000 | 9230000 | 9470000 | 9330000 | 2660000 | 2520000 | 2650000 | 287000 | 290000 | 303000 | 418000 | 427000 | 407000 | 10400000 | 10600000 | 10700000 | 3120000 | 2980000 | 3150000
| MQC 2 | 400000 | 426000 | 421000 | 240000 | 271000 | 239000 | 10200000 | 10000000 | 10000000 | 2730000 | 2800000 | 2830000 | 290000 | 294000 | 279000 | 402000 | 434000 | 388000 | 10500000 | 10800000 | 10400000 | 2890000 | 3000000 | 3070000
| MQC 3 | 431000 | 422000 | 420000 | 252000 | 245000 | 254000 | 10300000 | 10200000 | 10100000 | 2800000 | 2720000 | 2730000 | 280000 | 278000 | 293000 | 400000 | 404000 | 416000 | 10700000 | 10300000 | 10600000 | 2940000 | 2890000 | 2890000
| MQC 4 | 405000 | 411000 | 413000 | 230000 | 237000 | 225000 | 9690000 | 9710000 | 9130000 | 2610000 | 2670000 | 2620000 | 300000 | 292000 | 304000 | 417000 | 399000 | 418000 | 10700000 | 10700000 | 10300000 | 3030000 | 3030000 | 3030000
| MQC 5 | 409000 | 411000 | 406000 | 229000 | 231000 | 225000 | 9190000 | 9430000 | 9280000 | 2570000 | 2570000 | 2550000 | 277000 | 287000 | 300000 | 381000 | 390000 | 382000 | 9630000 | 10300000 | 10000000 | 2900000 | 2900000 | 2930000
| MQC 6 | 405000 | 401000 | 407000 | 205000 | 224000 | 218000 | 9100000 | 8920000 | 9000000 | 2480000 | 2680000 | 2550000 | 277000 | 283000 | 285000 | 403000 | 383000 | 391000 | 9790000 | 10300000 | 9850000 | 2810000 | 2990000 | 2910000
| High QC | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| HQC 1 | 405000 | 412000 | 394000 | 266000 | 260000 | 243000 | 7310000 | 7140000 | 6730000 | 3240000 | 3010000 | 2980000 | 342000 | 333000 | 334000 | 395000 | 395000 | 363000 | 9360000 | 9170000 | 8870000 | 3140000 | 3080000 | 3060000
| HQC 2 | 410000 | 415000 | 406000 | 260000 | 257000 | 265000 | 7190000 | 7000000 | 7260000 | 3110000 | 3100000 | 3160000 | 338000 | 342000 | 352000 | 355000 | 383000 | 400000 | 8610000 | 8780000 | 8720000 | 2950000 | 3030000 | 3130000
| HQC 3 | 556000 | 527000 | 552000 | 286000 | 290000 | 299000 | 6800000 | 6580000 | 7360000 | 3090000 | 2970000 | 3030000 | 362000 | 352000 | 344000 | 370000 | 354000 | 394000 | 8640000 | 8930000 | 9250000 | 3070000 | 3060000 | 3140000
| HQC 4 | 434000 | 423000 | 424000 | 293000 | 279000 | 283000 | 7670000 | 6980000 | 7120000 | 3400000 | 3190000 | 3030000 | 354000 | 347000 | 351000 | 388000 | 385000 | 391000 | 9560000 | 8740000 | 8880000 | 3220000 | 3020000 | 3100000
| HQC 5 | 433000 | 418000 | 407000 | 280000 | 281000 | 269000 | 7620000 | 7550000 | 7100000 | 3360000 | 3370000 | 3200000 | 344000 | 335000 | 328000 | 409000 | 376000 | 372000 | 9540000 | 9220000 | 8780000 | 3030000 | 3140000 | 2980000
| HQC 6 | 389000 | 387000 | 372000 | 246000 | 237000 | 237000 | 6390000 | 5900000 | 6120000 | 3020000 | 2840000 | 2730000 | 349000 | 341000 | 339000 | 382000 | 367000 | 373000 | 8520000 | 8410000 | 8430000 | 3090000 | 2820000 | 2870000
| Highest Level QC | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| HLQC 1 | 363000 | 368000 | 361000 | 278000 | 266000 | 250000 | 5650000 | 5500000 | 5670000 | 3320000 | 3260000 | 3240000 | 326000 | 355000 | 334000 | 367000 | 352000 | 353000 | 8410000 | 8300000 | 8160000 | 2930000 | 2900000 | 2900000
| HLQC 2 | 371000 | 362000 | 361000 | 263000 | 265000 | 259000 | 5600000 | 5440000 | 5610000 | 3160000 | 3260000 | 3180000 | 348000 | 348000 | 340000 | 370000 | 370000 | 345000 | 8310000 | 8110000 | 8630000 | 2910000 | 2920000 | 3010000
| HLQC 3 | 366000 | 378000 | 395000 | 267000 | 259000 | 284000 | 6050000 | 5700000 | 5890000 | 3280000 | 3290000 | 3170000 | 350000 | 364000 | 364000 | 362000 | 367000 | 389000 | 8660000 | 8420000 | 8560000 | 3050000 | 2910000 | 2890000
| HLQC 4 | 389000 | 397000 | 397000 | 292000 | 275000 | 286000 | 6020000 | 6360000 | 6270000 | 3270000 | 3320000 | 3500000 | 343000 | 358000 | 343000 | 359000 | 362000 | 369000 | 8130000 | 8410000 | 8290000 | 2860000 | 2900000 | 2880000
| HLQC 5 | 408000 | 413000 | 393000 | 280000 | 265000 | 263000 | 6150000 | 6290000 | 6310000 | 3280000 | 3310000 | 3220000 | 354000 | 358000 | 345000 | 379000 | 374000 | 358000 | 8610000 | 8420000 | 8570000 | 2920000 | 2900000 | 3000000
| HLQC 6 | 408000 | 422000 | 397000 | 264000 | 264000 | 276000 | 5990000 | 6060000 | 5900000 | 3230000 | 3180000 | 3220000 | 345000 | 354000 | 338000 | 369000 | 374000 | 351000 | 8240000 | 8310000 | 8120000 | 2840000 | 2770000 | 2860000
| Blank QC | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| Total Blank 1 | 30300 |  |  | 24400 |  |  | 133000 |  |  | 11500
| Total Blank 2 | 15700 |  |  | 14800 |  |  | 84700 |  |  | 4760 | [8 empty] | 2550
| Total Blank 3 | 10700 |  |  | 10200 |  |  | 120000 |  |  | 24900 | [8 empty] | 2060
Note on B112: Author: total blank, not matrix blank. Measured 3 times in the whole sequence

## Raw Peak Areas
Sample_Name | Cer d18:1/16:0 [M>264] | Cer d18:1/16:0-d7 [M>264] | Cer d18:1/18:0 [M>264] | Cer d18:1/18:0-d7 [M>264] | Cer d18:1/24:0 [M>264] | Cer d18:1/24:0-d7 [M>264] | Cer d18:1/24:1 [M>264] | Cer d18:1/24:1-d7 [M>264]
Total Blank 1 | 30300 |  | 24400 |  | 133000 |  | 11500
Matrix blank | 673000 |  | 563000 |  | 19000000 |  | 5420000
Std 6a | 65000 | 687000 | 98600 | 832000 | 1410000 | 21800000 | 1110000 | 7000000
Std 6b | 57100 | 506000 | 88100 | 804000 | 1230000 | 20800000 | 1020000 | 6810000
Std 6c | 48100 | 485000 | 87300 | 793000 | 1340000 | 20800000 | 1050000 | 6750000
Std 5a | 67200 | 577000 | 98700 | 775000 | 1610000 | 20700000 | 1260000 | 6410000
Std 5b | 58500 | 479000 | 102000 | 727000 | 1570000 | 20600000 | 1280000 | 6460000
Std 5c | 58600 | 473000 | 103000 | 715000 | 1610000 | 20600000 | 1360000 | 6360000
Std 4a | 139000 | 617000 | 218000 | 770000 | 3170000 | 21100000 | 2670000 | 6760000
Std 4b | 120000 | 582000 | 213000 | 790000 | 2920000 | 20800000 | 2480000 | 6490000
Std 4c | 109000 | 484000 | 209000 | 782000 | 3050000 | 20800000 | 2660000 | 6680000
Std 3a | 563000 | 494000 | 1080000 | 774000 | 14900000 | 21500000 | 13100000 | 6760000
Std 3b | 563000 | 534000 | 1030000 | 787000 | 14000000 | 20400000 | 12400000 | 6750000
Std 3c | 577000 | 508000 | 1080000 | 810000 | 13800000 | 20000000 | 13000000 | 6600000
Std 2a | 5130000 | 501000 | 9690000 | 767000 | 119000000 | 18900000 | 108000000 | 6380000
Std 2b | 5280000 | 514000 | 9680000 | 747000 | 114000000 | 19200000 | 105000000 | 6150000
Std 2c | 5500000 | 516000 | 9600000 | 748000 | 116000000 | 18800000 | 106000000 | 5930000
Std 1a | 10600000 | 507000 | 21500000 | 890000 | 209000000 | 18700000 | 190000000 | 6140000
Std 1b | 12700000 | 624000 | 18800000 | 754000 | 193000000 | 17000000 | 175000000 | 5600000
Std 1c | 11400000 | 580000 | 18200000 | 686000 | 197000000 | 17200000 | 174000000 | 5370000
pooled QC | 547000 | 409000 | 438000 | 772000 | 16900000 | 19700000 | 5130000 | 5890000
pooled QC | 569000 | 426000 | 427000 | 717000 | 17400000 | 18800000 | 5040000 | 5610000
SRM 1950-1a | 517000 | 338000 | 425000 | 714000 | 16000000 | 19100000 | 5020000 | 5500000
SRM 1950-1b | 496000 | 330000 | 421000 | 713000 | 16000000 | 18200000 | 4880000 | 5510000
SRM 1950-1c | 493000 | 332000 | 403000 | 682000 | 15800000 | 18600000 | 4870000 | 5580000
SRM 1950-2a | 464000 | 341000 | 410000 | 702000 | 15000000 | 17500000 | 4690000 | 5370000
SRM 1950-2b | 507000 | 351000 | 399000 | 715000 | 15300000 | 18200000 | 4720000 | 5390000
SRM 1950-2c | 505000 | 341000 | 383000 | 723000 | 14900000 | 17500000 | 4600000 | 5330000
SRM 1950-3a | 476000 | 328000 | 400000 | 708000 | 14600000 | 17500000 | 4560000 | 5240000
SRM 1950-3b | 499000 | 338000 | 410000 | 737000 | 15300000 | 18500000 | 4630000 | 5410000
SRM 1950-3c | 467000 | 333000 | 425000 | 720000 | 15100000 | 18000000 | 4470000 | 5470000
SRM 1950-4a | 471000 | 306000 | 399000 | 694000 | 14700000 | 17000000 | 4350000 | 5170000
SRM 1950-4b | 464000 | 317000 | 401000 | 679000 | 14200000 | 16900000 | 4340000 | 5090000
SRM 1950-4c | 474000 | 319000 | 392000 | 689000 | 15000000 | 16700000 | 4410000 | 4970000
SRM 1950-5a | 454000 | 306000 | 386000 | 670000 | 14000000 | 17300000 | 4310000 | 5080000
SRM 1950-5b | 464000 | 306000 | 390000 | 691000 | 13900000 | 16800000 | 4330000 | 4880000
SRM 1950-5c | 467000 | 300000 | 397000 | 683000 | 14700000 | 17200000 | 4290000 | 5080000
SRM 1950-6a | 450000 | 300000 | 372000 | 658000 | 14000000 | 16300000 | 4300000 | 4760000
SRM 1950-6b | 484000 | 308000 | 362000 | 661000 | 14500000 | 16600000 | 4200000 | 4980000
SRM 1950-6c | 465000 | 320000 | 409000 | 689000 | 14900000 | 17300000 | 4240000 | 5000000
pooled QC | 518000 | 388000 | 380000 | 694000 | 15700000 | 17500000 | 4270000 | 4740000
pooled QC | 529000 | 387000 | 357000 | 649000 | 15100000 | 16500000 | 4270000 | 4820000
T1D-1a | 430000 | 322000 | 437000 | 661000 | 14400000 | 16900000 | 4380000 | 4980000
T1D-1b | 448000 | 331000 | 423000 | 629000 | 13900000 | 16000000 | 4360000 | 4890000
T1D-1c | 441000 | 317000 | 423000 | 658000 | 13700000 | 15800000 | 4320000 | 4730000
T1D-2a | 426000 | 313000 | 407000 | 628000 | 13000000 | 14400000 | 4030000 | 4480000
T1D-2b | 434000 | 311000 | 389000 | 604000 | 13200000 | 15000000 | 4010000 | 4550000
T1D-2c | 428000 | 311000 | 389000 | 581000 | 12900000 | 14600000 | 4090000 | 4530000
T1D-3a | 410000 | 284000 | 400000 | 592000 | 13000000 | 15000000 | 3920000 | 4470000
T1D-3b | 411000 | 308000 | 395000 | 575000 | 13000000 | 14800000 | 3960000 | 4450000
T1D-3c | 435000 | 297000 | 380000 | 580000 | 12800000 | 14300000 | 3980000 | 4350000
T1D-4a | 409000 | 303000 | 396000 | 593000 | 12900000 | 15200000 | 3800000 | 4450000
T1D-4b | 411000 | 297000 | 398000 | 581000 | 12700000 | 15100000 | 3930000 | 4370000
T1D-4c | 419000 | 289000 | 369000 | 576000 | 12600000 | 14900000 | 3830000 | 4390000
T1D-5a | 411000 | 293000 | 367000 | 554000 | 12800000 | 14000000 | 3950000 | 4250000
T1D-5b | 421000 | 292000 | 368000 | 545000 | 12000000 | 14000000 | 3840000 | 4080000
T1D-5c | 408000 | 307000 | 374000 | 540000 | 12100000 | 13900000 | 3740000 | 4150000
T1D-6a | 410000 | 296000 | 386000 | 552000 | 12300000 | 13900000 | 3620000 | 4040000
T1D-6b | 426000 | 297000 | 352000 | 528000 | 12000000 | 13700000 | 3640000 | 4050000
T1D-6c | 466000 | 343000 | 367000 | 559000 | 12100000 | 13800000 | 3690000 | 4100000
total blank 2 | 15700 |  | 14800 |  | 84700 | 2550 | 4760
pooled QC | 484000 | 361000 | 309000 | 564000 | 12500000 | 13900000 | 3520000 | 4010000
pooled QC | 490000 | 361000 | 290000 | 532000 | 11900000 | 13700000 | 3400000 | 3920000
Young AA new-1a | 311000 | 324000 | 188000 | 521000 | 7810000 | 12100000 | 1980000 | 3710000
Young AA new-1b | 316000 | 333000 | 176000 | 496000 | 7590000 | 11700000 | 2030000 | 3590000
Young AA new-1c | 338000 | 352000 | 175000 | 496000 | 7680000 | 11800000 | 1990000 | 3680000
Young AA new-2a | 311000 | 328000 | 183000 | 480000 | 7920000 | 11400000 | 2030000 | 3570000
Young AA new-2b | 327000 | 322000 | 183000 | 498000 | 7910000 | 11600000 | 2030000 | 3560000
Young AA new-2c | 312000 | 328000 | 184000 | 486000 | 7920000 | 11500000 | 2010000 | 3580000
Young AA new-3a | 307000 | 312000 | 176000 | 463000 | 7660000 | 10900000 | 1990000 | 3480000
Young AA new-3b | 310000 | 315000 | 173000 | 480000 | 8090000 | 12100000 | 2000000 | 3540000
Young AA new-3c | 323000 | 321000 | 184000 | 547000 | 7820000 | 11800000 | 2050000 | 3650000
Young AA new-4a | 331000 | 323000 | 181000 | 486000 | 7780000 | 11400000 | 2090000 | 3630000
Young AA new-4b | 325000 | 323000 | 183000 | 475000 | 7860000 | 11200000 | 1980000 | 3490000
Young AA new-4c | 333000 | 313000 | 190000 | 533000 | 8070000 | 12400000 | 2070000 | 3640000
Young AA new-5a | 332000 | 321000 | 183000 | 500000 | 7620000 | 11200000 | 1980000 | 3550000
Young AA new-5b | 318000 | 324000 | 198000 | 522000 | 7810000 | 11700000 | 2120000 | 3670000
Young AA new-5c | 321000 | 330000 | 169000 | 461000 | 7160000 | 11000000 | 2010000 | 3610000
Young AA new-6a | 316000 | 319000 | 174000 | 430000 | 7440000 | 10600000 | 1940000 | 3320000
Young AA new-6b | 316000 | 314000 | 174000 | 469000 | 7460000 | 10400000 | 1970000 | 3450000
Young AA new-6c | 327000 | 312000 | 175000 | 455000 | 7370000 | 10600000 | 1890000 | 3330000
Pooled QC | 480000 | 361000 | 255000 | 495000 | 11300000 | 12100000 | 3100000 | 3500000
Pooled QC | 470000 | 364000 | 280000 | 478000 | 10900000 | 12100000 | 3150000 | 3390000
hTAG-1a | 476000 | 268000 | 320000 | 483000 | 16000000 | 11900000 | 4250000 | 3250000
hTAG-1b | 464000 | 259000 | 326000 | 464000 | 14900000 | 11600000 | 4300000 | 3330000
hTAG-1c | 472000 | 256000 | 333000 | 471000 | 15300000 | 11300000 | 4050000 | 3200000
hTAG-2a | 453000 | 262000 | 335000 | 471000 | 15000000 | 11400000 | 3990000 | 3140000
hTAG-2b | 440000 | 248000 | 315000 | 444000 | 14400000 | 11100000 | 4080000 | 3190000
hTAG-2c | 444000 | 263000 | 307000 | 461000 | 14600000 | 10900000 | 4080000 | 3190000
hTAG-3a | 458000 | 257000 | 322000 | 475000 | 15700000 | 11700000 | 3990000 | 3140000
hTAG-3b | 451000 | 254000 | 298000 | 421000 | 14500000 | 11000000 | 4030000 | 3120000
hTAG-3c | 451000 | 258000 | 295000 | 432000 | 14500000 | 10800000 | 3930000 | 3220000
hTAG-4a | 436000 | 268000 | 298000 | 447000 | 14200000 | 10700000 | 3920000 | 3130000
hTAG-4b | 446000 | 260000 | 294000 | 451000 | 14900000 | 11200000 | 4060000 | 3240000
hTAG-4c | 445000 | 267000 | 320000 | 445000 | 14600000 | 10900000 | 3970000 | 3210000
hTAG-5a | 446000 | 249000 | 304000 | 424000 | 14200000 | 10700000 | 3900000 | 3100000
hTAG-5b | 452000 | 241000 | 296000 | 449000 | 14700000 | 11000000 | 3740000 | 2990000
hTAG-5c | 462000 | 254000 | 292000 | 440000 | 13800000 | 10600000 | 3950000 | 3150000
hTAG-6a | 471000 | 246000 | 310000 | 437000 | 14600000 | 10900000 | 3920000 | 3080000
hTAG-6b | 466000 | 241000 | 284000 | 404000 | 14200000 | 10600000 | 3960000 | 3080000
hTAG-6c | 498000 | 268000 | 284000 | 417000 | 14500000 | 10700000 | 3840000 | 3070000
Pooled QC | 504000 | 373000 | 249000 | 458000 | 10500000 | 11200000 | 3070000 | 3330000
Pooled QC | 514000 | 385000 | 259000 | 433000 | 10200000 | 11300000 | 2930000 | 3220000
LLQC 1a | 225000 | 368000 | 102000 | 415000 | 4220000 | 10900000 | 1270000 | 3500000
LLQC 1b | 237000 | 373000 | 110000 | 410000 | 4280000 | 10700000 | 1210000 | 3590000
LLQC 1c | 227000 | 368000 | 108000 | 424000 | 4230000 | 10600000 | 1260000 | 3650000
LLQC 2a | 196000 | 361000 | 103000 | 440000 | 3780000 | 10700000 | 1170000 | 3730000
LLQC 2b | 206000 | 377000 | 91300 | 419000 | 3810000 | 10600000 | 1130000 | 3530000
LLQC 2c | 199000 | 387000 | 92200 | 409000 | 3720000 | 10600000 | 1120000 | 3590000
LLQC 3a | 275000 | 344000 | 144000 | 450000 | 6080000 | 11300000 | 1780000 | 3710000
LLQC 3b | 291000 | 370000 | 139000 | 474000 | 6110000 | 10700000 | 1770000 | 3680000
LLQC 3c | 275000 | 362000 | 139000 | 441000 | 5600000 | 10600000 | 1660000 | 3630000
LLQC 4a | 263000 | 366000 | 131000 | 433000 | 4930000 | 10700000 | 1460000 | 3530000
LLQC 4b | 268000 | 380000 | 133000 | 431000 | 4900000 | 10800000 | 1460000 | 3610000
LLQC 4c | 273000 | 381000 | 138000 | 443000 | 4970000 | 10700000 | 1400000 | 3660000
LLQC 5a | 199000 | 381000 | 96400 | 451000 | 4200000 | 11400000 | 1180000 | 3710000
LLQC 5b | 214000 | 394000 | 96900 | 441000 | 3910000 | 10500000 | 1160000 | 3610000
LLQC 5c | 209000 | 383000 | 104000 | 439000 | 4160000 | 11100000 | 1240000 | 3660000
LLQC 6a | 201000 | 372000 | 95100 | 421000 | 3930000 | 10800000 | 1070000 | 3560000
LLQC 6b | 195000 | 359000 | 92500 | 420000 | 3780000 | 11100000 | 1110000 | 3560000
LLQC 6c | 200000 | 385000 | 103000 | 433000 | 3960000 | 10800000 | 1200000 | 3690000
Pooled QC | 526000 | 422000 | 243000 | 442000 | 10100000 | 11600000 | 3030000 | 3300000
Pooled QC | 518000 | 412000 | 257000 | 430000 | 10700000 | 11400000 | 3130000 | 3390000
LQC 1a | 303000 | 356000 | 165000 | 470000 | 6530000 | 11900000 | 1960000 | 3720000
LQC 1b | 293000 | 347000 | 168000 | 459000 | 6350000 | 11800000 | 2010000 | 3630000
LQC 1c | 311000 | 355000 | 158000 | 458000 | 6550000 | 11600000 | 2000000 | 3710000
LQC 2a | 283000 | 363000 | 141000 | 457000 | 5690000 | 11400000 | 1730000 | 3710000
LQC 2b | 279000 | 372000 | 147000 | 476000 | 5390000 | 11800000 | 1710000 | 3680000
LQC 2c | 292000 | 390000 | 146000 | 464000 | 5550000 | 11800000 | 1700000 | 3740000
LQC 3a | 331000 | 359000 | 174000 | 461000 | 6420000 | 11600000 | 1870000 | 3600000
LQC 3b | 343000 | 382000 | 177000 | 510000 | 6530000 | 12400000 | 1880000 | 3730000
LQC 3c | 330000 | 357000 | 170000 | 473000 | 6360000 | 11900000 | 1930000 | 3590000
LQC 4a | 303000 | 356000 | 156000 | 440000 | 6330000 | 11700000 | 1870000 | 3580000
LQC 4b | 308000 | 355000 | 156000 | 464000 | 6620000 | 11700000 | 1840000 | 3550000
LQC 4c | 308000 | 364000 | 155000 | 463000 | 6160000 | 11300000 | 1860000 | 3510000
LQC 5a | 312000 | 353000 | 157000 | 458000 | 6300000 | 11400000 | 1840000 | 3450000
LQC 5b | 332000 | 354000 | 164000 | 434000 | 6290000 | 11300000 | 1820000 | 3460000
LQC 5c | 345000 | 365000 | 176000 | 448000 | 6420000 | 11400000 | 2000000 | 3680000
LQC 6a | 305000 | 359000 | 153000 | 438000 | 6210000 | 11300000 | 1830000 | 3550000
LQC 6b | 300000 | 360000 | 151000 | 437000 | 5910000 | 11200000 | 1810000 | 3560000
LQC 6c | 307000 | 353000 | 155000 | 427000 | 6080000 | 11500000 | 1800000 | 3540000
Pooled QC | 543000 | 414000 | 238000 | 422000 | 10200000 | 10800000 | 2770000 | 3060000
Pooled QC | 554000 | 428000 | 256000 | 431000 | 10400000 | 10700000 | 2850000 | 3220000
MQC 1a | 390000 | 287000 | 222000 | 418000 | 9230000 | 10400000 | 2660000 | 3120000
MQC 1b | 409000 | 290000 | 223000 | 427000 | 9470000 | 10600000 | 2520000 | 2980000
MQC 1c | 403000 | 303000 | 230000 | 407000 | 9330000 | 10700000 | 2650000 | 3150000
MQC 2a | 400000 | 290000 | 240000 | 402000 | 10200000 | 10500000 | 2730000 | 2890000
MQC 2b | 426000 | 294000 | 271000 | 434000 | 10000000 | 10800000 | 2800000 | 3000000
MQC 2c | 421000 | 279000 | 239000 | 388000 | 10000000 | 10400000 | 2830000 | 3070000
MQC 3a | 431000 | 280000 | 252000 | 400000 | 10300000 | 10700000 | 2800000 | 2940000
MQC 3b | 422000 | 278000 | 245000 | 404000 | 10200000 | 10300000 | 2720000 | 2890000
MQC 3c | 420000 | 293000 | 254000 | 416000 | 10100000 | 10600000 | 2730000 | 2890000
MQC 4a | 405000 | 300000 | 230000 | 417000 | 9690000 | 10700000 | 2610000 | 3030000
MQC 4b | 411000 | 292000 | 237000 | 399000 | 9710000 | 10700000 | 2670000 | 3030000
MQC 4c | 413000 | 304000 | 225000 | 418000 | 9130000 | 10300000 | 2620000 | 3030000
MQC 5a | 409000 | 277000 | 229000 | 381000 | 9190000 | 9630000 | 2570000 | 2900000
MQC 5b | 411000 | 287000 | 231000 | 390000 | 9430000 | 10300000 | 2570000 | 2900000
MQC 5c | 406000 | 300000 | 225000 | 382000 | 9280000 | 10000000 | 2550000 | 2930000
MQC 6a | 405000 | 277000 | 205000 | 403000 | 9100000 | 9790000 | 2480000 | 2810000
MQC 6b | 401000 | 283000 | 224000 | 383000 | 8920000 | 10300000 | 2680000 | 2990000
MQC 6c | 407000 | 285000 | 218000 | 391000 | 9000000 | 9850000 | 2550000 | 2910000
Pooled QC | 584000 | 446000 | 236000 | 394000 | 8760000 | 9390000 | 2420000 | 2740000
Pooled QC | 717000 | 566000 | 225000 | 389000 | 9050000 | 10200000 | 2500000 | 2760000
HQC 1a | 405000 | 342000 | 266000 | 395000 | 7310000 | 9360000 | 3240000 | 3140000
HQC 1b | 412000 | 333000 | 260000 | 395000 | 7140000 | 9170000 | 3010000 | 3080000
HQC 1c | 394000 | 334000 | 243000 | 363000 | 6730000 | 8870000 | 2980000 | 3060000
HQC 2a | 410000 | 338000 | 260000 | 355000 | 7190000 | 8610000 | 3110000 | 2950000
HQC 2b | 415000 | 342000 | 257000 | 383000 | 7000000 | 8780000 | 3100000 | 3030000
HQC 2c | 406000 | 352000 | 265000 | 400000 | 7260000 | 8720000 | 3160000 | 3130000
HQC 3a | 556000 | 362000 | 286000 | 370000 | 6800000 | 8640000 | 3090000 | 3070000
HQC 3b | 527000 | 352000 | 290000 | 354000 | 6580000 | 8930000 | 2970000 | 3060000
HQC 3c | 552000 | 344000 | 299000 | 394000 | 7360000 | 9250000 | 3030000 | 3140000
HQC 4a | 434000 | 354000 | 293000 | 388000 | 7670000 | 9560000 | 3400000 | 3220000
HQC 4b | 423000 | 347000 | 279000 | 385000 | 6980000 | 8740000 | 3190000 | 3020000
HQC 4c | 424000 | 351000 | 283000 | 391000 | 7120000 | 8880000 | 3030000 | 3100000
HQC 5a | 433000 | 344000 | 280000 | 409000 | 7620000 | 9540000 | 3360000 | 3030000
HQC 5b | 418000 | 335000 | 281000 | 376000 | 7550000 | 9220000 | 3370000 | 3140000
HQC 5c | 407000 | 328000 | 269000 | 372000 | 7100000 | 8780000 | 3200000 | 2980000
HQC 6a | 389000 | 349000 | 246000 | 382000 | 6390000 | 8520000 | 3020000 | 3090000
HQC 6b | 387000 | 341000 | 237000 | 367000 | 5900000 | 8410000 | 2840000 | 2820000
HQC 6c | 372000 | 339000 | 237000 | 373000 | 6120000 | 8430000 | 2730000 | 2870000
Pooled QC | 675000 | 536000 | 219000 | 384000 | 8990000 | 9830000 | 2470000 | 2760000
Pooled QC | 763000 | 577000 | 223000 | 368000 | 8640000 | 9370000 | 2440000 | 2700000
HLQC 1a | 363000 | 326000 | 278000 | 367000 | 5650000 | 8410000 | 3320000 | 2930000
HLQC 1b | 368000 | 355000 | 266000 | 352000 | 5500000 | 8300000 | 3260000 | 2900000
HLQC 1c | 361000 | 334000 | 250000 | 353000 | 5670000 | 8160000 | 3240000 | 2900000
HLQC 2a | 371000 | 348000 | 263000 | 370000 | 5600000 | 8310000 | 3160000 | 2910000
HLQC 2b | 362000 | 348000 | 265000 | 370000 | 5440000 | 8110000 | 3260000 | 2920000
HLQC 2c | 361000 | 340000 | 259000 | 345000 | 5610000 | 8630000 | 3180000 | 3010000
HLQC 3a | 366000 | 350000 | 267000 | 362000 | 6050000 | 8660000 | 3280000 | 3050000
HLQC 3b | 378000 | 364000 | 259000 | 367000 | 5700000 | 8420000 | 3290000 | 2910000
HLQC 3c | 395000 | 364000 | 284000 | 389000 | 5890000 | 8560000 | 3170000 | 2890000
HLQC 4a | 389000 | 343000 | 292000 | 359000 | 6020000 | 8130000 | 3270000 | 2860000
HLQC 4b | 397000 | 358000 | 275000 | 362000 | 6360000 | 8410000 | 3320000 | 2900000
HLQC 4c | 397000 | 343000 | 286000 | 369000 | 6270000 | 8290000 | 3500000 | 2880000
HLQC 5a | 408000 | 354000 | 280000 | 379000 | 6150000 | 8610000 | 3280000 | 2920000
HLQC 5b | 413000 | 358000 | 265000 | 374000 | 6290000 | 8420000 | 3310000 | 2900000
HLQC 5c | 393000 | 345000 | 263000 | 358000 | 6310000 | 8570000 | 3220000 | 3000000
HLQC 6a | 408000 | 345000 | 264000 | 369000 | 5990000 | 8240000 | 3230000 | 2840000
HLQC 6b | 422000 | 354000 | 264000 | 374000 | 6060000 | 8310000 | 3180000 | 2770000
HLQC 6c | 397000 | 338000 | 276000 | 351000 | 5900000 | 8120000 | 3220000 | 2860000
Pooled QC | 728000 | 560000 | 207000 | 346000 | 8350000 | 9110000 | 2380000 | 2540000
Pooled QC | 748000 | 558000 | 198000 | 360000 | 8470000 | 9050000 | 2400000 | 2640000
Std 6a | 40200 | 386000 | 39400 | 321000 | 592000 | 9180000 | 484000 | 2880000
Std 6b | 30500 | 301000 | 32400 | 331000 | 552000 | 9030000 | 431000 | 2940000
Std 6c | 31800 | 304000 | 39100 | 333000 | 573000 | 9150000 | 478000 | 2910000
Std 5a | 39000 | 365000 | 39600 | 318000 | 697000 | 9090000 | 571000 | 2910000
Std 5b | 43300 | 306000 | 50400 | 332000 | 698000 | 9090000 | 557000 | 2930000
Std 5c | 40000 | 304000 | 41500 | 326000 | 689000 | 8800000 | 596000 | 2770000
Std 4a | 109000 | 393000 | 92200 | 344000 | 1360000 | 8940000 | 1190000 | 2890000
Std 4b | 89400 | 360000 | 84300 | 342000 | 1260000 | 9030000 | 1110000 | 2950000
Std 4c | 78400 | 298000 | 91200 | 341000 | 1270000 | 8910000 | 1130000 | 2940000
Std 3a | 392000 | 307000 | 472000 | 352000 | 6470000 | 9200000 | 6210000 | 3100000
Std 3b | 418000 | 352000 | 449000 | 332000 | 6210000 | 9090000 | 5740000 | 2920000
Std 3c | 406000 | 323000 | 452000 | 322000 | 6200000 | 8810000 | 5680000 | 2850000
Std 2a | 3800000 | 329000 | 4320000 | 352000 | 58300000 | 9010000 | 55000000 | 2900000
Std 2b | 3610000 | 311000 | 4000000 | 322000 | 55000000 | 8590000 | 48700000 | 2780000
Std 2c | 3850000 | 323000 | 4180000 | 310000 | 56400000 | 8510000 | 54000000 | 2790000
Std 1a | 7280000 | 318000 | 8210000 | 324000 | 107000000 | 9020000 | 97200000 | 2740000
Std 1b | 9230000 | 422000 | 8280000 | 334000 | 113000000 | 8720000 | 96200000 | 2670000
Std 1c | 8690000 | 391000 | 8150000 | 319000 | 108000000 | 8480000 | 96600000 | 2590000
Total Blank 3 | 10700 |  | 10200 |  | 120000 | 2060 | 24900
Matrix blank | 668000 |  | 302000 |  | 9500000 |  | 2600000