## Participant Information
Participant Information
Participant Information |  | Group Name
Lab code
| Institution | Institution name
Street
City
Postal code
Country
| Contact person | Name
Address
Email
Phone

## Summary
Summary
Summary | This protocol describes the extraction and LC-MRM procedure for absolute quantification of Cer d18:1/16:0, Cer d18:1/18:0, Cer d18:1/24:0 and Cer d18:1/24:1 in human plasma according to Kauhanen et al, Anal Bioanal Chem (2016) 408:3475–3483.
A calibration curve is built by injecting several dilutions of the non-labelled standards mixed with a fixed concentration of labelled standards. Each calibration point is prepared and analysed in triplicates (n=3).
To establish the endogenous concentration of the four selected ceramides in the study samples (NIST SRM1950, NIST high TAG, NIST T1D, and NIST young A-A) each sample is extracted and analyzed in six replicates (n =6).
Prepare the pooled QC by mixing 100 mL each of NIST SRM1950, NIST high TAG, NIST T1D, and NIST young AA (total 400 mL).
We suggest proceeding first to measure the ceramides concentration in the four study samples and then continue with the intra-assay validation to avoid wasting material.

## Extraction - Overview
Extraction Overview
Extraction Overview | Table 1 - Extraction - Materials
Chemicals, solvents and materials
Ethyl acetate (EtAc)
2-propanol (IPA)
5% Bovine Serum Albumin (BSA)
96-well plates/tubes
Sealing foil
2 mL DeepWell plates/tubes/LC vials
Table 2 - Extraction - Sample Types
| Sample type | Components
| Pooled Quality Control samples (QC) | 10 µL pooled QC |  |  | 20 µL Labelled IS |  | 570 µL EtAc:IPA 2:8 (v/v)
| Matrix Blank (MB) | 10 µL Matrix (pooled QC) only |  |  | - |  | 590 µL EtAc:IPA 2:8
| Study Samples (S) | 10 µL Human plasma (NIST SRM1950, high TAG, T1D, young AA) |  |  | 20 µL Labelled IS |  | 570 µL EtAc:IPA 2:8
| Total Blank (TB) | 10 µL 5% BSA |  |  | - |  | 590 µL EtAc:IPA 2:8
| Calibration line standards (STD) | 10 µL 5% BSA + 20 µL non labelled STD mixture (Table 5) |  |  | 20 µL Labelled IS |  | 550 µL EtAc:IPA 2:8
Table 3 - Extraction - Labelled Internal Standard Mixture Solution (provided)
| Labelled IS mixture in EtAc:IPA 2:8 (v/v) | Concentration |  |  |  |  | Unit | Vendor
| D7-Cer d18:1/16:0 | 0.125 |  |  |  |  | pmol/µL [µM] | Avanti Polar Lipids
| D7-Cer d18:1/18:0 | 0.050
| D7-Cer d18:1/24:0 | 1500
| D7-Cer d18:1/24:1 | 0.500
Table 4 - Extraction - Non-Labelled Standards - for STD1 Mixture Solution (provided)
| Non-labelled standards in EtAc:IPA 2:8 (v/v) | Concentration |  |  |  |  | Unit | Vendor
| Cer d18:1/16:0 | 2.0 |  |  |  |  | pmol/µL [µM] | Avanti Polar Lipids
| Cer d18:1/18:0 | 2.0
| Cer d18:1/24:0 | 20.0
| Cer d18:1/24:1 | 20.0
Table 5 - Serial dilutions of Non-Labelled Standards from STD1 Mixture Solution for calibration line
| Add 20 µL of | STD1 Conc. | STD2 Conc. | STD3 Conc. | STD4 Conc. | STD5 Conc. | STD6 Conc. | Unit
| Cer d18:1/16:0 | 2.0 | 1.0 | 0.10 | 0.02 | 0.01 | 0.008 | pmol/µL [µM]
| Cer d18:1/18:0 | 2.0 | 1.0 | 0.10 | 0.02 | 0.01 | 0.008
| Cer d18:1/24:0 | 20.0 | 10.0 | 1.0 | 0.2 | 0.1 | 0.08
| Cer d18:1/24:1 | 20.0 | 10.0 | 1.0 | 0.2 | 0.1 | 0.08

## Extraction - Procedure
Extraction Procedure
Extraction Procedure | Extraction Procedure
| Step | Description | Reference Table
| 1 | Prepare different dilutions of STD1 by diluting the provided mixture (Extraction - Overview, Table 4) with EtAc:IPA 2:8 as suggested in table | Extraction - Overview'!B34
| 2 | Thaw plasma on ice (+4°C) and bring to room temperature prior to extraction. Mix by pipetting before use.
| 3 | Prepare samples in plates/tubes manually or with a robot, according to specifications in column C of table | Extraction - Overview'!C14
| 4 | Add 20 µL of labelled internal standard mixture to samples according to column F of table | Extraction - Overview'!F14
| 5 | For the calibration line, add 20 µL of different dilutions of the non-labelled mixture (STD1-6).
| 6 | Add 590 µL of ethyl acetate:isopropanol (2:8, vol/vol) to Total Blank (TB) and Matrix Blank (MB); 550 µL to calibration line (STD); 570 µL to study samples and QCs according to column H of table | Extraction - Overview'!H14
| 7 | Mix samples 10 min by pipetting or vortexing
| 8 | Centrifuge samples for 10 min at 3000 x gat room temperature
| 9 | Transfer 50 µL of the clear supernatant to MS-vials or MS plates, close/seal, store samples at -20 °C prior to LC-MRM analysis.

## LC-MRM - Overview
LC-MRM Overview
LC-MRM Overview | Table 6 Chemicals and solvents
Ammonium formate
Formic acid
Ultra-pure water
Acetonitrile
Table 7 LC-MRM LC-MS Details
| Property | Specifications
| Mass spectrometry | QQQ
| LC | UPLC
| Injection volume (indicative) | 5 µL
| Weak wash | Water
| Strong wash | Acetonitrile
| Pre-column | Waters Acquity BEH C18, 1.7 µm VanGuard Pre-Column
| Column | Waters Acquity BEH C18, 2.1 × 50 mm id. 1.7 µm
| Column temperature | 60 °C
| Mobile phase A | 10 mM ammonium acetate in water with 0.1% formic acid
| Mobile phase B | 10 mM ammonium acetate in acetonitrile:2-propanol (4:3, vol/vol) with 0.1% formic acid
Table 8 LC-MRM LC gradient
| Time (minutes) | A% | B% |  | Flow (µL/min)
| 0.00 | 15 | 85 |  | 500
| 0.50 | 15 | 85 |  | 500
| 1.50 | 0 | 100 |  | 500
| 4.00 | 0 | 100 |  | 500
| 4.10 | 15 | 85 |  | 500
| 5.00 | 15 | 85 |  | 500
Table 9 LC-MRM MS conditions (EXAMPLE: this only applies to Sciex QTRAP, please adapt)
| Polarity mode | Positive
| Ion spray | 5000 V
| Curtain gas | 25 psi
| Source temperature | 300 °C
| Gas 1 | 50 psi
| Gas 2 | 35 psi
| DP | 30 V
| EP | 10 V
| Collision cell exit potential | 20 V
| Collision energy (CE) | 40 eV
Table 10 LC-MRM Mass transitions used for absolute quantification. On different systems, RT and CE might be different!
| Analyte | Mass transitions | IS | Mass transitions | RT (min) | CE
| Cer d18:1/16:0 | 538.5 – 264.3 | D7- Cer d18:1/16:0 | 545.6 - 271.3 | 1.9-2.0 | 40
C34H68NO3 – C18H34N
| Cer d18:1/18:0 | 566.6 – 264.3 | D7- Cer d18:1/18:0 | 573.6 – 271.3 | 2.1-2.2 | 40
C36H72NO3 – C18H34N
| Cer d18:1/24:0 | 650.6 – 264.3 | D7- Cer d18:1/24:0 | 657.7 – 271.3 | 2.4-2.5 | 40
C42H84NO3 – C18H34N
| Cer d18:1/24:1 | 648.6 – 264.3 | D7- Cer d18:1/24:1 | 655.7 – 271.3 | 2.3-2.4 | 40
C42H82NO3 – C18H34N
Table 11 Data processing
| Processing software | (Add here name and version of the software used)

## LC-MRM - Sample Sequence
Sample Sequence (Example)
Sample Sequence (Example) | Sample Sequence 1
| Number | Sample Name | Remarks
| 1 | Total Blank 1
| 2 | Matrix blank
| 3 | Calibration line 1 from STD6 to 1 (n=3)
| 4 | […]
| 5 | Pooled QC
| 6 | Pooled QC
| 7 | SRM 1950-1
| 8 | SRM 1950-2
| 9 | SRM 1950-3
| 10 | SRM 1950-4
| 11 | SRM 1950-5
| 12 | SRM 1950-6
| 13 | Pooled QC
| 14 | Pooled QC
| 15 | T1D-1
| 16 | T1D-2
| 17 | T1D-3
| 18 | T1D-4
| 19 | T1D-5
| 20 | T1D-6
| 21 | Total Blank 2
| 22 | Pooled QC
| 23 | Pooled QC
| 24 | Young AA-1
| 25 | Young AA-2
| 26 | Young AA-3
| 27 | Young AA-4
| 28 | Young AA-5
| 29 | Young AA-6
| 30 | Pooled QC
| 31 | Pooled QC
| 32 | hTAG-1
| 33 | hTAG-2
| 34 | hTAG-3
| 35 | hTAG-4
| 36 | hTAG-5
| 37 | hTAG-6
| 38 | Pooled QC
| 39 | Pooled QC
| 40 | Calibration line 2 from STD1 to 6 (n=3)
| 41 | […]
| 42 | Matrix blank
| 43 | Total Blank 3

## LC-MRM - Intra Assay QC Proc
Intra Assay Procedure
Intra Assay Procedure | Quality Controls for Intra-Assay Variation
| Step | Description | Reference Table
| 1 | Prepare the Pooled NIST QC sample containing equal volumes of the four NIST reference plasma you received.
| 2 | The pooled QC representing middle quality control (MQC), low QC (LQC), and lower limit of quantification QC (LLOQ) is prepared by diluting MQC (e.g. two- and three-fold) with water.
| 3 | High QC (HQC) and upper limit of quantification (ULOQ) QC are prepared by spiking in 10 µL and 20 µL respectively of STD3 endogenous standards from table | Extraction - Overview'!B34
| 4 | Internal standards are added to all QC samples.
Intra-Assay Variation
| 6 | The precision and accuracy of the assay is determined for each of the four ceramides measured in LLOQ, LQC, MQC, HQC, and HLOQ.
| 7 | Intra-assay variation, precision, and accuracy is calculated for each of the ceramide in replicates of six (n = 6) at each QC concentration independently on three separate occasions.
| 8 | The intra-assay precision (percentage coefficient variance, %CV) and accuracy (percentage accuracy, %Accuracy) are calculated from the nominal concentrations according to the formulas reported in corresponding sheets and tables of this document.
| 9 | Recovery estimation is not necessary for the published method but only for different ones. Stability estimation is not necessary for the published method but only for different ones. Calibration line is prepared in 5% BSA.

## LC-MRM - Intra Assay QC Res
INTRA ASSAY - QC RESULTS
INTRA ASSAY - QC RESULTS | Table 12 LC-MRM Calibration Line 1 Data (Beginning). Please report the values of each of the three replicates separately for each compound and standard.
| Calibration Line 1 | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| STD 1 | 6238229.0128 | 6903427.3567 | 6950469.2018 | 7098547.848 | 7753889.3752 | 7877571.2428 | 87147649.6579 | 92687560.7336 | 93613476.6525 | 73479605.5366 | 81308303.413 | 81619964.089 | 284903.947 | 295256.5655 | 287465.7644 | 267696.2032 | 291079.1533 | 284777.7654 | 6087516.1429 | 6309437.6325 | 6285396.4273 | 1838414.6758 | 1920371.9479 | 1875851.4437
| STD 2 | 3308744.6377 | 3204919.9472 | 3199304.2623 | 3480150.5656 | 3605950.5724 | 3698598.8353 | 40866401.7793 | 38815049.2119 | 40018281.8684 | 37374053.4088 | 36022197.255 | 36091955.383 | 293647.003 | 293899.6294 | 310745.0326 | 268432.441 | 302715.75 | 304977.3578 | 5355096.1485 | 5602402.993 | 5428492.1328 | 1732081.8191 | 1755788.6576 | 1793895.2393
| STD 3 | 304585.3622 | 307149.5373 | 316478.0852 | 324580.892 | 338604.849 | 359040.6607 | 3445668.7461 | 3453512.2287 | 3462042.9984 | 3321000.6165 | 3143166.2807 | 3241448.061 | 293291.6948 | 293873.0905 | 308760.5697 | 280076.5382 | 284790.3703 | 307580.4981 | 5123336.4682 | 5152749.0295 | 5213699.2059 | 1707427.2442 | 1732349.0307 | 1689439.7476
| STD 4 | 67849.5961 | 64951.2646 | 69157.2294 | 76449.3069 | 70713.2163 | 75726.1704 | 703488.1976 | 730442.1832 | 711253.7785 | 701536.9953 | 694369.0239 | 678364.8928 | 290420.0421 | 302947.8215 | 300586.9638 | 272093.4308 | 287549.0945 | 311712.0364 | 4938769.2626 | 5248247.2469 | 5039751.9135 | 1700045.9114 | 1760339.8598 | 1739077.5504
| STD 5 | 35155.553 | 31125.5404 | 34601.0528 | 38212.8857 | 34567.9771 | 41461.1998 | 340531.8072 | 325935.6413 | 340213.7801 | 321323.3847 | 333401.8469 | 337295.822 | 290217.1382 | 290191.5641 | 306953.6232 | 266116.6556 | 274851.6621 | 297187.2152 | 4773893.0437 | 4966361.118 | 5147563.3648 | 1628456.7729 | 1685110.8148 | 1755092.0903
| STD 6 | 29234.7905 | 35134.4322 | 29710.6934 | 30743.3593 | 32969.1417 | 30392.1571 | 276421.6923 | 282661.3213 | 285096.2949 | 242999.4098 | 275230.8719 | 273895.4055 | 290389.386 | 290849.6211 | 305786.6961 | 278403.8317 | 286311.1073 | 293929.21 | 4900021.1974 | 5121714.8789 | 5137815.4713 | 1734655.2555 | 1689298.4461 | 1792114.0394
Table 13 LC-MRM Calibration Line 2 Data (End). Please report the values of each of the three replicates separately for each compound and standard.
| Calibration Line 2 | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| STD 1 | 6790728.3839 | 7334724.8254 | 7359838.6733 | 8042847.0426 | 8123667.8698 | 8724932.0466 | 92810163.5185 | 97245094.5024 | 95138225.9462 | 77486726.9584 | 80016694.5813 | 83578926.0541 | 306310.0837 | 336974.7863 | 304755.4654 | 292354.2687 | 296886.0282 | 311387.5491 | 6340861.1486 | 6452544.211 | 6225812.3861 | 1951893.6415 | 2118333.2597 | 1957270.0698
| STD 2 | 3463807.3403 | 3348286.3812 | 3411076.0924 | 3572862.8706 | 3842769.9648 | 3581256.707 | 42009061.0787 | 41764067.3852 | 41902400.8878 | 37617113.2103 | 37368702.2152 | 37503112.2938 | 316632.1141 | 319212.7195 | 325467.9845 | 273691.7151 | 297346.3715 | 293053.4519 | 5539320.9667 | 5840001.0122 | 5769198.3402 | 1881752.1864 | 1810948.7761 | 1851158.9342
| STD 3 | 326245.9548 | 321552.6742 | 321913.1536 | 381139.7836 | 352104.2343 | 368800.3772 | 3669419.2645 | 3385686.008 | 3552488.3305 | 3549784.3371 | 3388860.638 | 3442850.944 | 329699.658 | 330573.4712 | 318867.2981 | 297790.56 | 322162.6442 | 315463.3381 | 5347297.3169 | 5140361.0787 | 5469279.186 | 1860705.0622 | 1858890.2613 | 1891549.8148
| STD 4 | 71688.6056 | 73734.938 | 70427.8327 | 80862.8386 | 73714.246 | 78030.0328 | 739315.0385 | 775978.4604 | 753641.1691 | 702596.8428 | 751916.179 | 723445.3359 | 313374.7812 | 319200.1478 | 319044.1619 | 301438.7737 | 304719.6632 | 317978.2672 | 5244393.1063 | 5322052.8766 | 5560182.2423 | 1881088.7707 | 1807401.7957 | 1871795.205
| STD 5 | 33922.4891 | 35145.4718 | 34658.4656 | 37404.3761 | 34592.07 | 37522.3261 | 354470.0975 | 357693.372 | 376909.9318 | 341503.1884 | 348659.6691 | 369691.9901 | 298160.0936 | 298689.7447 | 329051.6881 | 291885.7413 | 294257.6948 | 302010.975 | 5061010.6544 | 5211471.8065 | 5374445.746 | 1755589.6434 | 1789122.7951 | 1855844.1361
| STD 6 | 26205.7196 | 27118.2534 | 27160.5368 | 31457.8109 | 31853.0145 | 32778.4659 | 296167.685 | 284426.8628 | 300005.534 | 274811.0291 | 278053.753 | 288215.9671 | 303370.0312 | 311276.0217 | 317760.8165 | 314364.7275 | 294832.1636 | 290903.8125 | 5177706.8721 | 5208035.5553 | 5419834.2425 | 1873410.1752 | 1761209.6895 | 1784461.4194
Table 14 LC-MRM Sample Data. Please report the values of each of the three replicates separately for each compound and standard.
| NIST SRM | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| SRM 1 | 339465.2767 | 338670.8516 | 334001.4656 | 183544.602 | 174681.6054 | 170510.4819 | 5847002.0189 | 5563779.0755 | 5635248.5079 | 1814163.954 | 1779961.3769 | 1748815.3024 | 246218.6958 | 241231.2884 | 237305.8959 | 304258.0704 | 289040.2043 | 301393.2821 | 6234104.5882 | 5727413.4296 | 5793272.6327 | 1859004.1522 | 1849097.089 | 1854504.8465
| SRM 2 | 321623.7495 | 310059.4679 | 322842.1711 | 165963.7913 | 162406.1026 | 163324.8087 | 5269252.0268 | 5032905.9102 | 5183632.3652 | 1580594.0889 | 1625250.0304 | 1536356.1655 | 251884.9351 | 251640.4902 | 251414.0847 | 313568.0748 | 311507.428 | 326314.8335 | 6060114.2001 | 6115462.4914 | 5592161.1094 | 1898226.9544 | 1887995.5294 | 1884616.5605
| SRM 3 | 327871.8764 | 334613.6167 | 330227.4022 | 167327.7229 | 171428.8938 | 165901.8703 | 5601539.165 | 5542735.1175 | 4990928.1501 | 1661061.5499 | 1630754.8306 | 1620826.3847 | 240071.4664 | 241635.3035 | 248692.9213 | 292131.3067 | 298924.7979 | 300755.2104 | 6194520.4734 | 6101542.1628 | 5784814.4855 | 1896854.0988 | 1897909.3938 | 1847231.3326
| SRM 4 | 316416.2567 | 323442.7682 | 320914.1855 | 162520.6662 | 174333.4021 | 166805.5283 | 5398538.2072 | 5257989.4382 | 5257696.5754 | 1578358.8294 | 1629772.2482 | 1602073.8788 | 249674.7742 | 259259.1493 | 242452.6252 | 314548.5856 | 321449.9428 | 309993.0284 | 6235230.4141 | 6215552.7473 | 5966513.9552 | 1826725.9761 | 1825879.2881 | 1842970.3038
| SRM 5 | 360027.1388 | 363033.8799 | 379418.3941 | 179782.3289 | 183388.6098 | 185064.1392 | 6044171.0973 | 6377920.0395 | 6087576.329 | 1803336.1977 | 1846494.4063 | 1830075.2023 | 261115.8251 | 257719.0228 | 252510.9831 | 312149.0478 | 306094.9698 | 306772.874 | 6319873.256 | 6356453.3101 | 6451167.5294 | 1943618.2894 | 1998483.1681 | 1944500.0678
| SRM 6 | 362605.8412 | 357288.8658 | 338884.685 | 172839.3085 | 171373.0553 | 167033.8367 | 6046700.3139 | 5939921.8265 | 5315140.5523 | 1758361.3583 | 1828775.8866 | 1674550.3313 | 266598.9215 | 263265.708 | 250058.4296 | 308775.7224 | 301993.8629 | 298722.5931 | 6703460.3759 | 6628178.2131 | 5650693.3884 | 2037978.1928 | 1970799.9271 | 1839103.5223
| NIST hTAG | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| hTAG 1 | 395198.8822 | 384274.4671 | 384013.6844 | 208732.0879 | 207366.4521 | 208416.0785 | 9843831.242 | 9225162.2682 | 9849574.8609 | 2515240.2899 | 2423180.063 | 2438630.0253 | 254596.2382 | 252909.6914 | 253108.7971 | 300460.0897 | 318070.2301 | 321609.3335 | 7317294.8858 | 6874897.7985 | 7262675.9392 | 2050637.1286 | 1912209.8357 | 2012913.5122
| hTAG 2 | 375449.0085 | 378166.1843 | 385022.6939 | 201113.619 | 203693.1798 | 192643.7524 | 9293889.2555 | 9352913.2568 | 8657420.752 | 2359209.2864 | 2365560.7802 | 2293270.0094 | 260729.3555 | 267289.7406 | 258884.4007 | 340969.8671 | 325090.5625 | 315312.0431 | 7530403.0656 | 7254001.0011 | 6788859.9567 | 1915410.5538 | 1969050.3526 | 1879944.2656
| hTAG 3 | 389091.924 | 395750.8945 | 381805.8576 | 211748.2706 | 204488.1629 | 196953.1498 | 9790237.6409 | 9178236.7776 | 9151914.4201 | 2463749.4335 | 2482187.0078 | 2417576.5535 | 247537.7332 | 258549.8024 | 262571.0529 | 330366.3436 | 328099.4131 | 318197.5167 | 7649737.1631 | 7464495.5397 | 7178078.6123 | 2015255.7665 | 2028392.9776 | 2001506.9195
| hTAG 4 | 403494.3843 | 379995.7056 | 380724.5269 | 194936.5381 | 200446.6362 | 202526.7315 | 9506120.4564 | 9607370.4113 | 9563533.1338 | 2334576.9887 | 2418327.9953 | 2451135.6158 | 259208.3482 | 259355.6004 | 255453.0883 | 318931.1563 | 323898.8234 | 323094.5884 | 7701182.8508 | 7373095.9938 | 7515969.0068 | 1969582.1312 | 2032035.8364 | 2046243.1182
| hTAG 5 | 377578.597 | 394272.1414 | 383127.3136 | 209258.65 | 210434.8443 | 198955.0927 | 10365440.4995 | 9674773.5724 | 9723430.0674 | 2698596.5796 | 2336312.2809 | 2361285.2788 | 263338.9192 | 267088.7795 | 252912.7252 | 331346.3853 | 332049.4114 | 313991.2416 | 7984097.552 | 7590240.5283 | 7294337.1493 | 2097890.0788 | 1993595.527 | 1979632.0306
| hTAG 6 | 396289.8207 | 397529.6962 | 389627.4047 | 202800.1173 | 196053.7972 | 183599.6635 | 9721759.1589 | 10052854.7131 | 9901033.211 | 2473981.8262 | 2437138.7159 | 2402725.8176 | 259884.547 | 261588.9336 | 268080.0645 | 337273.6184 | 306384.2117 | 297769.3485 | 7614844.7728 | 7557047.3949 | 7585604.5375 | 2115283.8306 | 2010107.8507 | 1986916.4813
| NIST T1D | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| T1D 1 | 310397.6003 | 301627.0111 | 311642.3914 | 211933.0071 | 201802.7709 | 197568.8198 | 5274452.0391 | 5495870.2606 | 5583169.2884 | 1719681.6727 | 1673893.8015 | 1713248.5146 | 255254.2394 | 263304.7625 | 260922.554 | 319524.4295 | 309513.6721 | 300178.259 | 6075652.167 | 6157310.0263 | 6183576.8619 | 1917513.4264 | 1953408.2289 | 2060097.1718
| T1D 2 | 318659.337 | 329998.4685 | 298215.1843 | 215444.9323 | 209210.2838 | 198092.3875 | 5839648.5307 | 5455459.663 | 5493325.2804 | 1683425.3062 | 1725153.8931 | 1607010.522 | 279050.3741 | 275559.6917 | 265740.6903 | 333265.5695 | 322408.9945 | 326178.2652 | 6859861.1826 | 6449536.4317 | 6289788.1404 | 2051697.7467 | 2034071.0626 | 1912443.0274
| T1D 3 | 325113.7927 | 293692.2494 | 311864.9631 | 193586.6658 | 193512.5953 | 193816.9394 | 5631590.65 | 5187534.9182 | 5389910.4773 | 1729199.7784 | 1612491.339 | 1680706.9815 | 271969.7355 | 264405.1875 | 282956.2181 | 321431.9925 | 329659.9068 | 322732.0106 | 6622338.649 | 6220726.6175 | 6423228.3127 | 2065716.9586 | 1952317.0994 | 2068690.0853
| T1D 4 | 318237.461 | 302902.0339 | 318649.7562 | 206497.4138 | 204787.5814 | 200424.0085 | 6019521.5353 | 5321257.2308 | 5875054.5232 | 1733841.169 | 1624911.8582 | 1772538.2784 | 288357.1568 | 267980.6264 | 277822.7446 | 329467.6331 | 319736.1345 | 337098.7319 | 6691257.8366 | 6238650.2936 | 6791047.406 | 2139222.0074 | 1995111.8626 | 2133950.5624
| T1D 5 | 304693.9952 | 330049.3743 | 308697.0805 | 209537.9321 | 190760.6184 | 191530.3113 | 5411582.7454 | 5274960.5056 | 5296330.3994 | 1650600.2292 | 1632220.8866 | 1629296.4492 | 258273.8165 | 272288.4632 | 269029.4052 | 316966.8922 | 321248.0609 | 315192.9192 | 6232192.5317 | 6237096.403 | 6135278.0368 | 1984779.4482 | 1923076.311 | 1929939.8526
| T1D 6 | 311105.4568 | 299034.6851 | 301496.0464 | 192415.9058 | 191450.2527 | 196054.4752 | 5743060.5929 | 5227530.4256 | 5173436.9901 | 1698313.4049 | 1555515.2335 | 1622668.4853 | 280232.6584 | 259130.9345 | 265871.8148 | 323054.7162 | 316632.1128 | 335126.0985 | 6720857.7059 | 6045080.6571 | 6118107.0171 | 2065412.8886 | 1921868.5011 | 1992178.1451
| NIST Young AA | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| yA-A 1 | 250816.5424 | 256622.6622 | 243032.0401 | 111828.3902 | 102787.3773 | 107700.3147 | 3779972.1442 | 3625803.1444 | 3472165.1893 | 1098825.3759 | 1124044.4663 | 1121196.7259 | 275418.1272 | 303595.8375 | 275951.6116 | 326682.6184 | 330963.7599 | 323882.7833 | 6061251.6696 | 5734075.8868 | 5654822.3306 | 2128027.9074 | 2054174.7483 | 2095036.8438
| yA-A 2 | 236700.5883 | 241456.6331 | 220993.5135 | 98076.7768 | 104665.3115 | 100767.1864 | 3457000.7947 | 3584861.5727 | 3293128.3448 | 1038452.0787 | 1068082.7006 | 1042929.1325 | 291354.4278 | 295377.7026 | 283939.0429 | 321673.6734 | 323974.978 | 328194.6117 | 5817762.4073 | 6300629.2198 | 5495569.0363 | 1965208.5011 | 2132921.7232 | 2063021.9176
| yA-A 3 | 266708.2373 | 254306.6582 | 266833.7474 | 112134.0971 | 114891.4056 | 102954.8974 | 4207346.3818 | 4150017.0473 | 3967605.7514 | 1299379.6707 | 1243839.675 | 1308117.8571 | 287949.4163 | 274950.6699 | 286207.8731 | 327235.0118 | 305462.5518 | 326274.3924 | 6088024.6814 | 6033801.3907 | 6143487.141 | 2149098.5322 | 2128718.757 | 2100943.3162
| yA-A 4 | 229655.9509 | 222767.7354 | 216738.2924 | 104593.2413 | 100931.6061 | 96379.7805 | 3068028.1092 | 3163521.7114 | 3316599.0222 | 1020276.7844 | 1017946.8419 | 1098197.0899 | 286547.1136 | 304098.2796 | 267952.7362 | 334060.4154 | 332825.5061 | 324946.8359 | 5909764.7581 | 5677110.7951 | 5904378.7167 | 1983761.165 | 1969891.4815 | 2078432.3046
| yA-A 5 | 226784.7947 | 226993.8269 | 244840.997 | 109941.4899 | 103872.5144 | 102890.8485 | 3449017.1437 | 3258291.9591 | 3294567.5973 | 1029017.7665 | 1102910.8604 | 1089255.6835 | 272255.8495 | 282324.2525 | 272904.2665 | 325066.2748 | 330930.729 | 326509.6483 | 5774216.5998 | 5613481.1341 | 5679751.3844 | 1999751.5429 | 2010718.5653 | 1953815.3186
| yA-A 6 | 228573.0204 | 227756.561 | 225071.472 | 98213.1988 | 105081.8966 | 97248.3883 | 3460590.3562 | 3441414.6158 | 3382845.1045 | 1168272.1768 | 1051920.7047 | 1025557.3199 | 296181.5962 | 291195.3862 | 273550.6752 | 320669.8384 | 330178.8207 | 322436.7853 | 6056117.8789 | 5862190.9044 | 5651442.1039 | 2104803.1879 | 2074890.456 | 1926200.8672
Table 15 LC-MRM Intra Assay QC Data. Please report the values of each of the three replicates separately for each compound and standard.
| Lowest Level QC | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| LLQC 1 | 79689.9708 | 86296.84 | 90174.3255 | 38795.1489 | 38852.6142 | 31191.8877 | 1072050.0048 | 1030835.0345 | 1050676.6558 | 333151.1486 | 345034.8072 | 354256.9083 | 333853.8934 | 328066.3145 | 339464.5161 | 339373.3514 | 330077.842 | 308110.2263 | 5708384.8908 | 5723956.0425 | 5767313.0284 | 1880862.2067 | 1889280.1664 | 1969743.0701
| LLQC 2 | 84672.0562 | 94394.4172 | 85798.4644 | 35606.0228 | 35402.079 | 40657.7544 | 1122124.3341 | 1088066.7953 | 1064218.3771 | 365984.5487 | 351300.3423 | 352031.4614 | 340061.4087 | 344911.7397 | 319779.1689 | 297588.3667 | 311282.5784 | 345974.0963 | 5702749.103 | 5815387.3308 | 5701451.8045 | 2057149.5147 | 2055660.1803 | 1915272.0272
| LLQC 3 | 86836.8849 | 82166.3419 | 81847.3725 | 39653.5409 | 38238.6297 | 36263.561 | 1051925.7653 | 1041983.7486 | 965746.7842 | 344814.6302 | 333166.2207 | 314799.4852 | 312227.0953 | 332360.9015 | 314088.5449 | 334964.8896 | 343228.2081 | 304091.4221 | 5621909.0573 | 5602555.6123 | 5159842.1467 | 1973493.4489 | 1913564.3351 | 1831356.3252
| LLQC 4 | 88498.8937 | 91053.1884 | 86684.0355 | 36411.1667 | 34771.7979 | 40766.0584 | 982182.6503 | 998294.8681 | 981581.7925 | 326228.3561 | 305818.5112 | 313187.2576 | 337995.2231 | 329804.5458 | 332640.6378 | 320645.9001 | 332164.9023 | 332728.9249 | 5476617.0011 | 5542666.4586 | 5480362.0772 | 1906244.6862 | 1857212.3403 | 1907251.654
| LLQC 5 | 78581.4617 | 77092.9003 | 72639.0405 | 34155.5335 | 32941.7125 | 30947.235 | 918191.2626 | 947373.4415 | 885551.9466 | 300300.7232 | 301149.9318 | 284723.7132 | 360662.8199 | 335737.6245 | 338674.4256 | 341049.4438 | 336815.1912 | 315199.8705 | 5548608.8754 | 5559979.9304 | 5396035.0903 | 1952505.721 | 1918739.1233 | 1919490.7602
| LLQC 6 | 79649.1462 | 75629.4235 | 79418.7121 | 29762.4491 | 33482.9962 | 29685.4547 | 967066.3693 | 922670.5053 | 961702.6726 | 315537.2299 | 307244.343 | 308969.6316 | 339003.2882 | 339822.128 | 328972.2593 | 315627.421 | 332747.3834 | 323187.4725 | 5548315.9663 | 5397999.3774 | 5645411.0206 | 1921997.2047 | 1918091.2553 | 1884539.5827
| Low QC | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| LQC 1 | 164204.8634 | 165387.8079 | 162335.4402 | 77831.3301 | 83308.8701 | 71435.0174 | 2393442.5547 | 2318376.682 | 2239612.9005 | 713721.1568 | 741791.6734 | 707890.7538 | 315377.7506 | 322986.0343 | 301383.6806 | 362715.8228 | 342361.7161 | 315835.6827 | 6169742.4955 | 5966809.3935 | 5653126.2202 | 2048142.993 | 2005720.1914 | 1988814.0239
| LQC 2 | 159942.5808 | 165084.1806 | 157624.5674 | 74208.5567 | 74023.9016 | 76784.3692 | 2335653.916 | 2208342.2475 | 2208826.4612 | 715125.0278 | 744717.1525 | 704222.2442 | 313497.9464 | 317969.5282 | 295026.8128 | 331084.2728 | 342678.2308 | 343307.9102 | 5823864.188 | 5731331.3604 | 5767784.1823 | 1988890.0248 | 1983006.0367 | 1948613.4852
| LQC 3 | 158343.0017 | 163859.8598 | 159042.6916 | 70484.0526 | 76547.8677 | 73062.6165 | 2238829.5129 | 2215657.388 | 2275483.0892 | 702883.4417 | 668834.8685 | 714345.5641 | 321169.7221 | 312175.6964 | 321755.7251 | 336637.3203 | 341651.3903 | 326331.0648 | 5885669.4769 | 5986066.0406 | 5998240.0901 | 1949298.7262 | 1896763.9562 | 2104295.262
| LQC 4 | 171339.889 | 164343.6875 | 161978.0859 | 79362.7385 | 71744.5673 | 73211.2849 | 2310643.6832 | 2053326.5537 | 2071398.0481 | 700191.4635 | 683213.7173 | 684992.4875 | 329749.2292 | 306628.4656 | 301882.1682 | 333932.1339 | 335684.376 | 327456.6167 | 5861196.0989 | 5381258.2078 | 5350193.5717 | 1977975.8383 | 1876665.3637 | 1885385.0713
| LQC 5 | 171459.9645 | 164553.2803 | 158622.6436 | 80970.8058 | 74125.3626 | 67626.8357 | 2318810.8007 | 2270784.3689 | 2077031.7659 | 719715.7953 | 695264.6777 | 697689.1518 | 310950.8387 | 303663.9649 | 299046.2647 | 334921.2468 | 333979.4419 | 289705.1105 | 5936038.2549 | 5688016.4109 | 5206390.9908 | 1938242.7619 | 1944984.5687 | 1794474.9706
| LQC 6 | 157414.1322 | 156331.3739 | 154593.009 | 63007.9987 | 74620.7687 | 70655.3347 | 2106136.3241 | 2197783.1405 | 2128134.0954 | 657873.7375 | 665643.8596 | 673741.2014 | 290949.7413 | 310278.0106 | 317295.03 | 317117.2899 | 333226.1176 | 325555.8711 | 5732218.4344 | 5839521.2719 | 5815079.0934 | 1898703.8971 | 1912712.43 | 1911866.7387
| Middle QC | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| MQC 1 | 287734.5612 | 291832.4138 | 285930.973 | 158218.0011 | 145611.8966 | 142391.4493 | 4964300.1588 | 4898263.3338 | 4483335.5574 | 1420516.6797 | 1476310.041 | 1410529.8159 | 287642.6738 | 272422.1071 | 271262.2736 | 342126.6942 | 315251.2722 | 304398.341 | 6595962.5156 | 6580564.4212 | 5712877.0551 | 2012135.1457 | 2031845.9507 | 1873846.2285
| MQC 2 | 295889.5897 | 301426.9902 | 305487.5739 | 138902.7811 | 142061.3492 | 128792.2789 | 5294745.6843 | 5017390.5913 | 4382525.789 | 1489165.8803 | 1476226.8372 | 1410001.2713 | 282991.7101 | 275325.3221 | 271957.0679 | 305922.1054 | 292326.3866 | 286972.8325 | 6587382.3025 | 6178148.4002 | 5271572.9978 | 2024192.1043 | 2037792.7162 | 1917259.5583
| MQC 3 | 297967.0943 | 311793.8105 | 287131.5795 | 146702.0234 | 138061.0481 | 138538.6525 | 5056492.7052 | 5310733.4525 | 4852285.8357 | 1509432.1166 | 1501406.5343 | 1407493.6215 | 292524.1602 | 292810.0051 | 278138.0205 | 311454.7119 | 299583.664 | 285383.1401 | 6488398.6 | 6172256.3548 | 6225109.4049 | 2035595.7973 | 1969353.4214 | 1908442.4087
| MQC 4 | 307707.0285 | 313855.1261 | 306734.598 | 160578.5057 | 150143.087 | 141747.6349 | 5356466.9234 | 5229088.3278 | 5534455.3606 | 1512622.905 | 1546180.8369 | 1655772.9833 | 275373.2526 | 292834.8848 | 304590.2353 | 337421.0919 | 342690.8556 | 331900.3362 | 6574511.1804 | 6588282.7159 | 6625361.0416 | 2022297.0989 | 2035922.5515 | 2162929.8811
| MQC 5 | 304438.1035 | 299395.6436 | 283817.1574 | 158049.1709 | 155014.8995 | 148853.3415 | 5400506.7984 | 5196651.6973 | 4420376.0612 | 1578378.7391 | 1558726.5575 | 1418044.5337 | 289282.218 | 280907.706 | 275918.6294 | 350039.5725 | 350682.4645 | 325767.5426 | 6649669.058 | 6248333.0361 | 5377483.4478 | 2074032.5923 | 2039788.1617 | 1963788.2366
| MQC 6 | 304457.985 | 292870.043 | 306691.3062 | 151379.773 | 153815.9685 | 150042.7881 | 5287409.2562 | 5275787.1963 | 5044835.9141 | 1532777.765 | 1528761.9877 | 1475078.7001 | 287475.3145 | 282380.1737 | 285186.0726 | 335702.6582 | 333203.835 | 333332.6527 | 6593612.1668 | 6622290.0504 | 6307678.3109 | 1950559.5933 | 2068314.1746 | 2048657.3752
| High QC | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| HQC 1 | 421837.2747 | 432297.3111 | 440830.8221 | 346101.3011 | 355873.8097 | 354842.1843 | 7187660.8477 | 7109313.0328 | 7110991.6991 | 3266183.9815 | 3372202.6348 | 3334027.6021 | 288300.1577 | 292225.4601 | 295621.8263 | 341832.9303 | 333076.8075 | 353831.8199 | 6720056.916 | 6395595.207 | 6325593.3523 | 2023164.2758 | 2070649.5043 | 2059612.9802
| HQC 2 | 451076.7501 | 448888.6843 | 443645.01 | 380324.3954 | 370202.6754 | 359127.7635 | 8490770.534 | 8188508.8307 | 6638814.6915 | 3758180.1219 | 3689453.9712 | 3488002.5974 | 297334.0697 | 296494.776 | 291966.9967 | 351469.7575 | 366996.756 | 314559.2062 | 7326650.2688 | 7222469.1673 | 5860069.1719 | 2170555.5109 | 2206968.5612 | 2042385.5878
| HQC 3 | 445480.1012 | 435040.9373 | 452348.5819 | 359344.9792 | 362122.4694 | 370335.1943 | 7034136.9854 | 7585923.3156 | 7227056.2883 | 3460661.6062 | 3450915.2993 | 3500786.8157 | 304218.4891 | 279652.8356 | 292925.0106 | 350403.2364 | 353147.1398 | 357855.7289 | 6621483.9242 | 6727630.6361 | 6443622.1616 | 2035672.6619 | 2036194.0689 | 2110504.1358
| HQC 4 | 459973.1225 | 439688.8741 | 435249.4047 | 374730.3673 | 350361.8647 | 320750.4401 | 7541396.9795 | 7471856.0798 | 6764794.8102 | 3467023.0224 | 3488686.5652 | 3248714.0157 | 285044.6449 | 284203.9331 | 280910.0381 | 334974.6394 | 336285.872 | 335068.0279 | 6779606.2723 | 6557860.4619 | 5830798.1324 | 2041487.2125 | 1989439.4722 | 2000320.275
| HQC 5 | 441663.562 | 449069.3544 | 444342.3615 | 354339.2254 | 352549.9574 | 357246.0082 | 7875754.2084 | 7702269.9873 | 7541252.2418 | 3392279.7325 | 3459358.9706 | 3413183.0263 | 289901.6987 | 301444.4042 | 295298.7881 | 343682.2558 | 342491.0276 | 343815.2743 | 7030862.8635 | 7111436.7768 | 7029298.7203 | 2141299.8703 | 2095233.0018 | 2147371.0258
| HQC 6 | 436174.7902 | 444993.0467 | 427594.9201 | 364239.0336 | 358434.8736 | 302866.2367 | 7747977.9999 | 7378568.1973 | 5723340.8843 | 3514804.4242 | 3373673.5391 | 3115620.709 | 285132.8722 | 292015.4538 | 275659.4339 | 357730.1287 | 349877.6193 | 292685.244 | 6851324.1604 | 6411261.1711 | 4982937.2749 | 2115183.9107 | 2040519.1577 | 1847078.8188
| Highest Level QC | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| HLQC 1 | 579111.3369 | 583376.0697 | 573670.0519 | 531492.0808 | 549551.6613 | 485517.8121 | 9351081.1074 | 9863222.3501 | 8575556.1975 | 5352683.1416 | 5292517.3447 | 5256001.8213 | 287510.414 | 284998.3398 | 284282.6733 | 323800.1701 | 328031.6116 | 307904.5935 | 6191180.2288 | 6467577.0282 | 6072540.274 | 2027907.3787 | 1955364.6171 | 2049904.943
| HLQC 2 | 594309.7669 | 591583.2389 | 571192.8187 | 533925.3743 | 534859.6866 | 579715.9327 | 9825310.6122 | 10113402.5109 | 9681688.1249 | 5269053.2577 | 5587958.3764 | 5316569.3115 | 295285.4446 | 292930.8837 | 289442.4493 | 315119.4017 | 319429.4227 | 351636.618 | 6910222.2011 | 7012415.2043 | 6445561.7132 | 2134557.4566 | 2136819.5492 | 2143806.4208
| HLQC 3 | 565234.2838 | 579047.0642 | 566351.6041 | 559293.1833 | 525071.2726 | 546206.2212 | 9699832.008 | 9104162.9435 | 9077130.5071 | 5269558.3663 | 5286806.4705 | 5244133.7185 | 288358.9746 | 289454.347 | 285189.2822 | 347195.2419 | 340731.7441 | 337304.8132 | 6601821.797 | 6272928.4835 | 6075267.1329 | 1984521.1192 | 1991082.3376 | 1977607.1164
| HLQC 4 | 606291.1447 | 616021.7674 | 601779.3458 | 574822.5012 | 572527.3025 | 595049.408 | 10097979.8268 | 10333976.7178 | 10507664.7119 | 5529580.722 | 5613988.2271 | 5535502.744 | 296359.8806 | 292568.0315 | 296772.5403 | 347356.8464 | 361101.6518 | 354880.319 | 6811755.8791 | 7141145.4038 | 6854582.9797 | 2101139.4311 | 2156356.3502 | 2159211.0158
| HLQC 5 | 586529.9586 | 583759.9807 | 581919.5849 | 517539.3212 | 570894.227 | 571485.6687 | 9713253.2329 | 9782036.8587 | 9620617.5163 | 5282442.456 | 5518429.874 | 5443896.6692 | 309620.7333 | 294734.6807 | 284267.8032 | 321523.3873 | 356174.1974 | 368430.8546 | 6962414.8719 | 6917080.6429 | 6806071.9165 | 2089775.0733 | 2117558.0724 | 2135455.4736
| HLQC 6 | 591229.6763 | 581671.4569 | 572952.0676 | 494703.7311 | 553256.1179 | 552246.7668 | 10294860.7506 | 10229610.6122 | 9565911.7469 | 5368788.0704 | 5371160.7884 | 5441798.3095 | 294949.9427 | 298256.2545 | 292228.0385 | 306901.6436 | 358237.7556 | 335274.1628 | 7024332.4821 | 7014515.968 | 6635389.3175 | 2077181.0375 | 2076824.3806 | 2121174.5454
| Blank QC | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| Blank | 2185.3102 | 2758.1202 | 3103.5235 | 834.7428 | 1092.0818 | 559.3388 | 4594.1614 | 4916.0736 | 4027.9966 | 390.3044 | N/A | 704.3229 | N/A | N/A | N/A | N/A | N/A | N/A | 493.6808 | N/A | 555.2228 | N/A | N/A | 218.6129
| Blank | 5389.7746 | 5244.1561 | 4374.9975 | 3374.0851 | 4147.8351 | 2247.2701 | 19612.0193 | 10341.2245 | 9657.9033 | 9807.7538 | 5407.2095 | 3831.4044 | N/A | N/A | N/A | N/A | 135.1942 | N/A | N/A | N/A | N/A | 569.8981 | 174.9673 | N/A
| Blank | 4374.9975 | 7247.3777 | N/A | 2247.2701 | 6626.1288 | 1193.4904 | 9657.9033 | 54796.5899 | 13888.7312 | 3831.4044 | 49934.3265 | 8566.7948 | N/A | N/A | N/A | N/A | 219.1989 | N/A | N/A | N/A | 849.1149 | N/A | 904.7407 | 295.5787
| Note for Blank | Refer to 'blue' highlights in Main raw peak areas'
| Note for intra-assay QC series | a 2nd run was repeated (with same extract) on a separate day (refer to Run 2xxxx tabs)

## Run 2 LC-MRM-Intra Assay dQCs
INTRA ASSAY - QC RESULTS
INTRA ASSAY - QC RESULTS | Table 15 LC-MRM Intra Assay QC Data. Please report the values of each of the three replicates separately for each compound and standard.
| Lowest Level QC | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| LLQC 1 | 78234.1246 | 73303.1245 | 75444.4749 | 27460.2743 | 23301.0379 | 27665.3008 | 1220777.8614 | 1185731.9516 | 1201879.1553 | 367258.1097 | 364723.563 | 360522.0803 | 273483.4624 | 289990.2734 | 284520.8528 | 226004.9266 | 230660.0769 | 235622.4606 | 6044300.1552 | 5878094.5345 | 6051803.3274 | 1931640.0908 | 1976932.8398 | 1912400.0546
| LLQC 2 | 72895.3238 | 76121.9462 | 82289.8885 | 25346.5673 | 27784.7047 | 26169.5871 | 1198431.6914 | 1217823.5076 | 1206326.8142 | 367733.575 | 359541.1305 | 357837.0102 | 277853.0066 | 285395.6283 | 277659.0909 | 229236.9215 | 227375.5661 | 228281.5746 | 5929123.9395 | 5997849.746 | 5714819.7353 | 1880214.9997 | 1851260.0019 | 1887389.5762
| LLQC 3 | 74785.98 | 72362.9257 | 72777.5527 | 24247.8868 | 24499.3589 | 25741.6082 | 1186067.661 | 1182112.2175 | 1208273.8448 | 363858.0676 | 350906.2265 | 352132.5578 | 265056.7273 | 279571.2915 | 274391.4531 | 229232.5217 | 233021.5226 | 217221.5451 | 5877135.2316 | 6066153.0297 | 5975577.0654 | 1913505.7481 | 1920346.8496 | 1922498.1059
| LLQC 4 | 76755.0655 | 67512.4496 | 75425.4349 | 25652.4915 | 27173.2099 | 24207.6686 | 1151989.3127 | 1126454.333 | 1144088.33 | 342700.3738 | 344640.4792 | 345864.4524 | 272146.7917 | 271830.2847 | 273110.1462 | 221560.8846 | 228399.6677 | 229227.1082 | 5954026.8584 | 6031490.421 | 5845368.4493 | 1889710.3773 | 1878750.8138 | 1863768.7187
| LLQC 5 | 67739.0582 | 66572.4391 | 69657.3016 | 23933.3813 | 23740.8106 | 23642.6892 | 1078884.5054 | 1046463.4515 | 1043178.5373 | 320129.7745 | 305759.7913 | 317074.1917 | 266462.5798 | 275670.9326 | 264308.8693 | 228071.9914 | 224337.8953 | 219354.5913 | 6036313.3324 | 5808900.2556 | 5868128.6392 | 1928579.9294 | 1870685.7261 | 1911726.2093
| LLQC 6 | 69035.3116 | 68617.1696 | 74517.5135 | 24768.2713 | 24554.6935 | 27158.9864 | 1114784.3791 | 1075977.8076 | 1115469.8401 | 334724.3396 | 321641.3104 | 328291.5155 | 272116.5239 | 280597.172 | 276072.9229 | 233322.9255 | 228820.5957 | 236859.748 | 5975774.9449 | 5900548.713 | 5993792.1935 | 1918451.4238 | 1879453.231 | 1897024.5917
| Low QC | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| LQC 1 | 173205.2575 | 174343.0596 | 174857.9611 | 64737.0771 | 62448.6803 | 60992.1701 | 2782118.2373 | 2755659.7918 | 2674842.1449 | 799827.0214 | 764814.6632 | 766174.8416 | 292079.5422 | 311124.9627 | 299469.1086 | 245468.7978 | 252359.2652 | 240224.364 | 6669014.9795 | 6490008.3622 | 6411342.661 | 2010465.1884 | 1973772.6443 | 1927292.5134
| LQC 2 | 167021.7946 | 168442.8296 | 166519.7115 | 60810.7559 | 55925.4794 | 55022.5863 | 2679254.805 | 2608159.5367 | 2646291.3133 | 770120.4804 | 753344.3745 | 766227.4237 | 297008.7101 | 295457.2468 | 293092.2072 | 241580.4727 | 239007.34 | 244284.1376 | 6428344.2422 | 6198189.417 | 6292280.1041 | 1968653.0939 | 1929159.7321 | 1938779.7733
| LQC 3 | 165178.6678 | 154798.9167 | 160074.187 | 53454.8282 | 51385.0016 | 53652.3762 | 2610240.2705 | 2580556.6392 | 2515721.1512 | 734834.9727 | 727956.2529 | 736013.2465 | 298386.9763 | 288244.8191 | 291287.9622 | 248975.6153 | 242438.1339 | 247304.8054 | 6362766.9711 | 6342127.1373 | 6176804.5804 | 1975954.0745 | 1915166.9986 | 1940851.859
| LQC 4 | 170594.3287 | 168621.1184 | 169803.7351 | 56664.7618 | 55440.5835 | 60360.2353 | 2702746.7752 | 2698065.3857 | 2699150.933 | 756908.155 | 769756.3039 | 776261.0773 | 300095.8565 | 287980.0347 | 295232.9775 | 248372.265 | 251280.9284 | 250732.6961 | 6422152.7282 | 6437437.9105 | 6328566.1005 | 1973061.6269 | 1988347.4277 | 1957736.6626
| LQC 5 | 171222.4649 | 165867.6043 | 172871.2463 | 58215.6941 | 58745.2679 | 58603.9043 | 2772610.4758 | 2677256.5155 | 2609855.7527 | 805822.1737 | 759322.0719 | 754852.2141 | 294478.5662 | 290379.3896 | 294092.3101 | 245498.9973 | 239132.2079 | 247721.5054 | 6364959.0198 | 6075908.7783 | 5949742.844 | 1959274.3462 | 1896477.4259 | 1846620.578
| LQC 6 | 166407.1178 | 161883.3482 | 162267.0912 | 58246.1025 | 54331.9132 | 54094.0233 | 2512161.4155 | 2507653.874 | 2520675.0413 | 748364.3914 | 741365.3916 | 735079.5467 | 294080.8084 | 302283.0001 | 291481.9796 | 245011.1942 | 236362.6612 | 252093.5258 | 6214863.2221 | 6127640.4916 | 6332252.5832 | 1923638.6627 | 1944366.5161 | 1902273.0927
| Middle QC | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| MQC 1 | 341092.1229 | 354598.1001 | 351749.1324 | 116319.9751 | 123724.4929 | 119268.7075 | 6243467.035 | 6253192.0168 | 6003948.0826 | 1665114.4299 | 1613008.9926 | 1583894.2745 | 332064.9316 | 318729.649 | 320604.1079 | 271960.38 | 261545.6998 | 269131.8235 | 7241337.019 | 7310393.5195 | 7182780.2276 | 2081547.4001 | 2101650.9531 | 2046734.7562
| MQC 2 | 368206.9927 | 369264.5025 | 359059.1146 | 130265.4231 | 126776.8476 | 127271.4473 | 6669784.2597 | 6850568.7643 | 6628524.1857 | 1780153.7312 | 1753695.4839 | 1724571.26 | 328958.1623 | 319733.7322 | 315855.5644 | 281636.7087 | 268039.3615 | 264846.1363 | 7770983.9265 | 7788553.5469 | 7565977.4929 | 2216006.716 | 2120401.1955 | 2178276.434
| MQC 3 | 363933.5074 | 354772.4389 | 360164.5663 | 133768.2699 | 134389.4414 | 123330.7965 | 6462779.3646 | 6486059.7981 | 6378899.7176 | 1739502.7294 | 1720241.344 | 1696589.7962 | 310613.4447 | 324470.2332 | 335860.5312 | 258212.3483 | 272488.7447 | 267260.1772 | 7593171.4988 | 7598719.0886 | 7330873.2346 | 2124366.0503 | 2120153.2642 | 2052433.059
| MQC 4 | 376295.3907 | 364344.518 | 346420.5646 | 131202.3558 | 128897.3244 | 123698.9785 | 6711295.0455 | 6656954.6045 | 5716891.3374 | 1765891.0073 | 1665687.6138 | 1587466.1683 | 327922.9864 | 327315.6431 | 304599.2595 | 273177.7099 | 261715.1454 | 253683.7279 | 7591540.1327 | 7392832.3386 | 6682775.0943 | 2174318.5559 | 2100413.4295 | 1949786.7656
| MQC 5 | 367289.0521 | 367065.9208 | 358434.7855 | 124759.5288 | 126834.3826 | 122003.4301 | 6505165.6678 | 6465216.0857 | 6009201.4125 | 1698893.7059 | 1699037.3172 | 1613254.9044 | 324130.6939 | 323204.707 | 319021.5749 | 253485.3747 | 267579.6395 | 249968.6367 | 7554411.5731 | 7557568.1263 | 7026145.6443 | 2161053.1964 | 2150765.7015 | 2011025.7538
| MQC 6 | 373648.8656 | 360000.8898 | 365968.243 | 133364.9697 | 123410.3209 | 121188.512 | 6355319.7966 | 6333608.0624 | 6214426.7953 | 1770045.3872 | 1681124.4691 | 1674899.6483 | 336653.6712 | 313503.0964 | 311552.2685 | 263766.2455 | 265269.1428 | 251654.1848 | 7460001.4181 | 7239501.0776 | 7205741.2345 | 2160604.6737 | 2069246.6721 | 2076744.5421
| High QC | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| HQC 1 | 525233.8486 | 508072.6017 | 504784.59 | 290361.3445 | 283578.5074 | 285459.4003 | 8821755.3745 | 8832226.1498 | 8328377.6064 | 3683882.247 | 3775224.3409 | 3768108.1395 | 323512.4639 | 336514.2532 | 324126.7355 | 267468.6803 | 268206.3028 | 253774.0681 | 7473558.4859 | 7605783.9423 | 7473685.3977 | 2172248.5167 | 2190058.4726 | 2134904.2898
| HQC 2 | 562979.0195 | 561808.9479 | 576787.6797 | 313838.4324 | 323316.8253 | 318540.6212 | 10833054.4396 | 10402953.2659 | 10440972.6305 | 4268415.9848 | 4271482.0405 | 4196883.4249 | 346275.0817 | 348584.9952 | 336266.5093 | 283429.3963 | 286056.1191 | 272264.2222 | 8540794.2256 | 8390728.1315 | 8487952.6616 | 2358509.4645 | 2340972.2294 | 2305373.3254
| HQC 3 | 533767.6253 | 519572.043 | 517427.3562 | 288682.0441 | 274621.9127 | 280630.227 | 9547452.6017 | 8968311.8009 | 9025589.1382 | 3909784.9175 | 3623081.7397 | 3806952.0053 | 327967.8958 | 317922.1327 | 324054.8595 | 268013.9264 | 258739.8394 | 263840.9553 | 7947084.2803 | 7213472.7353 | 7339075.7803 | 2244427.3968 | 2058452.61 | 2095264.6778
| HQC 4 | 551418.2746 | 519211.8626 | 534682.3556 | 299737.5868 | 285272.5746 | 290803.3923 | 9709260.159 | 9224030.5114 | 9288965.8741 | 3947170.8336 | 3784864.5427 | 3814463.3181 | 323445.6945 | 323969.8686 | 323550.1324 | 262303.0805 | 255107.2325 | 267767.8017 | 7880154.5202 | 7567342.485 | 7422618.1527 | 2230530.4423 | 2106030.1731 | 2125467.295
| HQC 5 | 550794.4691 | 542796.3477 | 533834.2123 | 300223.2353 | 295332.5062 | 285126.493 | 9686823.0037 | 9660339.0503 | 9779150.9947 | 3950209.9224 | 4000418.991 | 3903916.7321 | 332447.7966 | 337375.3184 | 333389.8984 | 270094.3275 | 268336.6861 | 267923.0448 | 8067027.0965 | 7917781.2747 | 7884048.96 | 2249374.735 | 2253080.642 | 2200964.2041
| HQC 6 | 537186.0466 | 532091.8522 | 528079.4035 | 295415.8807 | 288488.2766 | 282860.3002 | 9603751.225 | 9256493.5713 | 9110779.3763 | 3950004.4646 | 3802759.7024 | 3771767.1707 | 321131.1736 | 314786.3114 | 313677.2099 | 263526.2441 | 255322.5778 | 256819.9384 | 7790303.9786 | 7507708.0541 | 7289031.3697 | 2150385.6835 | 2092363.2156 | 2065005.7127
| Highest Level QC | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| HLQC 1 | 694504.0234 | 683277.8127 | 678033.1463 | 462328.9142 | 462031.1534 | 433499.1349 | 12030145.3369 | 11982722.7637 | 11502069.6303 | 6067357.1822 | 6063512.0659 | 5925912.6649 | 318925.2246 | 320935.9141 | 320007.7853 | 271572.3335 | 264419.6272 | 250918.2243 | 7601796.6584 | 7703858.8608 | 7368183.546 | 2115130.9433 | 2149183.887 | 2045118.7421
| HLQC 2 | 703072.5124 | 696408.2955 | 713990.7741 | 458300.7985 | 464302.7417 | 459978.9975 | 12476930.7951 | 12276134.5869 | 12362443.9918 | 6103173.2452 | 6253316.1839 | 6262494.7123 | 341039.0999 | 327650.2207 | 331833.3646 | 260688.5339 | 260938.5144 | 269322.0809 | 7929937.0384 | 7923674.8463 | 7905690.0689 | 2162059.5813 | 2193358.9146 | 2224973.3103
| HLQC 3 | 685404.1696 | 678832.0061 | 665415.4547 | 451892.2873 | 436358.7121 | 436108.372 | 11935406.9988 | 12107232.0486 | 11398011.4228 | 6073696.9814 | 5954194.6841 | 5720071.1667 | 318818.984 | 320329.059 | 316342.6167 | 262565.5872 | 260525.2517 | 251636.9195 | 7688301.821 | 7576975.4577 | 7181768.6564 | 2141375.5376 | 2095863.0515 | 1987578.2535
| HLQC 4 | 720503.2849 | 735312.4968 | 710119.5049 | 462003.1915 | 468037.5473 | 471986.6309 | 12965521.1054 | 12730618.6285 | 12853440.8668 | 6466926.7593 | 6260511.7785 | 6366523.4336 | 348527.2024 | 341087.0353 | 326273.8149 | 285629.9268 | 276956.2766 | 278163.2361 | 8066172.5495 | 8075623.8702 | 7953944.1533 | 2258291.3043 | 2246425.8746 | 2277433.2177
| HLQC 5 | 709212.4107 | 685671.293 | 673772.4512 | 469556.4849 | 444381.0938 | 427422.4355 | 11778674.462 | 11851671.9023 | 11058544.3667 | 6006211.2044 | 5931669.1124 | 5754242.2808 | 341750.0304 | 328787.3222 | 334057.4627 | 266163.6492 | 265541.5306 | 260278.7339 | 8048505.8 | 7669255.176 | 7455891.2136 | 2253205.1179 | 2150440.0038 | 2065942.75
| HLQC 6 | 698699.5097 | 701426.8309 | 696999.1034 | 454059.7276 | 464368.876 | 448011.9486 | 12698385.2056 | 12281249.5949 | 12510458.3112 | 6199955.6613 | 6263772.6312 | 5994907.731 | 334889.1773 | 329816.5055 | 331106.063 | 271386.6179 | 272673.1777 | 268231.4252 | 8112516.6848 | 8100125.078 | 8084449.8829 | 2279472.2634 | 2226587.4675 | 2208074.244

## Main Raw Peak Areas
Sample_Name | Cer d18:1/16:0 [M>264] | Cer d18:1/16:0-d7 [M>264] | Cer d18:1/18:0 [M>264] | Cer d18:1/18:0-d7 [M>264] | Cer d18:1/24:0 [M>264] | Cer d18:1/24:0-d7 [M>264] | Cer d18:1/24:1 [M>264] | Cer d18:1/24:1-d7 [M>264]
Total blank 5uL_r1_t1 | 2185.3102 | 834.7428 | 4594.1614 | 390.3044 | N/A | N/A | 493.6808 | N/A
Total blank 5uL_r1_t2 | 2758.1202 | 1092.0818 | 4916.0736 | N/A | N/A | N/A | N/A | N/A
Total blank 5uL_r1_t3 | 3103.5235 | 559.3388 | 4027.9966 | 704.3229 | N/A | N/A | 555.2228 | 218.6129
Matrix blank 5uL_r1_t1 | 258032.9381 | 125444.0433 | 4187598.6087 | 1236581.0339 | N/A | N/A | 139.1157 | N/A
Total blank 5uL_r1_t4 | N/A | N/A | 5598.9136 | 1311.5453 | 318.2154 | N/A | N/A | 311.8367
Total blank 5uL_r1_t5 | N/A | 372.4782 | 1839.487 | 643.6051 | N/A | N/A | 366.2274 | N/A
Cer std 7_r1_t1 | 1474.7233 | 546.6983 | 2939.7046 | 302.3993 | 300766.9033 | 297385.0222 | 4951735.9755 | 1752621.1455
Cer std 6_r1_t1 | 29234.7905 | 30743.3593 | 276421.6923 | 242999.4098 | 290389.386 | 278403.8317 | 4900021.1974 | 1734655.2555
Cer std 5_r1_t1 | 35155.553 | 38212.8857 | 340531.8072 | 321323.3847 | 290217.1382 | 266116.6556 | 4773893.0437 | 1628456.7729
Cer std 4_r1_t1 | 67849.5961 | 76449.3069 | 703488.1976 | 701536.9953 | 290420.0421 | 272093.4308 | 4938769.2626 | 1700045.9114
Cer std 3_r1_t1 | 304585.3622 | 324580.892 | 3445668.7461 | 3321000.6165 | 293291.6948 | 280076.5382 | 5123336.4682 | 1707427.2442
Cer std 2_r1_t1 | 3308744.6377 | 3480150.5656 | 40866401.7793 | 37374053.4088 | 293647.003 | 268432.441 | 5355096.1485 | 1732081.8191
Cer std 1_r1_t1 | 6238229.0128 | 7098547.848 | 87147649.6579 | 73479605.5366 | 284903.947 | 267696.2032 | 6087516.1429 | 1838414.6758
Total blank 5uL_r1_t6 | N/A | 5807.6433 | 60248.8006 | 55439.7633 | N/A | 272.9025 | N/A | 2723.8091
Total blank 5uL_r2_t1 | 5389.7746 | 3374.0851 | 19612.0193 | 9807.7538 | N/A | N/A | N/A | 569.8981
Cer std 7_r2_t1 | 4123.0863 | 1986.8197 | 10927.6442 | 5600.5284 | 310033.9811 | 293702.2774 | 4933765.6393 | 1760131.0876
Cer std 6_r2_t1 | 35134.4322 | 32969.1417 | 282661.3213 | 275230.8719 | 290849.6211 | 286311.1073 | 5121714.8789 | 1689298.4461
Cer std 5_r2_t1 | 31125.5404 | 34567.9771 | 325935.6413 | 333401.8469 | 290191.5641 | 274851.6621 | 4966361.118 | 1685110.8148
Cer std 4_r2_t1 | 64951.2646 | 70713.2163 | 730442.1832 | 694369.0239 | 302947.8215 | 287549.0945 | 5248247.2469 | 1760339.8598
Cer std 3_r2_t1 | 307149.5373 | 338604.849 | 3453512.2287 | 3143166.2807 | 293873.0905 | 284790.3703 | 5152749.0295 | 1732349.0307
Cer std 2_r2_t1 | 3204919.9472 | 3605950.5724 | 38815049.2119 | 36022197.255 | 293899.6294 | 302715.75 | 5602402.993 | 1755788.6576
Cer std 1_r2_t1 | 6903427.3567 | 7753889.3752 | 92687560.7336 | 81308303.413 | 295256.5655 | 291079.1533 | 6309437.6325 | 1920371.9479
EA.IPA_r1_t1 | N/A | 5123.8534 | 65215.4307 | 62205.3149 | N/A | 375.4227 | 6053.2197 | 2741.7388
EA.IPA_r1_t1 | N/A | 1716.5251 | 13311.626 | 9734.2151 | N/A | 331.8311 | 1901.2223 | 411.3527
Total blank 5uL_r2_t2 | 5244.1561 | 4147.8351 | 10341.2245 | 5407.2095 | N/A | 135.1942 | N/A | 174.9673
Total blank 5uL_r2_t3 | 4374.9975 | 2247.2701 | 9657.9033 | 3831.4044 | N/A | N/A | N/A | N/A
Cer std 7_r3_t1 | 5868.1562 | 2350.455 | 8725.8957 | 2784.7015 | 301920.1529 | 300942.5552 | 5069529.5755 | 1765506.5679
Cer std 6_r3_t1 | 29710.6934 | 30392.1571 | 285096.2949 | 273895.4055 | 305786.6961 | 293929.21 | 5137815.4713 | 1792114.0394
Cer std 5_r3_t1 | 34601.0528 | 41461.1998 | 340213.7801 | 337295.822 | 306953.6232 | 297187.2152 | 5147563.3648 | 1755092.0903
Cer std 4_r3_t1 | 69157.2294 | 75726.1704 | 711253.7785 | 678364.8928 | 300586.9638 | 311712.0364 | 5039751.9135 | 1739077.5504
Cer std 3_r3_t1 | 316478.0852 | 359040.6607 | 3462042.9984 | 3241448.061 | 308760.5697 | 307580.4981 | 5213699.2059 | 1689439.7476
Cer std 2_r3_t1 | 3199304.2623 | 3698598.8353 | 40018281.8684 | 36091955.383 | 310745.0326 | 304977.3578 | 5428492.1328 | 1793895.2393
Cer std 1_r3_t1 | 6950469.2018 | 7877571.2428 | 93613476.6525 | 81619964.089 | 287465.7644 | 284777.7654 | 6285396.4273 | 1875851.4437
Total blank 5uL_r2_t4 | 7247.3777 | 6626.1288 | 54796.5899 | 49934.3265 | N/A | 219.1989 | N/A | 904.7407
Total blank 5uL_r2_t5 | N/A | 1193.4904 | 13888.7312 | 8566.7948 | N/A | N/A | 849.1149 | 295.5787
Pooled QC EG037_r1_t1 | 340911.6945 | 114430.3932 | 3832109.3554 | 1500245.629 | N/A | 174.2282 | 1469.9409 | 754.1342
Pooled QC EG037_r1_t2 | 349615.4298 | 106872.5246 | 3677426.8287 | 1472794.6628 | N/A | N/A | 1117.6948 | 511.7338
Pooled QC EG037_r1_t3 | 342042.0554 | 109508.2696 | 3810900.457 | 1535379.1637 | N/A | N/A | 949.273 | 294.4695
Pooled QC EG037_r1_t4 | 357968.9764 | 101610.0384 | 3909832.3604 | 1515666.031 | N/A | N/A | 357.0334 | 454.0265
Total blank 5uL_r2_t6 | 1742.355 | 1167.1635 | 9732.7024 | 4984.302 | N/A | N/A | 812.5017 | N/A
Pooled QC EG037_r1_t5 | 349106.2733 | 110307.5405 | 3629253.3241 | 1447990.9787 | N/A | N/A | 1172.4992 | N/A
Pooled QC EG037_r1_t6 | 348830.6286 | 102549.9592 | 3648823.6691 | 1490085.8064 | N/A | 180.0405 | 316.9858 | 585.0359
Pooled QC_r1_t1 | 281263.0028 | 147836.4112 | 4571767.0168 | 1420292.046 | 262698.2389 | 312052.3512 | 5404395.7577 | 1849382.8599
Pooled QC_r1_t2 | 277441.7386 | 151322.6059 | 4910288.8341 | 1444423.5358 | 273464.7336 | 324994.6906 | 6200897.4726 | 1915739.0852
SRM1950_r1_t1 | 339465.2767 | 183544.602 | 5847002.0189 | 1814163.954 | 246218.6958 | 304258.0704 | 6234104.5882 | 1859004.1522
SRM1950_r1_t2 | 338670.8516 | 174681.6054 | 5563779.0755 | 1779961.3769 | 241231.2884 | 289040.2043 | 5727413.4296 | 1849097.089
SRM1950_r1_t3 | 334001.4656 | 170510.4819 | 5635248.5079 | 1748815.3024 | 237305.8959 | 301393.2821 | 5793272.6327 | 1854504.8465
SRM1950_r2_t1 | 321623.7495 | 165963.7913 | 5269252.0268 | 1580594.0889 | 251884.9351 | 313568.0748 | 6060114.2001 | 1898226.9544
SRM1950_r2_t2 | 310059.4679 | 162406.1026 | 5032905.9102 | 1625250.0304 | 251640.4902 | 311507.428 | 6115462.4914 | 1887995.5294
SRM1950_r2_t3 | 322842.1711 | 163324.8087 | 5183632.3652 | 1536356.1655 | 251414.0847 | 326314.8335 | 5592161.1094 | 1884616.5605
Pooled QC_r1_t3 | 287390.5388 | 154639.0245 | 5093792.4113 | 1467247.5359 | 265101.4892 | 312706.9856 | 6306796.7887 | 2036420.2909
Pooled QC_r1_t4 | 283465.3103 | 145624.326 | 4882089.3755 | 1442402.3092 | 268280.7981 | 319406.2769 | 6445083.8821 | 1946519.8302
SRM1950_r3_t1 | 327871.8764 | 167327.7229 | 5601539.165 | 1661061.5499 | 240071.4664 | 292131.3067 | 6194520.4734 | 1896854.0988
SRM1950_r3_t2 | 334613.6167 | 171428.8938 | 5542735.1175 | 1630754.8306 | 241635.3035 | 298924.7979 | 6101542.1628 | 1897909.3938
SRM1950_r3_t3 | 330227.4022 | 165901.8703 | 4990928.1501 | 1620826.3847 | 248692.9213 | 300755.2104 | 5784814.4855 | 1847231.3326
SRM1950_r4_t1 | 316416.2567 | 162520.6662 | 5398538.2072 | 1578358.8294 | 249674.7742 | 314548.5856 | 6235230.4141 | 1826725.9761
SRM1950_r4_t2 | 323442.7682 | 174333.4021 | 5257989.4382 | 1629772.2482 | 259259.1493 | 321449.9428 | 6215552.7473 | 1825879.2881
SRM1950_r4_t3 | 320914.1855 | 166805.5283 | 5257696.5754 | 1602073.8788 | 242452.6252 | 309993.0284 | 5966513.9552 | 1842970.3038
Pooled QC_r1_t5 | 270748.3546 | 148499.7574 | 5078076.6779 | 1428386.1278 | 273808.2372 | 317130.4104 | 6317995.0663 | 1932896.494
Pooled QC_r1_t6 | 293854.5485 | 145183.1412 | 4960886.7631 | 1539620.9733 | 277728.5038 | 315364.8932 | 6420536.2013 | 2049443.8393
SRM1950_r5_t1 | 360027.1388 | 179782.3289 | 6044171.0973 | 1803336.1977 | 261115.8251 | 312149.0478 | 6319873.256 | 1943618.2894
SRM1950_r5_t2 | 363033.8799 | 183388.6098 | 6377920.0395 | 1846494.4063 | 257719.0228 | 306094.9698 | 6356453.3101 | 1998483.1681
SRM1950_r5_t3 | 379418.3941 | 185064.1392 | 6087576.329 | 1830075.2023 | 252510.9831 | 306772.874 | 6451167.5294 | 1944500.0678
SRM1950_r6_t1 | 362605.8412 | 172839.3085 | 6046700.3139 | 1758361.3583 | 266598.9215 | 308775.7224 | 6703460.3759 | 2037978.1928
SRM1950_r6_t2 | 357288.8658 | 171373.0553 | 5939921.8265 | 1828775.8866 | 263265.708 | 301993.8629 | 6628178.2131 | 1970799.9271
SRM1950_r6_t3 | 338884.685 | 167033.8367 | 5315140.5523 | 1674550.3313 | 250058.4296 | 298722.5931 | 5650693.3884 | 1839103.5223
Pooled QC_r2_t1 | 265046.2923 | 142179.8366 | 4719823.9436 | 1353519.3065 | 297386.2711 | 340026.5514 | 6671049.9546 | 2093118.6164
Pooled QC_r2_t2 | 265524.339 | 136384.4889 | 4399507.7323 | 1350813.868 | 297960.2838 | 335779.1592 | 6412663.7994 | 2091726.0187
T1D_r1_t1 | 310397.6003 | 211933.0071 | 5274452.0391 | 1719681.6727 | 255254.2394 | 319524.4295 | 6075652.167 | 1917513.4264
T1D_r1_t2 | 301627.0111 | 201802.7709 | 5495870.2606 | 1673893.8015 | 263304.7625 | 309513.6721 | 6157310.0263 | 1953408.2289
T1D_r1_t3 | 311642.3914 | 197568.8198 | 5583169.2884 | 1713248.5146 | 260922.554 | 300178.259 | 6183576.8619 | 2060097.1718
T1D_r2_t1 | 318659.337 | 215444.9323 | 5839648.5307 | 1683425.3062 | 279050.3741 | 333265.5695 | 6859861.1826 | 2051697.7467
T1D_r2_t2 | 329998.4685 | 209210.2838 | 5455459.663 | 1725153.8931 | 275559.6917 | 322408.9945 | 6449536.4317 | 2034071.0626
T1D_r2_t3 | 298215.1843 | 198092.3875 | 5493325.2804 | 1607010.522 | 265740.6903 | 326178.2652 | 6289788.1404 | 1912443.0274
Pooled QC_r2_t3 | 264116.8768 | 142963.9871 | 4407505.2985 | 1390999.087 | 276026.9721 | 352527.7848 | 6109946.446 | 2013514.7722
Pooled QC_r2_t4 | 256440.9976 | 139835.8308 | 4430283.5365 | 1300024.5977 | 292451.4631 | 338030.0839 | 6109236.7831 | 1909639.0037
T1D_r3_t1 | 325113.7927 | 193586.6658 | 5631590.65 | 1729199.7784 | 271969.7355 | 321431.9925 | 6622338.649 | 2065716.9586
T1D_r3_t2 | 293692.2494 | 193512.5953 | 5187534.9182 | 1612491.339 | 264405.1875 | 329659.9068 | 6220726.6175 | 1952317.0994
T1D_r3_t3 | 311864.9631 | 193816.9394 | 5389910.4773 | 1680706.9815 | 282956.2181 | 322732.0106 | 6423228.3127 | 2068690.0853
T1D_r4_t1 | 318237.461 | 206497.4138 | 6019521.5353 | 1733841.169 | 288357.1568 | 329467.6331 | 6691257.8366 | 2139222.0074
T1D_r4_t2 | 302902.0339 | 204787.5814 | 5321257.2308 | 1624911.8582 | 267980.6264 | 319736.1345 | 6238650.2936 | 1995111.8626
T1D_r4_t3 | 318649.7562 | 200424.0085 | 5875054.5232 | 1772538.2784 | 277822.7446 | 337098.7319 | 6791047.406 | 2133950.5624
Pooled QC_r2_t5 | 266056.62 | 138421.6052 | 4787918.0114 | 1347621.761 | 294374.6535 | 339525.2902 | 6706334.0654 | 2043092.1753
Pooled QC_r2_t6 | 266723.4801 | 143755.3705 | 4484894.3148 | 1271682.5682 | 291254.7305 | 345923.7293 | 6212333.3056 | 1968300.1271
T1D_r5_t1 | 304693.9952 | 209537.9321 | 5411582.7454 | 1650600.2292 | 258273.8165 | 316966.8922 | 6232192.5317 | 1984779.4482
T1D_r5_t2 | 330049.3743 | 190760.6184 | 5274960.5056 | 1632220.8866 | 272288.4632 | 321248.0609 | 6237096.403 | 1923076.311
T1D_r5_t3 | 308697.0805 | 191530.3113 | 5296330.3994 | 1629296.4492 | 269029.4052 | 315192.9192 | 6135278.0368 | 1929939.8526
T1D_r6_t1 | 311105.4568 | 192415.9058 | 5743060.5929 | 1698313.4049 | 280232.6584 | 323054.7162 | 6720857.7059 | 2065412.8886
T1D_r6_t2 | 299034.6851 | 191450.2527 | 5227530.4256 | 1555515.2335 | 259130.9345 | 316632.1128 | 6045080.6571 | 1921868.5011
T1D_r6_t3 | 301496.0464 | 196054.4752 | 5173436.9901 | 1622668.4853 | 265871.8148 | 335126.0985 | 6118107.0171 | 1992178.1451
Total blank 5uL_r3_t1 | 4706.5019 | 2193.2707 | 10423.0994 | 3382.7053 | N/A | 527.1893 | 7733.1747 | 2924.5729
Pooled QC_r3_t1 | 281080.8002 | 153634.472 | 4954845.5577 | 1463469.0422 | 291903.5667 | 351969.6101 | 6208672.7893 | 1995043.4167
Pooled QC_r3_t2 | 284522.7056 | 159481.4373 | 5214463.0007 | 1464844.852 | 277423.5063 | 350464.2163 | 6289110.2165 | 2040113.2585
YoungAA_r1_t1 | 250816.5424 | 111828.3902 | 3779972.1442 | 1098825.3759 | 275418.1272 | 326682.6184 | 6061251.6696 | 2128027.9074
YoungAA_r1_t2 | 256622.6622 | 102787.3773 | 3625803.1444 | 1124044.4663 | 303595.8375 | 330963.7599 | 5734075.8868 | 2054174.7483
YoungAA_r1_t3 | 243032.0401 | 107700.3147 | 3472165.1893 | 1121196.7259 | 275951.6116 | 323882.7833 | 5654822.3306 | 2095036.8438
YoungAA_r2_t1 | 236700.5883 | 98076.7768 | 3457000.7947 | 1038452.0787 | 291354.4278 | 321673.6734 | 5817762.4073 | 1965208.5011
YoungAA_r2_t2 | 241456.6331 | 104665.3115 | 3584861.5727 | 1068082.7006 | 295377.7026 | 323974.978 | 6300629.2198 | 2132921.7232
YoungAA_r2_t3 | 220993.5135 | 100767.1864 | 3293128.3448 | 1042929.1325 | 283939.0429 | 328194.6117 | 5495569.0363 | 2063021.9176
Pooled QC_r3_t3 | 296392.9081 | 156054.7003 | 5192123.4239 | 1463143.5025 | 275906.9055 | 333984.6504 | 6720782.06 | 2016273.07
Pooled QC_r3_t4 | 305087.534 | 146208.9528 | 5315067.055 | 1577592.2147 | 290054.892 | 333800.5815 | 6632445.774 | 2117951.0114
YoungAA_r3_t1 | 266708.2373 | 112134.0971 | 4207346.3818 | 1299379.6707 | 287949.4163 | 327235.0118 | 6088024.6814 | 2149098.5322
YoungAA_r3_t2 | 254306.6582 | 114891.4056 | 4150017.0473 | 1243839.675 | 274950.6699 | 305462.5518 | 6033801.3907 | 2128718.757
YoungAA_r3_t3 | 266833.7474 | 102954.8974 | 3967605.7514 | 1308117.8571 | 286207.8731 | 326274.3924 | 6143487.141 | 2100943.3162
YoungAA_r4_t1 | 229655.9509 | 104593.2413 | 3068028.1092 | 1020276.7844 | 286547.1136 | 334060.4154 | 5909764.7581 | 1983761.165
YoungAA_r4_t2 | 222767.7354 | 100931.6061 | 3163521.7114 | 1017946.8419 | 304098.2796 | 332825.5061 | 5677110.7951 | 1969891.4815
YoungAA_r4_t3 | 216738.2924 | 96379.7805 | 3316599.0222 | 1098197.0899 | 267952.7362 | 324946.8359 | 5904378.7167 | 2078432.3046
Pooled QC_r3_t5 | 289728.1578 | 155375.5772 | 4980383.3219 | 1466033.8535 | 276339.1082 | 347747.2714 | 6435553.3443 | 2094262.1072
Pooled QC_r3_t6 | 286717.7381 | 153668.8777 | 5043013.1112 | 1507971.7059 | 283283.1836 | 328478.8856 | 6315906.188 | 1981128.6061
YoungAA_r5_t1 | 226784.7947 | 109941.4899 | 3449017.1437 | 1029017.7665 | 272255.8495 | 325066.2748 | 5774216.5998 | 1999751.5429
YoungAA_r5_t2 | 226993.8269 | 103872.5144 | 3258291.9591 | 1102910.8604 | 282324.2525 | 330930.729 | 5613481.1341 | 2010718.5653
YoungAA_r5_t3 | 244840.997 | 102890.8485 | 3294567.5973 | 1089255.6835 | 272904.2665 | 326509.6483 | 5679751.3844 | 1953815.3186
YoungAA_r6_t1 | 228573.0204 | 98213.1988 | 3460590.3562 | 1168272.1768 | 296181.5962 | 320669.8384 | 6056117.8789 | 2104803.1879
YoungAA_r6_t2 | 227756.561 | 105081.8966 | 3441414.6158 | 1051920.7047 | 291195.3862 | 330178.8207 | 5862190.9044 | 2074890.456
YoungAA_r6_t3 | 225071.472 | 97248.3883 | 3382845.1045 | 1025557.3199 | 273550.6752 | 322436.7853 | 5651442.1039 | 1926200.8672
Pooled QC_r4_t1 | 294924.104 | 149009.5059 | 4950673.9835 | 1507377.8558 | 300877.5557 | 311465.3577 | 6851260.951 | 2112677.6462
Pooled QC_r4_t2 | 270042.3301 | 130575.859 | 3957075.7652 | 1319452.422 | 280194.7618 | 318763.4261 | 5192759.6354 | 1964023.5745
highTAG_r1_t1 | 395198.8822 | 208732.0879 | 9843831.242 | 2515240.2899 | 254596.2382 | 300460.0897 | 7317294.8858 | 2050637.1286
highTAG_r1_t2 | 384274.4671 | 207366.4521 | 9225162.2682 | 2423180.063 | 252909.6914 | 318070.2301 | 6874897.7985 | 1912209.8357
highTAG_r1_t3 | 384013.6844 | 208416.0785 | 9849574.8609 | 2438630.0253 | 253108.7971 | 321609.3335 | 7262675.9392 | 2012913.5122
highTAG_r2_t1 | 375449.0085 | 201113.619 | 9293889.2555 | 2359209.2864 | 260729.3555 | 340969.8671 | 7530403.0656 | 1915410.5538
highTAG_r2_t2 | 378166.1843 | 203693.1798 | 9352913.2568 | 2365560.7802 | 267289.7406 | 325090.5625 | 7254001.0011 | 1969050.3526
highTAG_r2_t3 | 385022.6939 | 192643.7524 | 8657420.752 | 2293270.0094 | 258884.4007 | 315312.0431 | 6788859.9567 | 1879944.2656
Pooled QC_r4_t3 | 284140.086 | 151806.9208 | 5014617.7336 | 1495687.5585 | 298653.5569 | 337014.2814 | 6742900.6563 | 2187643.6137
Pooled QC_r4_t4 | 276613.9528 | 144341.6317 | 4704661.2143 | 1408505.2755 | 285103.6778 | 344817.6585 | 5963006.5524 | 2038313.4133
highTAG_r3_t1 | 389091.924 | 211748.2706 | 9790237.6409 | 2463749.4335 | 247537.7332 | 330366.3436 | 7649737.1631 | 2015255.7665
highTAG_r3_t2 | 395750.8945 | 204488.1629 | 9178236.7776 | 2482187.0078 | 258549.8024 | 328099.4131 | 7464495.5397 | 2028392.9776
highTAG_r3_t3 | 381805.8576 | 196953.1498 | 9151914.4201 | 2417576.5535 | 262571.0529 | 318197.5167 | 7178078.6123 | 2001506.9195
highTAG_r4_t1 | 403494.3843 | 194936.5381 | 9506120.4564 | 2334576.9887 | 259208.3482 | 318931.1563 | 7701182.8508 | 1969582.1312
highTAG_r4_t2 | 379995.7056 | 200446.6362 | 9607370.4113 | 2418327.9953 | 259355.6004 | 323898.8234 | 7373095.9938 | 2032035.8364
highTAG_r4_t3 | 380724.5269 | 202526.7315 | 9563533.1338 | 2451135.6158 | 255453.0883 | 323094.5884 | 7515969.0068 | 2046243.1182
Pooled QC_r4_t5 | 286914.6885 | 153497.1497 | 4822116.98 | 1400348.4175 | 283972.47 | 350868.4697 | 6431170.259 | 2106716.2067
Pooled QC_r4_t6 | 281306.7115 | 134851.4646 | 4594132.9765 | 1460061.7249 | 286126.1689 | 303628.715 | 6107128.9721 | 2055865.524
highTAG_r5_t1 | 377578.597 | 209258.65 | 10365440.4995 | 2698596.5796 | 263338.9192 | 331346.3853 | 7984097.552 | 2097890.0788
highTAG_r5_t2 | 394272.1414 | 210434.8443 | 9674773.5724 | 2336312.2809 | 267088.7795 | 332049.4114 | 7590240.5283 | 1993595.527
highTAG_r5_t3 | 383127.3136 | 198955.0927 | 9723430.0674 | 2361285.2788 | 252912.7252 | 313991.2416 | 7294337.1493 | 1979632.0306
highTAG_r6_t1 | 396289.8207 | 202800.1173 | 9721759.1589 | 2473981.8262 | 259884.547 | 337273.6184 | 7614844.7728 | 2115283.8306
highTAG_r6_t2 | 397529.6962 | 196053.7972 | 10052854.7131 | 2437138.7159 | 261588.9336 | 306384.2117 | 7557047.3949 | 2010107.8507
highTAG_r6_t3 | 389627.4047 | 183599.6635 | 9901033.211 | 2402725.8176 | 268080.0645 | 297769.3485 | 7585604.5375 | 1986916.4813
Pooled QC_r5_t1 | 291131.7177 | 133699.617 | 5327143.0198 | 1447507.1842 | 303554.1277 | 326492.6686 | 6854426.4494 | 2127924.2806
Pooled QC_r5_t2 | 288499.751 | 130452.1522 | 4671178.8723 | 1401530.4994 | 299737.3757 | 322945.5187 | 6367852.8471 | 2127015.6644
EA.IPA_r1_t1 | N/A | 415.114 | 6491.2253 | 2027.5959 | N/A | N/A | 5653.5141 | 2029.4052
EA.IPA_r1_t1 | N/A | 274.7348 | 6500.416 | 983.733 | N/A | 176.984 | 2533.6309 | 1012.8063
Total blank 5uL_r3_t2 | 4849.9694 | 1132.4328 | 6850.4506 | 1232.5166 | N/A | N/A | 1554.5614 | 433.1176
Total blank 5uL_r3_t3 | 4176.5054 | 1609.1086 | 8189.4098 | 935.2701 | N/A | N/A | 1443.79 | 618.539
Cer std 7_r1_t1 | 7804.5433 | 3469.2479 | 11968.5703 | 2126.105 | 332936.6717 | 316733.8291 | 5345323.0526 | 1881353.473
Cer std 6_r1_t1 | 26205.7196 | 31457.8109 | 296167.685 | 274811.0291 | 303370.0312 | 314364.7275 | 5177706.8721 | 1873410.1752
Cer std 5_r1_t1 | 33922.4891 | 37404.3761 | 354470.0975 | 341503.1884 | 298160.0936 | 291885.7413 | 5061010.6544 | 1755589.6434
Cer std 4_r1_t1 | 71688.6056 | 80862.8386 | 739315.0385 | 702596.8428 | 313374.7812 | 301438.7737 | 5244393.1063 | 1881088.7707
Cer std 3_r1_t1 | 326245.9548 | 381139.7836 | 3669419.2645 | 3549784.3371 | 329699.658 | 297790.56 | 5347297.3169 | 1860705.0622
Cer std 2_r1_t1 | 3463807.3403 | 3572862.8706 | 42009061.0787 | 37617113.2103 | 316632.1141 | 273691.7151 | 5539320.9667 | 1881752.1864
Cer std 1_r1_t1 | 6790728.3839 | 8042847.0426 | 92810163.5185 | 77486726.9584 | 306310.0837 | 292354.2687 | 6340861.1486 | 1951893.6415
EA.IPA_r1_t1 | 5832.3125 | 5384.474 | 54640.1897 | 56483.287 | N/A | N/A | 4926.4508 | 1881.7108
EA.IPA_r1_t1 | 2077.7677 | 1878.0633 | 23972.0534 | 15399.8372 | N/A | N/A | 2110.7952 | 1117.7119
Total blank 5uL_r3_t4 | 3057.2712 | 1451.5123 | 14363.8643 | 6714.9528 | N/A | N/A | 2262.9431 | 862.5686
Total blank 5uL_r3_t5 | 3091.3252 | 1642.8033 | 11564.2419 | 4454.3293 | N/A | N/A | 2259.7061 | 395.1926
Cer std 7_r2_t1 | 1816.5184 | 1080.7357 | 10981.1856 | 2849.5204 | 314137.4657 | 310372.5794 | 5300259.025 | 1801857.8925
Cer std 6_r2_t1 | 27118.2534 | 31853.0145 | 284426.8628 | 278053.753 | 311276.0217 | 294832.1636 | 5208035.5553 | 1761209.6895
Cer std 5_r2_t1 | 35145.4718 | 34592.07 | 357693.372 | 348659.6691 | 298689.7447 | 294257.6948 | 5211471.8065 | 1789122.7951
Cer std 4_r2_t1 | 73734.938 | 73714.246 | 775978.4604 | 751916.179 | 319200.1478 | 304719.6632 | 5322052.8766 | 1807401.7957
Cer std 3_r2_t1 | 321552.6742 | 352104.2343 | 3385686.008 | 3388860.638 | 330573.4712 | 322162.6442 | 5140361.0787 | 1858890.2613
Cer std 2_r2_t1 | 3348286.3812 | 3842769.9648 | 41764067.3852 | 37368702.2152 | 319212.7195 | 297346.3715 | 5840001.0122 | 1810948.7761
Cer std 1_r2_t1 | 7334724.8254 | 8123667.8698 | 97245094.5024 | 80016694.5813 | 336974.7863 | 296886.0282 | 6452544.211 | 2118333.2597
EA.IPA_r1_t1 | 5300.4467 | 5185.2156 | 52766.7319 | 50600.7972 | 449.8729 | 294.4912 | 4355.2885 | 1840.8123
EA.IPA_r1_t1 | 2423.1428 | 1782.8285 | 19472.2592 | 12521.483 | N/A | 509.6013 | 3380.3567 | 879.4628
Total blank 5uL_r3_t6 | N/A | 1932.6802 | 13308.3843 | 7072.0593 | N/A | N/A | 1973.4037 | 808.4478
Total blank 5uL_r4_t1 | N/A | 2011.3256 | 12253.91 | 6258.2718 | N/A | N/A | 1640.9585 | 589.5446
Cer std 7_r3_t1 | 1569.1637 | 801.1594 | 7966.1631 | 3756.5368 | 320713.9126 | 285704.6488 | 5052333.1269 | 1846099.952
Cer std 6_r3_t1 | 27160.5368 | 32778.4659 | 300005.534 | 288215.9671 | 317760.8165 | 290903.8125 | 5419834.2425 | 1784461.4194
Cer std 5_r3_t1 | 34658.4656 | 37522.3261 | 376909.9318 | 369691.9901 | 329051.6881 | 302010.975 | 5374445.746 | 1855844.1361
Cer std 4_r3_t1 | 70427.8327 | 78030.0328 | 753641.1691 | 723445.3359 | 319044.1619 | 317978.2672 | 5560182.2423 | 1871795.205
Cer std 3_r3_t1 | 321913.1536 | 368800.3772 | 3552488.3305 | 3442850.944 | 318867.2981 | 315463.3381 | 5469279.186 | 1891549.8148
Cer std 2_r3_t1 | 3411076.0924 | 3581256.707 | 41902400.8878 | 37503112.2938 | 325467.9845 | 293053.4519 | 5769198.3402 | 1851158.9342
Cer std 1_r3_t1 | 7359838.6733 | 8724932.0466 | 95138225.9462 | 83578926.0541 | 304755.4654 | 311387.5491 | 6225812.3861 | 1957270.0698
Total blank 5uL_r4_t2 | N/A | 6778.2075 | 75110.8838 | 58261.6013 | N/A | N/A | 5525.439 | 2243.6284
Total blank 5uL_r4_t3 | 4311.284 | 4146.4207 | 28119.5735 | 20351.4899 | N/A | N/A | 3859.2633 | 1568.4486
Pooled QC_r5_t3 | 288569.677 | 151477.929 | 5139777.377 | 1457336.8549 | 303693.9957 | 364141.9203 | 6653396.7619 | 2178842.9672
Matrix blank 5uL_r1_t2 | 273501.0254 | 144700.2413 | 4592603.3703 | 1384793.1311 | N/A | N/A | 6708.0749 | 2319.1562
Total blank 5uL_r4_t4 | 4286.165 | 2646.5953 | 13605.4853 | 5915.4759 | N/A | N/A | 2407.9867 | 1272.673
Total blank 5uL_r4_t5 | 3269.4154 | 1968.7735 | 9218.3648 | 2687.3984 | N/A | N/A | 1107.8347 | 356.1722
Pooled QC EG037_r1_t7 | 379373.519 | 118923.5206 | 4061837.4843 | 1618248.5118 | N/A | N/A | 1608.4344 | 1180.9541
Pooled QC EG037_r1_t8 | 374345.0469 | 112300.5956 | 3773296.316 | 1541293.1244 | N/A | N/A | 2175.9697 | 650.272
Pooled QC EG037_r1_t9 | 375818.829 | 109791.5576 | 3775728.0336 | 1502113.7181 | N/A | N/A | 1270.2481 | 616.4583
Pooled QC_r5_t4 | 284560.0156 | 147144.1345 | 4387338.6076 | 1439434.7054 | 306862.2365 | 346688.7606 | 5730235.6348 | 2130462.9131
Pooled QC_r5_t5 | 295579.9476 | 157323.1366 | 5018411.4169 | 1459655.4781 | 302961.4681 | 373235.3525 | 6847494.6299 | 2204813.3477
LLQC_r1_t1 | 79689.9708 | 38795.1489 | 1072050.0048 | 333151.1486 | 333853.8934 | 339373.3514 | 5708384.8908 | 1880862.2067
LLQC_r1_t2 | 86296.84 | 38852.6142 | 1030835.0345 | 345034.8072 | 328066.3145 | 330077.842 | 5723956.0425 | 1889280.1664
LLQC_r1_t3 | 90174.3255 | 31191.8877 | 1050676.6558 | 354256.9083 | 339464.5161 | 308110.2263 | 5767313.0284 | 1969743.0701
LLQC_r2_t1 | 84672.0562 | 35606.0228 | 1122124.3341 | 365984.5487 | 340061.4087 | 297588.3667 | 5702749.103 | 2057149.5147
LLQC_r2_t2 | 94394.4172 | 35402.079 | 1088066.7953 | 351300.3423 | 344911.7397 | 311282.5784 | 5815387.3308 | 2055660.1803
LLQC_r2_t3 | 85798.4644 | 40657.7544 | 1064218.3771 | 352031.4614 | 319779.1689 | 345974.0963 | 5701451.8045 | 1915272.0272
LLQC_r3_t1 | 86836.8849 | 39653.5409 | 1051925.7653 | 344814.6302 | 312227.0953 | 334964.8896 | 5621909.0573 | 1973493.4489
LLQC_r3_t2 | 82166.3419 | 38238.6297 | 1041983.7486 | 333166.2207 | 332360.9015 | 343228.2081 | 5602555.6123 | 1913564.3351
LLQC_r3_t3 | 81847.3725 | 36263.561 | 965746.7842 | 314799.4852 | 314088.5449 | 304091.4221 | 5159842.1467 | 1831356.3252
LLQC_r4_t1 | 88498.8937 | 36411.1667 | 982182.6503 | 326228.3561 | 337995.2231 | 320645.9001 | 5476617.0011 | 1906244.6862
LLQC_r4_t2 | 91053.1884 | 34771.7979 | 998294.8681 | 305818.5112 | 329804.5458 | 332164.9023 | 5542666.4586 | 1857212.3403
LLQC_r4_t3 | 86684.0355 | 40766.0584 | 981581.7925 | 313187.2576 | 332640.6378 | 332728.9249 | 5480362.0772 | 1907251.654
LLQC_r5_t1 | 78581.4617 | 34155.5335 | 918191.2626 | 300300.7232 | 360662.8199 | 341049.4438 | 5548608.8754 | 1952505.721
LLQC_r5_t2 | 77092.9003 | 32941.7125 | 947373.4415 | 301149.9318 | 335737.6245 | 336815.1912 | 5559979.9304 | 1918739.1233
LLQC_r5_t3 | 72639.0405 | 30947.235 | 885551.9466 | 284723.7132 | 338674.4256 | 315199.8705 | 5396035.0903 | 1919490.7602
LLQC_r6_t1 | 79649.1462 | 29762.4491 | 967066.3693 | 315537.2299 | 339003.2882 | 315627.421 | 5548315.9663 | 1921997.2047
LLQC_r6_t2 | 75629.4235 | 33482.9962 | 922670.5053 | 307244.343 | 339822.128 | 332747.3834 | 5397999.3774 | 1918091.2553
LLQC_r6_t3 | 79418.7121 | 29685.4547 | 961702.6726 | 308969.6316 | 328972.2593 | 323187.4725 | 5645411.0206 | 1884539.5827
LQC_r1_t1 | 164204.8634 | 77831.3301 | 2393442.5547 | 713721.1568 | 315377.7506 | 362715.8228 | 6169742.4955 | 2048142.993
LQC_r1_t2 | 165387.8079 | 83308.8701 | 2318376.682 | 741791.6734 | 322986.0343 | 342361.7161 | 5966809.3935 | 2005720.1914
LQC_r1_t3 | 162335.4402 | 71435.0174 | 2239612.9005 | 707890.7538 | 301383.6806 | 315835.6827 | 5653126.2202 | 1988814.0239
LQC_r2_t1 | 159942.5808 | 74208.5567 | 2335653.916 | 715125.0278 | 313497.9464 | 331084.2728 | 5823864.188 | 1988890.0248
LQC_r2_t2 | 165084.1806 | 74023.9016 | 2208342.2475 | 744717.1525 | 317969.5282 | 342678.2308 | 5731331.3604 | 1983006.0367
LQC_r2_t3 | 157624.5674 | 76784.3692 | 2208826.4612 | 704222.2442 | 295026.8128 | 343307.9102 | 5767784.1823 | 1948613.4852
LQC_r3_t1 | 158343.0017 | 70484.0526 | 2238829.5129 | 702883.4417 | 321169.7221 | 336637.3203 | 5885669.4769 | 1949298.7262
LQC_r3_t2 | 163859.8598 | 76547.8677 | 2215657.388 | 668834.8685 | 312175.6964 | 341651.3903 | 5986066.0406 | 1896763.9562
LQC_r3_t3 | 159042.6916 | 73062.6165 | 2275483.0892 | 714345.5641 | 321755.7251 | 326331.0648 | 5998240.0901 | 2104295.262
LQC_r4_t1 | 171339.889 | 79362.7385 | 2310643.6832 | 700191.4635 | 329749.2292 | 333932.1339 | 5861196.0989 | 1977975.8383
LQC_r4_t2 | 164343.6875 | 71744.5673 | 2053326.5537 | 683213.7173 | 306628.4656 | 335684.376 | 5381258.2078 | 1876665.3637
LQC_r4_t3 | 161978.0859 | 73211.2849 | 2071398.0481 | 684992.4875 | 301882.1682 | 327456.6167 | 5350193.5717 | 1885385.0713
LQC_r5_t1 | 171459.9645 | 80970.8058 | 2318810.8007 | 719715.7953 | 310950.8387 | 334921.2468 | 5936038.2549 | 1938242.7619
LQC_r5_t2 | 164553.2803 | 74125.3626 | 2270784.3689 | 695264.6777 | 303663.9649 | 333979.4419 | 5688016.4109 | 1944984.5687
LQC_r5_t3 | 158622.6436 | 67626.8357 | 2077031.7659 | 697689.1518 | 299046.2647 | 289705.1105 | 5206390.9908 | 1794474.9706
LQC_r6_t1 | 157414.1322 | 63007.9987 | 2106136.3241 | 657873.7375 | 290949.7413 | 317117.2899 | 5732218.4344 | 1898703.8971
LQC_r6_t2 | 156331.3739 | 74620.7687 | 2197783.1405 | 665643.8596 | 310278.0106 | 333226.1176 | 5839521.2719 | 1912712.43
LQC_r6_t3 | 154593.009 | 70655.3347 | 2128134.0954 | 673741.2014 | 317295.03 | 325555.8711 | 5815079.0934 | 1911866.7387
MQC_r1_t1 | 287734.5612 | 158218.0011 | 4964300.1588 | 1420516.6797 | 287642.6738 | 342126.6942 | 6595962.5156 | 2012135.1457
MQC_r1_t2 | 291832.4138 | 145611.8966 | 4898263.3338 | 1476310.041 | 272422.1071 | 315251.2722 | 6580564.4212 | 2031845.9507
MQC_r1_t3 | 285930.973 | 142391.4493 | 4483335.5574 | 1410529.8159 | 271262.2736 | 304398.341 | 5712877.0551 | 1873846.2285
MQC_r2_t1 | 295889.5897 | 138902.7811 | 5294745.6843 | 1489165.8803 | 282991.7101 | 305922.1054 | 6587382.3025 | 2024192.1043
MQC_r2_t2 | 301426.9902 | 142061.3492 | 5017390.5913 | 1476226.8372 | 275325.3221 | 292326.3866 | 6178148.4002 | 2037792.7162
MQC_r2_t3 | 305487.5739 | 128792.2789 | 4382525.789 | 1410001.2713 | 271957.0679 | 286972.8325 | 5271572.9978 | 1917259.5583
MQC_r3_t1 | 297967.0943 | 146702.0234 | 5056492.7052 | 1509432.1166 | 292524.1602 | 311454.7119 | 6488398.6 | 2035595.7973
MQC_r3_t2 | 311793.8105 | 138061.0481 | 5310733.4525 | 1501406.5343 | 292810.0051 | 299583.664 | 6172256.3548 | 1969353.4214
MQC_r3_t3 | 287131.5795 | 138538.6525 | 4852285.8357 | 1407493.6215 | 278138.0205 | 285383.1401 | 6225109.4049 | 1908442.4087
MQC_r4_t1 | 307707.0285 | 160578.5057 | 5356466.9234 | 1512622.905 | 275373.2526 | 337421.0919 | 6574511.1804 | 2022297.0989
MQC_r4_t2 | 313855.1261 | 150143.087 | 5229088.3278 | 1546180.8369 | 292834.8848 | 342690.8556 | 6588282.7159 | 2035922.5515
MQC_r4_t3 | 306734.598 | 141747.6349 | 5534455.3606 | 1655772.9833 | 304590.2353 | 331900.3362 | 6625361.0416 | 2162929.8811
MQC_r5_t1 | 304438.1035 | 158049.1709 | 5400506.7984 | 1578378.7391 | 289282.218 | 350039.5725 | 6649669.058 | 2074032.5923
MQC_r5_t2 | 299395.6436 | 155014.8995 | 5196651.6973 | 1558726.5575 | 280907.706 | 350682.4645 | 6248333.0361 | 2039788.1617
MQC_r5_t3 | 283817.1574 | 148853.3415 | 4420376.0612 | 1418044.5337 | 275918.6294 | 325767.5426 | 5377483.4478 | 1963788.2366
MQC_r6_t1 | 304457.985 | 151379.773 | 5287409.2562 | 1532777.765 | 287475.3145 | 335702.6582 | 6593612.1668 | 1950559.5933
MQC_r6_t2 | 292870.043 | 153815.9685 | 5275787.1963 | 1528761.9877 | 282380.1737 | 333203.835 | 6622290.0504 | 2068314.1746
MQC_r6_t3 | 306691.3062 | 150042.7881 | 5044835.9141 | 1475078.7001 | 285186.0726 | 333332.6527 | 6307678.3109 | 2048657.3752
HQC_r1_t1 | 421837.2747 | 346101.3011 | 7187660.8477 | 3266183.9815 | 288300.1577 | 341832.9303 | 6720056.916 | 2023164.2758
HQC_r1_t2 | 432297.3111 | 355873.8097 | 7109313.0328 | 3372202.6348 | 292225.4601 | 333076.8075 | 6395595.207 | 2070649.5043
HQC_r1_t3 | 440830.8221 | 354842.1843 | 7110991.6991 | 3334027.6021 | 295621.8263 | 353831.8199 | 6325593.3523 | 2059612.9802
HQC_r2_t1 | 451076.7501 | 380324.3954 | 8490770.534 | 3758180.1219 | 297334.0697 | 351469.7575 | 7326650.2688 | 2170555.5109
HQC_r2_t2 | 448888.6843 | 370202.6754 | 8188508.8307 | 3689453.9712 | 296494.776 | 366996.756 | 7222469.1673 | 2206968.5612
HQC_r2_t3 | 443645.01 | 359127.7635 | 6638814.6915 | 3488002.5974 | 291966.9967 | 314559.2062 | 5860069.1719 | 2042385.5878
HQC_r3_t1 | 445480.1012 | 359344.9792 | 7034136.9854 | 3460661.6062 | 304218.4891 | 350403.2364 | 6621483.9242 | 2035672.6619
HQC_r3_t2 | 435040.9373 | 362122.4694 | 7585923.3156 | 3450915.2993 | 279652.8356 | 353147.1398 | 6727630.6361 | 2036194.0689
HQC_r3_t3 | 452348.5819 | 370335.1943 | 7227056.2883 | 3500786.8157 | 292925.0106 | 357855.7289 | 6443622.1616 | 2110504.1358
HQC_r4_t1 | 459973.1225 | 374730.3673 | 7541396.9795 | 3467023.0224 | 285044.6449 | 334974.6394 | 6779606.2723 | 2041487.2125
HQC_r4_t2 | 439688.8741 | 350361.8647 | 7471856.0798 | 3488686.5652 | 284203.9331 | 336285.872 | 6557860.4619 | 1989439.4722
HQC_r4_t3 | 435249.4047 | 320750.4401 | 6764794.8102 | 3248714.0157 | 280910.0381 | 335068.0279 | 5830798.1324 | 2000320.275
HQC_r5_t1 | 441663.562 | 354339.2254 | 7875754.2084 | 3392279.7325 | 289901.6987 | 343682.2558 | 7030862.8635 | 2141299.8703
HQC_r5_t2 | 449069.3544 | 352549.9574 | 7702269.9873 | 3459358.9706 | 301444.4042 | 342491.0276 | 7111436.7768 | 2095233.0018
HQC_r5_t3 | 444342.3615 | 357246.0082 | 7541252.2418 | 3413183.0263 | 295298.7881 | 343815.2743 | 7029298.7203 | 2147371.0258
HQC_r6_t1 | 436174.7902 | 364239.0336 | 7747977.9999 | 3514804.4242 | 285132.8722 | 357730.1287 | 6851324.1604 | 2115183.9107
HQC_r6_t2 | 444993.0467 | 358434.8736 | 7378568.1973 | 3373673.5391 | 292015.4538 | 349877.6193 | 6411261.1711 | 2040519.1577
HQC_r6_t3 | 427594.9201 | 302866.2367 | 5723340.8843 | 3115620.709 | 275659.4339 | 292685.244 | 4982937.2749 | 1847078.8188
HLQC_r1_t1 | 579111.3369 | 531492.0808 | 9351081.1074 | 5352683.1416 | 287510.414 | 323800.1701 | 6191180.2288 | 2027907.3787
HLQC_r1_t2 | 583376.0697 | 549551.6613 | 9863222.3501 | 5292517.3447 | 284998.3398 | 328031.6116 | 6467577.0282 | 1955364.6171
HLQC_r1_t3 | 573670.0519 | 485517.8121 | 8575556.1975 | 5256001.8213 | 284282.6733 | 307904.5935 | 6072540.274 | 2049904.943
HLQC_r2_t1 | 594309.7669 | 533925.3743 | 9825310.6122 | 5269053.2577 | 295285.4446 | 315119.4017 | 6910222.2011 | 2134557.4566
HLQC_r2_t2 | 591583.2389 | 534859.6866 | 10113402.5109 | 5587958.3764 | 292930.8837 | 319429.4227 | 7012415.2043 | 2136819.5492
HLQC_r2_t3 | 571192.8187 | 579715.9327 | 9681688.1249 | 5316569.3115 | 289442.4493 | 351636.618 | 6445561.7132 | 2143806.4208
HLQC_r3_t1 | 565234.2838 | 559293.1833 | 9699832.008 | 5269558.3663 | 288358.9746 | 347195.2419 | 6601821.797 | 1984521.1192
HLQC_r3_t2 | 579047.0642 | 525071.2726 | 9104162.9435 | 5286806.4705 | 289454.347 | 340731.7441 | 6272928.4835 | 1991082.3376
HLQC_r3_t3 | 566351.6041 | 546206.2212 | 9077130.5071 | 5244133.7185 | 285189.2822 | 337304.8132 | 6075267.1329 | 1977607.1164
HLQC_r4_t1 | 606291.1447 | 574822.5012 | 10097979.8268 | 5529580.722 | 296359.8806 | 347356.8464 | 6811755.8791 | 2101139.4311
HLQC_r4_t2 | 616021.7674 | 572527.3025 | 10333976.7178 | 5613988.2271 | 292568.0315 | 361101.6518 | 7141145.4038 | 2156356.3502
HLQC_r4_t3 | 601779.3458 | 595049.408 | 10507664.7119 | 5535502.744 | 296772.5403 | 354880.319 | 6854582.9797 | 2159211.0158
HLQC_r5_t1 | 586529.9586 | 517539.3212 | 9713253.2329 | 5282442.456 | 309620.7333 | 321523.3873 | 6962414.8719 | 2089775.0733
HLQC_r5_t2 | 583759.9807 | 570894.227 | 9782036.8587 | 5518429.874 | 294734.6807 | 356174.1974 | 6917080.6429 | 2117558.0724
HLQC_r5_t3 | 581919.5849 | 571485.6687 | 9620617.5163 | 5443896.6692 | 284267.8032 | 368430.8546 | 6806071.9165 | 2135455.4736
HLQC_r6_t1 | 591229.6763 | 494703.7311 | 10294860.7506 | 5368788.0704 | 294949.9427 | 306901.6436 | 7024332.4821 | 2077181.0375
HLQC_r6_t2 | 581671.4569 | 553256.1179 | 10229610.6122 | 5371160.7884 | 298256.2545 | 358237.7556 | 7014515.968 | 2076824.3806
HLQC_r6_t3 | 572952.0676 | 552246.7668 | 9565911.7469 | 5441798.3095 | 292228.0385 | 335274.1628 | 6635389.3175 | 2121174.5454
Pooled QC_r5_t6 | 296923.3195 | 151829.4055 | 5228941.2953 | 1489888.8819 | 314394.9748 | 376368.3119 | 6971358.3387 | 2261941.9362
Matrix blank 5uL_r1_t3 | 279831.3985 | 133361.4067 | 4680302.4093 | 1391735.5302 | N/A | N/A | 6808.065 | 2756.1259
Total blank 5uL_r4_t6 | 4389.9094 | 2412.5238 | 11343.3854 | 4233.014 | N/A | N/A | 3084.0571 | 870.5452

## Run 2 Raw Peak Areas (dQCs)
Sample_Name | Cer d18:1/16:0 [M>264] | Cer d18:1/16:0-d7 [M>264] | Cer d18:1/18:0 [M>264] | Cer d18:1/18:0-d7 [M>264] | Cer d18:1/24:0 [M>264] | Cer d18:1/24:0-d7 [M>264] | Cer d18:1/24:1 [M>264] | Cer d18:1/24:1-d7 [M>264]
Pooled QC_r1_t1 | 357597.5012 | 309586.8152 | 128164.3648 | 267719.4303 | 6584576.8266 | 7516704.4172 | 1755716.9298 | 2165822.533
Pooled QC_r1_t2 | 354546.7785 | 317894.414 | 124988.3988 | 267814.4418 | 6415676.8694 | 7592596.0461 | 1744675.4232 | 2133470.3577
Pooled QC_r1_t3 | 346990.1141 | 312310.3734 | 120658.3606 | 257219.3045 | 5803259.7354 | 6780242.1694 | 1570526.1998 | 1995624.0272
LLQC_r1_t1 | 78234.1246 | 273483.4624 | 27460.2743 | 226004.9266 | 1220777.8614 | 6044300.1552 | 367258.1097 | 1931640.0908
LLQC_r1_t2 | 73303.1245 | 289990.2734 | 23301.0379 | 230660.0769 | 1185731.9516 | 5878094.5345 | 364723.563 | 1976932.8398
LLQC_r1_t3 | 75444.4749 | 284520.8528 | 27665.3008 | 235622.4606 | 1201879.1553 | 6051803.3274 | 360522.0803 | 1912400.0546
LLQC_r2_t1 | 72895.3238 | 277853.0066 | 25346.5673 | 229236.9215 | 1198431.6914 | 5929123.9395 | 367733.575 | 1880214.9997
LLQC_r2_t2 | 76121.9462 | 285395.6283 | 27784.7047 | 227375.5661 | 1217823.5076 | 5997849.746 | 359541.1305 | 1851260.0019
LLQC_r2_t3 | 82289.8885 | 277659.0909 | 26169.5871 | 228281.5746 | 1206326.8142 | 5714819.7353 | 357837.0102 | 1887389.5762
LLQC_r3_t1 | 74785.98 | 265056.7273 | 24247.8868 | 229232.5217 | 1186067.661 | 5877135.2316 | 363858.0676 | 1913505.7481
LLQC_r3_t2 | 72362.9257 | 279571.2915 | 24499.3589 | 233021.5226 | 1182112.2175 | 6066153.0297 | 350906.2265 | 1920346.8496
LLQC_r3_t3 | 72777.5527 | 274391.4531 | 25741.6082 | 217221.5451 | 1208273.8448 | 5975577.0654 | 352132.5578 | 1922498.1059
LLQC_r4_t1 | 76755.0655 | 272146.7917 | 25652.4915 | 221560.8846 | 1151989.3127 | 5954026.8584 | 342700.3738 | 1889710.3773
LLQC_r4_t2 | 67512.4496 | 271830.2847 | 27173.2099 | 228399.6677 | 1126454.333 | 6031490.421 | 344640.4792 | 1878750.8138
LLQC_r4_t3 | 75425.4349 | 273110.1462 | 24207.6686 | 229227.1082 | 1144088.33 | 5845368.4493 | 345864.4524 | 1863768.7187
LLQC_r5_t1 | 67739.0582 | 266462.5798 | 23933.3813 | 228071.9914 | 1078884.5054 | 6036313.3324 | 320129.7745 | 1928579.9294
LLQC_r5_t2 | 66572.4391 | 275670.9326 | 23740.8106 | 224337.8953 | 1046463.4515 | 5808900.2556 | 305759.7913 | 1870685.7261
LLQC_r5_t3 | 69657.3016 | 264308.8693 | 23642.6892 | 219354.5913 | 1043178.5373 | 5868128.6392 | 317074.1917 | 1911726.2093
LLQC_r6_t1 | 69035.3116 | 272116.5239 | 24768.2713 | 233322.9255 | 1114784.3791 | 5975774.9449 | 334724.3396 | 1918451.4238
LLQC_r6_t2 | 68617.1696 | 280597.172 | 24554.6935 | 228820.5957 | 1075977.8076 | 5900548.713 | 321641.3104 | 1879453.231
LLQC_r6_t3 | 74517.5135 | 276072.9229 | 27158.9864 | 236859.748 | 1115469.8401 | 5993792.1935 | 328291.5155 | 1897024.5917
LQC_r1_t1 | 173205.2575 | 292079.5422 | 64737.0771 | 245468.7978 | 2782118.2373 | 6669014.9795 | 799827.0214 | 2010465.1884
LQC_r1_t2 | 174343.0596 | 311124.9627 | 62448.6803 | 252359.2652 | 2755659.7918 | 6490008.3622 | 764814.6632 | 1973772.6443
LQC_r1_t3 | 174857.9611 | 299469.1086 | 60992.1701 | 240224.364 | 2674842.1449 | 6411342.661 | 766174.8416 | 1927292.5134
LQC_r2_t1 | 167021.7946 | 297008.7101 | 60810.7559 | 241580.4727 | 2679254.805 | 6428344.2422 | 770120.4804 | 1968653.0939
LQC_r2_t2 | 168442.8296 | 295457.2468 | 55925.4794 | 239007.34 | 2608159.5367 | 6198189.417 | 753344.3745 | 1929159.7321
LQC_r2_t3 | 166519.7115 | 293092.2072 | 55022.5863 | 244284.1376 | 2646291.3133 | 6292280.1041 | 766227.4237 | 1938779.7733
LQC_r3_t1 | 165178.6678 | 298386.9763 | 53454.8282 | 248975.6153 | 2610240.2705 | 6362766.9711 | 734834.9727 | 1975954.0745
LQC_r3_t2 | 154798.9167 | 288244.8191 | 51385.0016 | 242438.1339 | 2580556.6392 | 6342127.1373 | 727956.2529 | 1915166.9986
LQC_r3_t3 | 160074.187 | 291287.9622 | 53652.3762 | 247304.8054 | 2515721.1512 | 6176804.5804 | 736013.2465 | 1940851.859
LQC_r4_t1 | 170594.3287 | 300095.8565 | 56664.7618 | 248372.265 | 2702746.7752 | 6422152.7282 | 756908.155 | 1973061.6269
LQC_r4_t2 | 168621.1184 | 287980.0347 | 55440.5835 | 251280.9284 | 2698065.3857 | 6437437.9105 | 769756.3039 | 1988347.4277
LQC_r4_t3 | 169803.7351 | 295232.9775 | 60360.2353 | 250732.6961 | 2699150.933 | 6328566.1005 | 776261.0773 | 1957736.6626
LQC_r5_t1 | 171222.4649 | 294478.5662 | 58215.6941 | 245498.9973 | 2772610.4758 | 6364959.0198 | 805822.1737 | 1959274.3462
LQC_r5_t2 | 165867.6043 | 290379.3896 | 58745.2679 | 239132.2079 | 2677256.5155 | 6075908.7783 | 759322.0719 | 1896477.4259
LQC_r5_t3 | 172871.2463 | 294092.3101 | 58603.9043 | 247721.5054 | 2609855.7527 | 5949742.844 | 754852.2141 | 1846620.578
LQC_r6_t1 | 166407.1178 | 294080.8084 | 58246.1025 | 245011.1942 | 2512161.4155 | 6214863.2221 | 748364.3914 | 1923638.6627
LQC_r6_t2 | 161883.3482 | 302283.0001 | 54331.9132 | 236362.6612 | 2507653.874 | 6127640.4916 | 741365.3916 | 1944366.5161
LQC_r6_t3 | 162267.0912 | 291481.9796 | 54094.0233 | 252093.5258 | 2520675.0413 | 6332252.5832 | 735079.5467 | 1902273.0927
MQC_r1_t1 | 341092.1229 | 332064.9316 | 116319.9751 | 271960.38 | 6243467.035 | 7241337.019 | 1665114.4299 | 2081547.4001
MQC_r1_t2 | 354598.1001 | 318729.649 | 123724.4929 | 261545.6998 | 6253192.0168 | 7310393.5195 | 1613008.9926 | 2101650.9531
MQC_r1_t3 | 351749.1324 | 320604.1079 | 119268.7075 | 269131.8235 | 6003948.0826 | 7182780.2276 | 1583894.2745 | 2046734.7562
MQC_r2_t1 | 368206.9927 | 328958.1623 | 130265.4231 | 281636.7087 | 6669784.2597 | 7770983.9265 | 1780153.7312 | 2216006.716
MQC_r2_t2 | 369264.5025 | 319733.7322 | 126776.8476 | 268039.3615 | 6850568.7643 | 7788553.5469 | 1753695.4839 | 2120401.1955
MQC_r2_t3 | 359059.1146 | 315855.5644 | 127271.4473 | 264846.1363 | 6628524.1857 | 7565977.4929 | 1724571.26 | 2178276.434
MQC_r3_t1 | 363933.5074 | 310613.4447 | 133768.2699 | 258212.3483 | 6462779.3646 | 7593171.4988 | 1739502.7294 | 2124366.0503
MQC_r3_t2 | 354772.4389 | 324470.2332 | 134389.4414 | 272488.7447 | 6486059.7981 | 7598719.0886 | 1720241.344 | 2120153.2642
MQC_r3_t3 | 360164.5663 | 335860.5312 | 123330.7965 | 267260.1772 | 6378899.7176 | 7330873.2346 | 1696589.7962 | 2052433.059
MQC_r4_t1 | 376295.3907 | 327922.9864 | 131202.3558 | 273177.7099 | 6711295.0455 | 7591540.1327 | 1765891.0073 | 2174318.5559
MQC_r4_t2 | 364344.518 | 327315.6431 | 128897.3244 | 261715.1454 | 6656954.6045 | 7392832.3386 | 1665687.6138 | 2100413.4295
MQC_r4_t3 | 346420.5646 | 304599.2595 | 123698.9785 | 253683.7279 | 5716891.3374 | 6682775.0943 | 1587466.1683 | 1949786.7656
MQC_r5_t1 | 367289.0521 | 324130.6939 | 124759.5288 | 253485.3747 | 6505165.6678 | 7554411.5731 | 1698893.7059 | 2161053.1964
MQC_r5_t2 | 367065.9208 | 323204.707 | 126834.3826 | 267579.6395 | 6465216.0857 | 7557568.1263 | 1699037.3172 | 2150765.7015
MQC_r5_t3 | 358434.7855 | 319021.5749 | 122003.4301 | 249968.6367 | 6009201.4125 | 7026145.6443 | 1613254.9044 | 2011025.7538
MQC_r6_t1 | 373648.8656 | 336653.6712 | 133364.9697 | 263766.2455 | 6355319.7966 | 7460001.4181 | 1770045.3872 | 2160604.6737
MQC_r6_t2 | 360000.8898 | 313503.0964 | 123410.3209 | 265269.1428 | 6333608.0624 | 7239501.0776 | 1681124.4691 | 2069246.6721
MQC_r6_t3 | 365968.243 | 311552.2685 | 121188.512 | 251654.1848 | 6214426.7953 | 7205741.2345 | 1674899.6483 | 2076744.5421
HQC_r1_t1 | 525233.8486 | 323512.4639 | 290361.3445 | 267468.6803 | 8821755.3745 | 7473558.4859 | 3683882.247 | 2172248.5167
HQC_r1_t2 | 508072.6017 | 336514.2532 | 283578.5074 | 268206.3028 | 8832226.1498 | 7605783.9423 | 3775224.3409 | 2190058.4726
HQC_r1_t3 | 504784.59 | 324126.7355 | 285459.4003 | 253774.0681 | 8328377.6064 | 7473685.3977 | 3768108.1395 | 2134904.2898
HQC_r2_t1 | 562979.0195 | 346275.0817 | 313838.4324 | 283429.3963 | 10833054.4396 | 8540794.2256 | 4268415.9848 | 2358509.4645
HQC_r2_t2 | 561808.9479 | 348584.9952 | 323316.8253 | 286056.1191 | 10402953.2659 | 8390728.1315 | 4271482.0405 | 2340972.2294
HQC_r2_t3 | 576787.6797 | 336266.5093 | 318540.6212 | 272264.2222 | 10440972.6305 | 8487952.6616 | 4196883.4249 | 2305373.3254
HQC_r3_t1 | 533767.6253 | 327967.8958 | 288682.0441 | 268013.9264 | 9547452.6017 | 7947084.2803 | 3909784.9175 | 2244427.3968
HQC_r3_t2 | 519572.043 | 317922.1327 | 274621.9127 | 258739.8394 | 8968311.8009 | 7213472.7353 | 3623081.7397 | 2058452.61
HQC_r3_t3 | 517427.3562 | 324054.8595 | 280630.227 | 263840.9553 | 9025589.1382 | 7339075.7803 | 3806952.0053 | 2095264.6778
HQC_r4_t1 | 551418.2746 | 323445.6945 | 299737.5868 | 262303.0805 | 9709260.159 | 7880154.5202 | 3947170.8336 | 2230530.4423
HQC_r4_t2 | 519211.8626 | 323969.8686 | 285272.5746 | 255107.2325 | 9224030.5114 | 7567342.485 | 3784864.5427 | 2106030.1731
HQC_r4_t3 | 534682.3556 | 323550.1324 | 290803.3923 | 267767.8017 | 9288965.8741 | 7422618.1527 | 3814463.3181 | 2125467.295
HQC_r5_t1 | 550794.4691 | 332447.7966 | 300223.2353 | 270094.3275 | 9686823.0037 | 8067027.0965 | 3950209.9224 | 2249374.735
HQC_r5_t2 | 542796.3477 | 337375.3184 | 295332.5062 | 268336.6861 | 9660339.0503 | 7917781.2747 | 4000418.991 | 2253080.642
HQC_r5_t3 | 533834.2123 | 333389.8984 | 285126.493 | 267923.0448 | 9779150.9947 | 7884048.96 | 3903916.7321 | 2200964.2041
HQC_r6_t1 | 537186.0466 | 321131.1736 | 295415.8807 | 263526.2441 | 9603751.225 | 7790303.9786 | 3950004.4646 | 2150385.6835
HQC_r6_t2 | 532091.8522 | 314786.3114 | 288488.2766 | 255322.5778 | 9256493.5713 | 7507708.0541 | 3802759.7024 | 2092363.2156
HQC_r6_t3 | 528079.4035 | 313677.2099 | 282860.3002 | 256819.9384 | 9110779.3763 | 7289031.3697 | 3771767.1707 | 2065005.7127
HLQC_r1_t1 | 694504.0234 | 318925.2246 | 462328.9142 | 271572.3335 | 12030145.3369 | 7601796.6584 | 6067357.1822 | 2115130.9433
HLQC_r1_t2 | 683277.8127 | 320935.9141 | 462031.1534 | 264419.6272 | 11982722.7637 | 7703858.8608 | 6063512.0659 | 2149183.887
HLQC_r1_t3 | 678033.1463 | 320007.7853 | 433499.1349 | 250918.2243 | 11502069.6303 | 7368183.546 | 5925912.6649 | 2045118.7421
HLQC_r2_t1 | 703072.5124 | 341039.0999 | 458300.7985 | 260688.5339 | 12476930.7951 | 7929937.0384 | 6103173.2452 | 2162059.5813
HLQC_r2_t2 | 696408.2955 | 327650.2207 | 464302.7417 | 260938.5144 | 12276134.5869 | 7923674.8463 | 6253316.1839 | 2193358.9146
HLQC_r2_t3 | 713990.7741 | 331833.3646 | 459978.9975 | 269322.0809 | 12362443.9918 | 7905690.0689 | 6262494.7123 | 2224973.3103
HLQC_r3_t1 | 685404.1696 | 318818.984 | 451892.2873 | 262565.5872 | 11935406.9988 | 7688301.821 | 6073696.9814 | 2141375.5376
HLQC_r3_t2 | 678832.0061 | 320329.059 | 436358.7121 | 260525.2517 | 12107232.0486 | 7576975.4577 | 5954194.6841 | 2095863.0515
HLQC_r3_t3 | 665415.4547 | 316342.6167 | 436108.372 | 251636.9195 | 11398011.4228 | 7181768.6564 | 5720071.1667 | 1987578.2535
HLQC_r4_t1 | 720503.2849 | 348527.2024 | 462003.1915 | 285629.9268 | 12965521.1054 | 8066172.5495 | 6466926.7593 | 2258291.3043
HLQC_r4_t2 | 735312.4968 | 341087.0353 | 468037.5473 | 276956.2766 | 12730618.6285 | 8075623.8702 | 6260511.7785 | 2246425.8746
HLQC_r4_t3 | 710119.5049 | 326273.8149 | 471986.6309 | 278163.2361 | 12853440.8668 | 7953944.1533 | 6366523.4336 | 2277433.2177
HLQC_r5_t1 | 709212.4107 | 341750.0304 | 469556.4849 | 266163.6492 | 11778674.462 | 8048505.8 | 6006211.2044 | 2253205.1179
HLQC_r5_t2 | 685671.293 | 328787.3222 | 444381.0938 | 265541.5306 | 11851671.9023 | 7669255.176 | 5931669.1124 | 2150440.0038
HLQC_r5_t3 | 673772.4512 | 334057.4627 | 427422.4355 | 260278.7339 | 11058544.3667 | 7455891.2136 | 5754242.2808 | 2065942.75
HLQC_r6_t1 | 698699.5097 | 334889.1773 | 454059.7276 | 271386.6179 | 12698385.2056 | 8112516.6848 | 6199955.6613 | 2279472.2634
HLQC_r6_t2 | 701426.8309 | 329816.5055 | 464368.876 | 272673.1777 | 12281249.5949 | 8100125.078 | 6263772.6312 | 2226587.4675
HLQC_r6_t3 | 696999.1034 | 331106.063 | 448011.9486 | 268231.4252 | 12510458.3112 | 8084449.8829 | 5994907.731 | 2208074.244
Pooled QC_r1_t4 | 365472.6913 | 325185.4813 | 123955.3415 | 266222.1778 | 6478730.6961 | 7494497.4019 | 1689638.0775 | 2124291.2816
Pooled QC_r1_t5 | 366308.8241 | 323259.8616 | 128714.118 | 258957.8228 | 6584370.7597 | 7584401.7869 | 1710710.1932 | 2132975.4746
Pooled QC_r1_t6 | 355460.6709 | 308915.3545 | 128499.8271 | 262892.4074 | 6283087.9109 | 7427724.3682 | 1671464.4522 | 2109748.2772