## Participant Information
Participant Information
Participant Information |  | Group Name
|  | Lab code | assigned by coordinators
| Institution | Institution name
Street
City
Postal code
Country
| Contact person | Name
Address
Email
Phone

## Summary
Summary
Summary | This protocol describes the extraction and LC-MRM procedure for absolute quantification of Cer d18:1/16:0, Cer d18:1/18:0, Cer d18:1/24:0 and Cer d18:1/24:1 in human plasma according to Kauhanen et al, Anal Bioanal Chem (2016) 408:3475–3483.
A calibration curve is built by injecting several dilutions of the non-labelled standards mixed with a fixed concentration of labelled standards. Each calibration point is prepared and analysed in triplicates (n=3).
To establish the endogenous concentration of the four selected ceramides in the study samples (NIST SRM1950, NIST high TAG, NIST T1D, and NIST young A-A) each sample is extracted and analyzed in six replicates (n =6).
Prepare the pooled QC by mixing 100 mL each of NIST SRM1950, NIST high TAG, NIST T1D, and NIST young AA (total 400 mL).
We suggest proceeding first to measure the ceramides concentration in the four study samples and then continue with the intra-assay validation to avoid wasting material.

## Extraction - Overview
Extraction Overview
Extraction Overview | Table 1 - Extraction - Materials
Chemicals, solvents and materials
Ethyl acetate (EtAc)
2-propanol (IPA)
5% Bovine Serum Albumin (BSA)
96-well plates/tubes
Sealing foil
2 mL DeepWell plates/tubes/LC vials
Table 2 - Extraction - Sample Types
| Sample type | Components
| Pooled Quality Control samples (QC) | 10 µL pooled QC |  |  | 20 µL Labelled IS |  | 570 µL EtAc:IPA 2:8 (v/v)
| Matrix Blank (MB) | 10 µL Matrix (pooled QC) only |  |  | - |  | 590 µL EtAc:IPA 2:8
| Study Samples (S) | 10 µL Human plasma (NIST SRM1950, high TAG, T1D, young AA) |  |  | 20 µL Labelled IS |  | 570 µL EtAc:IPA 2:8
| Total Blank (TB) | 10 µL 5% BSA |  |  | - |  | 590 µL EtAc:IPA 2:8
| Calibration line standards (STD) | 10 µL 5% BSA + 20 µL non labelled STD mixture (Table 5) |  |  | 20 µL Labelled IS |  | 550 µL EtAc:IPA 2:8
Table 3 - Extraction - Labelled Internal Standard Mixture Solution (provided)
| Labelled IS mixture in EtAc:IPA 2:8 (v/v) | Concentration |  |  |  |  | Unit | Vendor
| D7-Cer d18:1/16:0 | 0.125 |  |  |  |  | pmol/µL [µM] | Avanti Polar Lipids
| D7-Cer d18:1/18:0 | 0.050
| D7-Cer d18:1/24:0 | 1500
| D7-Cer d18:1/24:1 | 0.500
Table 4 - Extraction - Non-Labelled Standards - for STD1 Mixture Solution (provided)
| Non-labelled standards in EtAc:IPA 2:8 (v/v) | Concentration |  |  |  |  | Unit | Vendor
| Cer d18:1/16:0 | 2.0 |  |  |  |  | pmol/µL [µM] | Avanti Polar Lipids
| Cer d18:1/18:0 | 2.0
| Cer d18:1/24:0 | 20.0
| Cer d18:1/24:1 | 20.0
Table 5 - Serial dilutions of Non-Labelled Standards from STD1 Mixture Solution for calibration line
| Add 20 µL of | STD1 Conc. | STD2 Conc. | STD3 Conc. | STD4 Conc. | STD5 Conc. | STD6 Conc. | Unit
| Cer d18:1/16:0 | 2.0 | 1.0 | 0.10 | 0.02 | 0.01 | 0.008 | pmol/µL [µM]
| Cer d18:1/18:0 | 2.0 | 1.0 | 0.10 | 0.02 | 0.01 | 0.008
| Cer d18:1/24:0 | 20.0 | 10.0 | 1.0 | 0.2 | 0.1 | 0.08
| Cer d18:1/24:1 | 20.0 | 10.0 | 1.0 | 0.2 | 0.1 | 0.08

## Extraction - Procedure
Extraction Procedure
Extraction Procedure | Extraction Procedure
| Step | Description | Reference Table
| 1 | Prepare different dilutions of STD1 by diluting the provided mixture (Extraction - Overview, Table 4) with EtAc:IPA 2:8 as suggested in table | Extraction - Overview'!B34
| 2 | Thaw plasma on ice (+4°C) and bring to room temperature prior to extraction. Mix by pipetting before use.
| 3 | Prepare samples in plates/tubes manually or with a robot, according to specifications in column C of table | Extraction - Overview'!C14
| 4 | Add 20 µL of labelled internal standard mixture to samples according to column F of table | Extraction - Overview'!F14
| 5 | For the calibration line, add 20 µL of different dilutions of the non-labelled mixture (STD1-6).
| 6 | Add 590 µL of ethyl acetate:isopropanol (2:8, vol/vol) to Total Blank (TB) and Matrix Blank (MB); 550 µL to calibration line (STD); 570 µL to study samples and QCs according to column H of table | Extraction - Overview'!H14
| 7 | Mix samples 10 min by pipetting or vortexing
| 8 | Centrifuge samples for 10 min at 3000 x gat room temperature
| 9 | Transfer 50 µL of the clear supernatant to MS-vials or MS plates, close/seal, store samples at -20 °C prior to LC-MRM analysis.

## LC-MRM - Overview
LC-MRM Overview
LC-MRM Overview | Table 6 Chemicals and solvents
Ammonium formate
Formic acid
Ultra-pure water
Acetonitrile
Table 7 LC-MRM LC-MS Details
| Property | Specifications
| Mass spectrometry | QQQ
| LC | UPLC
| Injection volume (indicative) | 5 µL
| Weak wash | Water
| Strong wash | Acetonitrile
| Pre-column | Waters Acquity BEH C18, 1.7 µm VanGuard Pre-Column
| Column | Waters Acquity BEH C18, 2.1 × 50 mm id. 1.7 µm
| Column temperature | 60 °C
| Mobile phase A | 10 mM ammonium acetate in water with 0.1% formic acid
| Mobile phase B | 10 mM ammonium acetate in acetonitrile:2-propanol (4:3, vol/vol) with 0.1% formic acid
Table 8 LC-MRM LC gradient
| Time (minutes) | A% | B% |  | Flow (µL/min)
| 0.00 | 15 | 85 |  | 500
| 0.50 | 15 | 85 |  | 500
| 1.50 | 0 | 100 |  | 500
| 4.00 | 0 | 100 |  | 500
| 4.10 | 15 | 85 |  | 500
| 5.00 | 15 | 85 |  | 500
Table 9 LC-MRM MS conditions (EXAMPLE: this only applies to Sciex QTRAP, please adapt)
| Polarity mode | Positive
| Ion spray | 5000 V
| Curtain gas | 25 psi
| Source temperature | 300 °C
| Gas 1 | 50 psi
| Gas 2 | 35 psi
| DP | 30 V
| EP | 10 V
| Collision cell exit potential | 20 V
| Collision energy (CE) | 40 eV
Table 10 LC-MRM Mass transitions used for absolute quantification. On different systems, RT and CE might be different!
| Analyte | Mass transitions | IS | Mass transitions | RT (min) | CE
| Cer d18:1/16:0 | 538.5 – 264.3 | D7- Cer d18:1/16:0 | 545.6 - 271.3 | 1.65 | 40
C34H68NO3 – C18H34N
| Cer d18:1/18:0 | 566.6 – 264.3 | D7- Cer d18:1/18:0 | 573.6 – 271.3 | 1.81 | 40
C36H72NO3 – C18H34N
| Cer d18:1/24:0 | 650.6 – 264.3 | D7- Cer d18:1/24:0 | 657.7 – 271.3 | 2.15 | 40
C42H84NO3 – C18H34N
| Cer d18:1/24:1 | 648.6 – 264.3 | D7- Cer d18:1/24:1 | 655.7 – 271.3 | 2.03 | 40
C42H82NO3 – C18H34N
Table 11 Data processing
| Processing software | Analyst 1.6.2- Quantitation Wizard

## LC-MRM - Sample Sequence
Sample Sequence (Example)
Sample Sequence (Example) | Sample Sequence 1
| Number | Sample Name | Remarks
| 1 | NoInj
| 2 | Total Blank 1 1 | Extracted Total Blank 1-3 times measured
| 3 | Total Blank 1 2
| 4 | Total Blank 1 3
| 5 | Matrix blank 1 1 | Extracted Matrix Blank 1-3 times measured
| 6 | Matrix blank 1 2
| 7 | Matrix blank 1 3
| 8 | Cal6 1 | Extracted calibration 1
| 9 | Cal6 2 | Extracted calibration 2
| 10 | Cal6 3 | Extracted calibration 3
| 11 | Cal5 1 | Extracted calibration 1
| 12 | Cal5 2 | Extracted calibration 2
| 13 | Cal5 3 | Extracted calibration 3
| 14 | Cal4 1 | Extracted calibration 1
| 15 | Cal4 2 | Extracted calibration 2
| 16 | Cal4 3 | Extracted calibration 3
| 17 | Cal3 1 | Extracted calibration 1
| 18 | Cal3 2 | Extracted calibration 2
| 19 | Cal3 3 | Extracted calibration 3
| 20 | Cal2 1 | Extracted calibration 1
| 21 | Cal2 2 | Extracted calibration 2
| 22 | Cal2 3 | Extracted calibration 3
| 23 | Cal1 1 | Extracted calibration 1
| 24 | Cal1 2 | Extracted calibration 2
| 25 | Cal1 3 | Extracted calibration 3
| 26 | Pooled QC 1
| 27 | Pooled QC 2
| 28 | SRM 1950-1 1 | Extract 1 - 3 times measured
| 29 | SRM 1950-1 2
| 30 | SRM 1950-1 3
| 31 | SRM 1950-2 1 | Extract 2 - 3 times measured
| 32 | SRM 1950-2 2
| 33 | SRM 1950-2 3
| 34 | SRM 1950-3 1 | Extract 3 - 3 times measured
| 35 | SRM 1950-3 2
| 36 | SRM 1950-3 3
| 37 | SRM 1950-4 1 | Extract 4 - 3 times measured
| 38 | SRM 1950-4 2
| 39 | SRM 1950-4 3
| 40 | SRM 1950-5 1 | Extract 5 - 3 times measured
| 41 | SRM 1950-5 2
| 42 | SRM 1950-5 3
| 43 | SRM 1950-6 1 | Extract 6 - 3 times measured
| 44 | SRM 1950-6 2
| 45 | SRM 1950-6 3
| 46 | Pooled QC 3
| 47 | Pooled QC 4
| 48 | T1D-1 1 | Extract 1 - 3 times measured
| 49 | T1D-1 2
| 50 | T1D-1 3
| 51 | T1D-2 1 | Extract 2 - 3 times measured
| 52 | T1D-2 2
| 53 | T1D-2 3
| 54 | T1D-3 1 | Extract 3 - 3 times measured
| 55 | T1D-3 2
| 56 | T1D-3 3
| 57 | T1D-4 1 | Extract 4 - 3 times measured
| 58 | T1D-4 2
| 59 | T1D-4 3
| 60 | T1D-5 1 | Extract 5 - 3 times measured
| 61 | T1D-5 2
| 62 | T1D-5 3
| 63 | T1D-6 1 | Extract 6 - 3 times measured
| 64 | T1D-6 2
| 65 | T1D-6 3
| 66 | Total Blank 2 1 | Extracted Total Blank 2-3 times measured
| 67 | Total Blank 2 2
| 68 | Total Blank 2 3
| 69 | Pooled QC 5
| 70 | Pooled QC 6
| 71 | Young AA-1 1 | Extract 1 - 3 times measured
| 72 | Young AA-1 2
| 73 | Young AA-1 3
| 74 | Young AA-2 1 | Extract 2 - 3 times measured
| 75 | Young AA-2 2
| 76 | Young AA-2 3
| 77 | Young AA-3 1 | Extract 3 - 3 times measured
| 78 | Young AA-3 2
| 79 | Young AA-3 3
| 80 | Young AA-4 1 | Extract 4 - 3 times measured
| 81 | Young AA-4 2
| 82 | Young AA-4 3
| 83 | Young AA-5 1 | Extract 5 - 3 times measured
| 84 | Young AA-5 2
| 85 | Young AA-5 3
| 86 | Young AA-6 1 | Extract 6 - 3 times measured
| 87 | Young AA-6 2
| 88 | Young AA-6 3
| 89 | Pooled QC 7
| 90 | Pooled QC 8
| 91 | hTAG-1 1 | Extract 1 - 3 times measured
| 92 | hTAG-1 2
| 93 | hTAG-1 3
| 94 | hTAG-2 1 | Extract 2 - 3 times measured
| 95 | hTAG-2 2
| 96 | hTAG-2 3
| 97 | hTAG-3 1 | Extract 3 - 3 times measured
| 98 | hTAG-3 2
| 99 | hTAG-3 3
| 100 | hTAG-4 1 | Extract 4 - 3 times measured
| 101 | hTAG-4 2
| 102 | hTAG-4 3
| 103 | hTAG-5 1 | Extract 5 - 3 times measured
| 104 | hTAG-5 2
| 105 | hTAG-5 3
| 106 | hTAG-6 1 | Extract 6 - 3 times measured
| 107 | hTAG-6 2
| 108 | hTAG-6 3
| 109 | Pooled QC 9
| 110 | Pooled QC 10
| 111 | Cal1 1 2 | Extracted calibration 1
| 112 | Cal1 2 2 | Extracted calibration 2
| 113 | Cal1 3 2 | Extracted calibration 3
| 114 | Cal2 1 2 | Extracted calibration 1
| 115 | Cal2 2 2 | Extracted calibration 2
| 116 | Cal2 3 2 | Extracted calibration 3
| 117 | Cal3 1 2 | Extracted calibration 1
| 118 | Cal3 2 2 | Extracted calibration 2
| 119 | Cal3 3 2 | Extracted calibration 3
| 120 | Cal4 1 2 | Extracted calibration 1
| 121 | Cal4 2 2 | Extracted calibration 2
| 122 | Cal4 3 2 | Extracted calibration 3
| 123 | Cal5 1 2 | Extracted calibration 1
| 124 | Cal5 2 2 | Extracted calibration 2
| 125 | Cal5 3 2 | Extracted calibration 3
| 126 | Cal6 1 2 | Extracted calibration 1
| 127 | Cal6 2 2 | Extracted calibration 2
| 128 | Cal6 3 2 | Extracted calibration 3
| 129 | Matrix blank 2 1 | Extracted Matrix Blank 2-3 times measured
| 130 | Matrix blank 2 2
| 131 | Matrix blank 2 3
| 132 | Total Blank 3 1 | Extracted Total Blank 3-3 times measured
| 133 | Total Blank 3 2
| 134 | Total Blank 3 3
| 135 | Total Blank 4 1 2 | Extracted Total Blank 1-3 times measured | This part was repeated incl. Extraction as HQC and HLQC were wrongly prepared (+ 3x3 Matrix blank)
| 136 | Total Blank 4 2 2
| 137 | Total Blank 4 3 2
| 138 | Pooled QC 11
| 139 | Pooled QC 12
| 140 | LLQC1 1 | Extract 1 - 3 times measured
| 141 | LLQC1 2
| 142 | LLQC1 3
| 143 | LLQC2 1 | Extract 2 - 3 times measured
| 144 | LLQC2 2
| 145 | LLQC2 3
| 146 | LLQC3 1 | Extract 3 - 3 times measured
| 147 | LLQC3 2
| 148 | LLQC3 3
| 149 | LLQC4 1 | Extract 4 - 3 times measured
| 150 | LLQC4 2
| 151 | LLQC4 3
| 152 | LLQC5 1 | Extract 5 - 3 times measured
| 153 | LLQC5 2
| 154 | LLQC5 3
| 155 | LLQC6 1 | Extract 6 - 3 times measured
| 156 | LLQC6 2
| 157 | LLQC6 3
| 158 | Pooled QC 13
| 159 | Pooled QC 14
| 160 | LQC1 1 | Extract 1 - 3 times measured
| 161 | LQC1 2
| 162 | LQC1 3
| 163 | LQC2 1 | Extract 2 - 3 times measured
| 164 | LQC2 2
| 165 | LQC2 3
| 166 | LQC3 1 | Extract 3 - 3 times measured
| 167 | LQC3 2
| 168 | LQC3 3
| 169 | LQC4 1 | Extract 4 - 3 times measured
| 170 | LQC4 2
| 171 | LQC4 3
| 172 | LQC5 1 | Extract 5 - 3 times measured
| 173 | LQC5 2
| 174 | LQC5 3
| 175 | LQC6 1 | Extract 6 - 3 times measured
| 176 | LQC6 2
| 177 | LQC6 3
| 178 | Total Blank 5 1 2 | Extracted Total Blank 2-3 times measured
| 179 | Total Blank 5 2 2
| 180 | Total Blank 5 3 2
| 181 | Pooled QC 15
| 182 | Pooled QC 16
| 183 | MQC1 1 | Extract 1 - 3 times measured
| 184 | MQC1 2
| 185 | MQC1 3
| 186 | MQC2 1 | Extract 2 - 3 times measured
| 187 | MQC2 2
| 188 | MQC2 3
| 189 | MQC3 1 | Extract 3 - 3 times measured
| 190 | MQC3 2
| 191 | MQC3 3
| 192 | MQC4 1 | Extract 4 - 3 times measured
| 193 | MQC4 2
| 194 | MQC4 3
| 195 | MQC5 1 | Extract 5 - 3 times measured
| 196 | MQC5 2
| 197 | MQC5 3
| 198 | MQC6 1 | Extract 6 - 3 times measured
| 199 | MQC6 2
| 200 | MQC6 3
| 201 | Pooled QC 17
| 202 | Pooled QC 18
| 203 | HQC1 1 | Extract 1 - 3 times measured
| 204 | HQC1 2
| 205 | HQC1 3
| 206 | HQC2 1 | Extract 2 - 3 times measured
| 207 | HQC2 2
| 208 | HQC2 3
| 209 | HQC3 1 | Extract 3 - 3 times measured
| 210 | HQC3 2
| 211 | HQC3 3
| 212 | HQC4 1 | Extract 4 - 3 times measured
| 213 | HQC4 2
| 214 | HQC4 3
| 215 | HQC5 1 | Extract 5 - 3 times measured
| 216 | HQC5 2
| 217 | HQC5 3
| 218 | HQC6 1 | Extract 6 - 3 times measured
| 219 | HQC6 2
| 220 | HQC6 3
| 221 | Pooled QC 19
| 222 | Pooled QC 20
| 223 | HLQC1 1 | Extract 1 - 3 times measured
| 224 | HLQC1 2
| 225 | HLQC1 3
| 226 | HLQC2 1 | Extract 2 - 3 times measured
| 227 | HLQC2 2
| 228 | HLQC2 3
| 229 | HLQC3 1 | Extract 3 - 3 times measured
| 230 | HLQC3 2
| 231 | HLQC3 3
| 232 | HLQC4 1 | Extract 4 - 3 times measured
| 233 | HLQC4 2
| 234 | HLQC4 3
| 235 | HLQC5 1 | Extract 5 - 3 times measured
| 236 | HLQC5 2
| 237 | HLQC5 3
| 238 | HLQC6 1 | Extract 6 - 3 times measured
| 239 | HLQC6 2
| 240 | HLQC6 3
| 241 | Pooled QC 21
| 242 | Pooled QC 22
| 243 | Total Blank 6 1 2 | Extracted Total Blank 3-3 times measured
| 244 | Total Blank 6 2 2
| 245 | Total Blank 6 3 2

## LC-MRM - Intra Assay QC Proc
Intra Assay Procedure
Intra Assay Procedure | Quality Controls for Intra-Assay Variation
| Step | Description | Reference Table
| 1 | Prepare the Pooled NIST QC sample containing equal volumes of the four NIST reference plasma you received.
| 2 | The pooled QC representing middle quality control (MQC), low QC (LQC), and lower limit of quantification QC (LLOQ) is prepared by diluting MQC (e.g. two- and three-fold) with water.
| 3 | High QC (HQC) and upper limit of quantification (ULOQ) QC are prepared by spiking in 10 µL and 20 µL respectively of STD3 endogenous standards from table | Extraction - Overview'!B34
| 4 | Internal standards are added to all QC samples.
Intra-Assay Variation
| 6 | The precision and accuracy of the assay is determined for each of the four ceramides measured in LLOQ, LQC, MQC, HQC, and HLOQ.
| 7 | Intra-assay variation, precision, and accuracy is calculated for each of the ceramide in replicates of six (n = 6) at each QC concentration independently on three separate occasions.
| 8 | The intra-assay precision (percentage coefficient variance, %CV) and accuracy (percentage accuracy, %Accuracy) are calculated from the nominal concentrations according to the formulas reported in corresponding sheets and tables of this document.
| 9 | Recovery estimation is not necessary for the published method but only for different ones. Stability estimation is not necessary for the published method but only for different ones. Calibration line is prepared in 5% BSA.

## LC-MRM - Intra Assay QC Res
INTRA ASSAY - QC RESULTS
INTRA ASSAY - QC RESULTS | Table 12 LC-MRM Calibration Line 1 Data (Beginning). Please report the values of each of the three replicates separately for each compound and standard.
| Calibration Line 1 | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| STD 1 | 2800000 | 2620000 | 2800000 | 3070000 | 2910000 | 3130000 | 50900000 | 45900000 | 50400000 | 39700000 | 36100000 | 39400000 | 140000 | 142000 | 135000 | 124000 | 128000 | 125000 | 3580000 | 3520000 | 3640000 | 995000 | 984000 | 946000
| STD 2 | 1390000 | 1380000 | 1400000 | 1580000 | 1500000 | 1540000 | 22700000 | 21300000 | 21400000 | 17100000 | 16800000 | 16700000 | 135000 | 133000 | 136000 | 124000 | 118000 | 120000 | 3150000 | 2990000 | 3070000 | 858000 | 808000 | 825000
| STD 3 | 142000 | 149000 | 143000 | 157000 | 167000 | 156000 | 1700000 | 1800000 | 1690000 | 1370000 | 1410000 | 1380000 | 134000 | 138000 | 139000 | 122000 | 123000 | 125000 | 2330000 | 2440000 | 2400000 | 648000 | 664000 | 672000
| STD 4 | 35600 | 31900 | 28500 | 31800 | 31900 | 32000 | 366000 | 346000 | 337000 | 274000 | 279000 | 255000 | 138000 | 137000 | 135000 | 121000 | 120000 | 118000 | 2410000 | 2340000 | 2320000 | 648000 | 641000 | 640000
| STD 5 | 15300 | 17500 | 17200 | 15900 | 16100 | 15500 | 177000 | 176000 | 187000 | 132000 | 131000 | 140000 | 136000 | 135000 | 141000 | 119000 | 121000 | 121000 | 2300000 | 2290000 | 2320000 | 657000 | 655000 | 648000
| STD 6 | 14200 | 14300 | 15000 | 14400 | 13000 | 12000 | 145000 | 146000 | 152000 | 110000 | 108000 | 112000 | 131000 | 136000 | 133000 | 121000 | 123000 | 121000 | 2270000 | 2310000 | 2330000 | 649000 | 641000 | 640000
Table 13 LC-MRM Calibration Line 2 Data (End). Please report the values of each of the three replicates separately for each compound and standard.
| Calibration Line 2 | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| STD 1 | 3620000 | 3360000 | 3560000 | 3840000 | 3660000 | 3920000 | 61900000 | 56300000 | 61500000 | 47800000 | 43700000 | 47500000 | 179000 | 184000 | 175000 | 155000 | 159000 | 157000 | 4310000 | 4330000 | 4170000 | 1140000 | 1150000 | 1150000
| STD 2 | 1740000 | 1760000 | 1820000 | 1910000 | 1890000 | 2010000 | 26000000 | 25500000 | 27800000 | 19600000 | 19100000 | 20600000 | 167000 | 173000 | 175000 | 151000 | 148000 | 162000 | 3590000 | 3490000 | 3880000 | 929000 | 952000 | 1010000
| STD 3 | 181000 | 194000 | 178000 | 197000 | 201000 | 195000 | 2000000 | 2040000 | 2070000 | 1580000 | 1580000 | 1560000 | 177000 | 179000 | 175000 | 152000 | 155000 | 154000 | 2730000 | 2740000 | 2750000 | 762000 | 739000 | 758000
| STD 4 | 40300 | 41500 | 48300 | 41600 | 40500 | 45100 | 410000 | 421000 | 437000 | 313000 | 317000 | 320000 | 180000 | 176000 | 179000 | 163000 | 152000 | 163000 | 2710000 | 2760000 | 2840000 | 742000 | 747000 | 735000
| STD 5 | 21800 | 21100 | 22100 | 22900 | 21400 | 19600 | 231000 | 223000 | 216000 | 168000 | 159000 | 159000 | 178000 | 183000 | 176000 | 161000 | 152000 | 153000 | 2740000 | 2710000 | 2740000 | 750000 | 753000 | 738000
| STD 6 | 17200 | 19300 | 18400 | 15900 | 19100 | 17800 | 183000 | 190000 | 185000 | 131000 | 129000 | 125000 | 182000 | 182000 | 183000 | 160000 | 162000 | 157000 | 2780000 | 2880000 | 2780000 | 747000 | 754000 | 754000
Table 14 LC-MRM Sample Data. Please report the values of each of the three replicates separately for each compound and standard.
| NIST SRM | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| SRM 1 | 212000 | 213000 | 205000 | 95200 | 92700 | 91300 | 3240000 | 3340000 | 3260000 | 866000 | 890000 | 880000 | 157000 | 160000 | 162000 | 178000 | 183000 | 174000 | 3620000 | 3780000 | 3740000 | 964000 | 1000000 | 999000
| SRM 2 | 202000 | 202000 | 204000 | 89400 | 92600 | 91900 | 3060000 | 3050000 | 3220000 | 820000 | 860000 | 856000 | 159000 | 161000 | 157000 | 176000 | 177000 | 178000 | 3520000 | 3560000 | 3670000 | 929000 | 936000 | 971000
| SRM 3 | 201000 | 197000 | 199000 | 92100 | 89800 | 87600 | 3130000 | 2900000 | 2950000 | 824000 | 785000 | 812000 | 152000 | 155000 | 156000 | 178000 | 171000 | 170000 | 3560000 | 3410000 | 3460000 | 958000 | 918000 | 909000
| SRM 4 | 204000 | 198000 | 203000 | 95300 | 87300 | 93600 | 3200000 | 3010000 | 3070000 | 844000 | 814000 | 794000 | 155000 | 157000 | 157000 | 181000 | 168000 | 176000 | 3670000 | 3500000 | 3580000 | 961000 | 933000 | 957000
| SRM 5 | 210000 | 212000 | 208000 | 92300 | 92600 | 96800 | 3210000 | 3290000 | 3250000 | 849000 | 875000 | 846000 | 163000 | 162000 | 165000 | 187000 | 179000 | 182000 | 3710000 | 3700000 | 3710000 | 997000 | 982000 | 993000
| SRM 6 | 206000 | 211000 | 212000 | 97200 | 93600 | 98600 | 3350000 | 3300000 | 3230000 | 861000 | 853000 | 851000 | 158000 | 157000 | 158000 | 182000 | 181000 | 181000 | 3670000 | 3500000 | 3490000 | 959000 | 947000 | 949000
| NIST hTAG | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| hTAG 1 | 230000 | 228000 | 231000 | 107000 | 111000 | 110000 | 5180000 | 5190000 | 5330000 | 1180000 | 1180000 | 1210000 | 164000 | 161000 | 167000 | 192000 | 198000 | 197000 | 4120000 | 4150000 | 4290000 | 1030000 | 1010000 | 1020000
| hTAG 2 | 231000 | 232000 | 232000 | 118000 | 118000 | 117000 | 5370000 | 5530000 | 5430000 | 1240000 | 1220000 | 1240000 | 157000 | 159000 | 163000 | 197000 | 190000 | 190000 | 4160000 | 4210000 | 4150000 | 1020000 | 998000 | 1030000
| hTAG 3 | 237000 | 233000 | 235000 | 119000 | 118000 | 116000 | 5350000 | 5400000 | 5510000 | 1230000 | 1220000 | 1230000 | 164000 | 164000 | 160000 | 199000 | 193000 | 200000 | 4120000 | 4080000 | 4140000 | 1040000 | 1010000 | 976000
| hTAG 4 | 234000 | 238000 | 234000 | 115000 | 123000 | 116000 | 5330000 | 5430000 | 5380000 | 1190000 | 1230000 | 1230000 | 162000 | 167000 | 161000 | 192000 | 201000 | 205000 | 4030000 | 4130000 | 4160000 | 1000000 | 1020000 | 1000000
| hTAG 5 | 235000 | 232000 | 238000 | 123000 | 120000 | 111000 | 5400000 | 5230000 | 5310000 | 1280000 | 1230000 | 1190000 | 166000 | 165000 | 165000 | 200000 | 203000 | 199000 | 4280000 | 4160000 | 4190000 | 1050000 | 1030000 | 998000
| hTAG 6 | 229000 | 236000 | 226000 | 115000 | 120000 | 117000 | 5160000 | 5320000 | 5160000 | 1190000 | 1230000 | 1200000 | 160000 | 166000 | 164000 | 197000 | 202000 | 202000 | 4110000 | 4230000 | 4100000 | 984000 | 1010000 | 1010000
| NIST T1D | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| T1D 1 | 215000 | 212000 | 217000 | 141000 | 147000 | 145000 | 3200000 | 3040000 | 3170000 | 854000 | 826000 | 838000 | 192000 | 196000 | 193000 | 245000 | 250000 | 241000 | 3620000 | 3590000 | 3620000 | 945000 | 895000 | 915000
| T1D 2 | 209000 | 207000 | 212000 | 144000 | 139000 | 138000 | 2980000 | 2940000 | 2950000 | 812000 | 788000 | 807000 | 191000 | 187000 | 197000 | 240000 | 237000 | 244000 | 3510000 | 3430000 | 3440000 | 924000 | 895000 | 926000
| T1D 3 | 208000 | 208000 | 209000 | 138000 | 134000 | 139000 | 2930000 | 2920000 | 2860000 | 773000 | 755000 | 744000 | 196000 | 200000 | 196000 | 244000 | 241000 | 243000 | 3410000 | 3590000 | 3540000 | 928000 | 936000 | 935000
| T1D 4 | 195000 | 208000 | 212000 | 137000 | 139000 | 135000 | 2830000 | 2850000 | 2890000 | 743000 | 751000 | 785000 | 195000 | 195000 | 199000 | 242000 | 251000 | 251000 | 3440000 | 3440000 | 3460000 | 920000 | 936000 | 919000
| T1D 5 | 216000 | 208000 | 210000 | 144000 | 138000 | 143000 | 2910000 | 2910000 | 2950000 | 789000 | 770000 | 795000 | 192000 | 193000 | 187000 | 237000 | 238000 | 241000 | 3490000 | 3360000 | 3490000 | 910000 | 887000 | 888000
| T1D 6 | 205000 | 209000 | 207000 | 144000 | 142000 | 138000 | 2820000 | 2940000 | 2980000 | 772000 | 775000 | 794000 | 194000 | 198000 | 193000 | 247000 | 244000 | 251000 | 3500000 | 3470000 | 3490000 | 928000 | 907000 | 927000
| NIST Young AA | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| yA-A 1 | 142000 | 136000 | 139000 | 56000 | 56100 | 57400 | 1960000 | 1950000 | 1960000 | 567000 | 555000 | 534000 | 183000 | 180000 | 179000 | 186000 | 184000 | 187000 | 3430000 | 3360000 | 3320000 | 1050000 | 1000000 | 1000000
| yA-A 2 | 140000 | 144000 | 139000 | 54100 | 56000 | 55300 | 1970000 | 1950000 | 1960000 | 558000 | 554000 | 550000 | 178000 | 181000 | 178000 | 185000 | 189000 | 180000 | 3280000 | 3390000 | 3300000 | 1030000 | 1010000 | 985000
| yA-A 3 | 142000 | 135000 | 142000 | 54000 | 54200 | 54600 | 1930000 | 1910000 | 1960000 | 549000 | 543000 | 561000 | 187000 | 177000 | 186000 | 189000 | 196000 | 197000 | 3400000 | 3390000 | 3440000 | 1030000 | 1090000 | 1050000
| yA-A 4 | 144000 | 135000 | 143000 | 54300 | 54800 | 55400 | 1980000 | 1920000 | 1980000 | 567000 | 546000 | 571000 | 181000 | 177000 | 177000 | 189000 | 187000 | 183000 | 3360000 | 3280000 | 3310000 | 1020000 | 998000 | 1010000
| yA-A 5 | 140000 | 135000 | 144000 | 54000 | 56600 | 56400 | 1950000 | 1960000 | 1950000 | 549000 | 546000 | 573000 | 185000 | 181000 | 182000 | 193000 | 183000 | 191000 | 3360000 | 3300000 | 3340000 | 1040000 | 1010000 | 1030000
| yA-A 6 | 144000 | 149000 | 144000 | 58500 | 57700 | 57400 | 1970000 | 1940000 | 1860000 | 544000 | 551000 | 531000 | 186000 | 184000 | 186000 | 198000 | 193000 | 195000 | 3390000 | 3420000 | 3310000 | 1040000 | 1050000 | 1030000
Table 15 LC-MRM Intra Assay QC Data. Please report the values of each of the three replicates separately for each compound and standard.
| Lowest Level QC | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| LLQC 1 | 74100 | 78500 | 79100 | 32000 | 32500 | 33200 | 957000 | 932000 | 946000 | 247000 | 233000 | 253000 | 194000 | 193000 | 191000 | 177000 | 174000 | 182000 | 2910000 | 2890000 | 2960000 | 799000 | 794000 | 806000
| LLQC 2 | 80400 | 81200 | 80200 | 33400 | 33200 | 33300 | 930000 | 927000 | 925000 | 250000 | 243000 | 243000 | 195000 | 197000 | 195000 | 188000 | 182000 | 183000 | 2960000 | 3040000 | 2990000 | 814000 | 806000 | 780000
| LLQC 3 | 82400 | 82100 | 78300 | 34000 | 32900 | 31800 | 926000 | 946000 | 937000 | 251000 | 238000 | 242000 | 196000 | 195000 | 198000 | 182000 | 183000 | 180000 | 2960000 | 3020000 | 2970000 | 807000 | 799000 | 812000
| LLQC 4 | 81200 | 75900 | 78600 | 31600 | 29800 | 32400 | 883000 | 899000 | 912000 | 230000 | 228000 | 240000 | 200000 | 199000 | 196000 | 181000 | 175000 | 178000 | 2950000 | 2920000 | 2950000 | 805000 | 776000 | 805000
| LLQC 5 | 81600 | 80700 | 79600 | 32000 | 32900 | 31500 | 928000 | 927000 | 929000 | 240000 | 238000 | 241000 | 199000 | 195000 | 200000 | 184000 | 187000 | 186000 | 2970000 | 2950000 | 2940000 | 786000 | 765000 | 785000
| LLQC 6 | 80900 | 78900 | 81400 | 32300 | 30300 | 32100 | 886000 | 907000 | 913000 | 235000 | 233000 | 243000 | 201000 | 203000 | 200000 | 186000 | 184000 | 179000 | 2810000 | 2940000 | 2930000 | 786000 | 775000 | 784000
| Low QC | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| LQC 1 | 124000 | 123000 | 126000 | 56000 | 53500 | 55500 | 1630000 | 1630000 | 1590000 | 416000 | 407000 | 421000 | 193000 | 194000 | 190000 | 196000 | 196000 | 193000 | 3260000 | 3170000 | 3140000 | 860000 | 852000 | 846000
| LQC 2 | 125000 | 132000 | 129000 | 55500 | 56500 | 53900 | 1720000 | 1710000 | 1680000 | 428000 | 432000 | 427000 | 200000 | 198000 | 191000 | 202000 | 204000 | 199000 | 3300000 | 3340000 | 3300000 | 879000 | 879000 | 857000
| LQC 3 | 119000 | 123000 | 118000 | 54000 | 52100 | 48800 | 1470000 | 1530000 | 1500000 | 379000 | 396000 | 385000 | 193000 | 198000 | 186000 | 194000 | 193000 | 190000 | 3070000 | 3080000 | 3080000 | 803000 | 807000 | 831000
| LQC 4 | 120000 | 120000 | 124000 | 52700 | 51800 | 51800 | 1540000 | 1520000 | 1470000 | 389000 | 381000 | 386000 | 201000 | 199000 | 198000 | 195000 | 197000 | 196000 | 3150000 | 3220000 | 3180000 | 828000 | 829000 | 829000
| LQC 5 | 116000 | 119000 | 122000 | 53600 | 51600 | 51400 | 1460000 | 1500000 | 1460000 | 369000 | 397000 | 388000 | 194000 | 199000 | 201000 | 190000 | 199000 | 201000 | 3060000 | 3140000 | 3160000 | 822000 | 859000 | 837000
| LQC 6 | 123000 | 123000 | 117000 | 52500 | 54800 | 52700 | 1500000 | 1500000 | 1480000 | 390000 | 380000 | 386000 | 203000 | 206000 | 210000 | 204000 | 201000 | 202000 | 3270000 | 3190000 | 3220000 | 860000 | 861000 | 853000
| Middle QC | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| MQC 1 | 207000 | 214000 | 206000 | 119000 | 115000 | 113000 | 3580000 | 3550000 | 3430000 | 914000 | 933000 | 931000 | 177000 | 182000 | 181000 | 224000 | 222000 | 224000 | 3820000 | 3840000 | 3780000 | 1020000 | 1040000 | 999000
| MQC 2 | 207000 | 211000 | 214000 | 114000 | 112000 | 114000 | 3620000 | 3550000 | 3600000 | 961000 | 941000 | 935000 | 183000 | 177000 | 182000 | 229000 | 229000 | 229000 | 3780000 | 3840000 | 3870000 | 1040000 | 1030000 | 1030000
| MQC 3 | 204000 | 202000 | 202000 | 106000 | 109000 | 107000 | 3210000 | 3250000 | 3270000 | 848000 | 855000 | 859000 | 176000 | 185000 | 176000 | 215000 | 219000 | 217000 | 3620000 | 3640000 | 3770000 | 957000 | 964000 | 997000
| MQC 4 | 203000 | 207000 | 213000 | 111000 | 110000 | 111000 | 3390000 | 3490000 | 3430000 | 878000 | 925000 | 930000 | 186000 | 186000 | 188000 | 226000 | 234000 | 234000 | 3890000 | 3910000 | 3870000 | 1020000 | 1030000 | 1060000
| MQC 5 | 218000 | 215000 | 222000 | 121000 | 123000 | 119000 | 3750000 | 3720000 | 3790000 | 976000 | 965000 | 947000 | 178000 | 178000 | 178000 | 231000 | 226000 | 240000 | 3940000 | 3820000 | 3770000 | 1040000 | 1020000 | 1060000
| MQC 6 | 212000 | 203000 | 208000 | 110000 | 112000 | 113000 | 3330000 | 3310000 | 3290000 | 854000 | 878000 | 858000 | 184000 | 181000 | 181000 | 224000 | 226000 | 230000 | 3730000 | 3730000 | 3730000 | 1000000 | 980000 | 992000
| High QC | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| HQC 1 | 308000 | 306000 | 304000 | 260000 | 262000 | 269000 | 5100000 | 5100000 | 5160000 | 2040000 | 2040000 | 2070000 | 177000 | 182000 | 184000 | 234000 | 234000 | 233000 | 3930000 | 3850000 | 3770000 | 1030000 | 1050000 | 1010000
| HQC 2 | 313000 | 324000 | 310000 | 267000 | 264000 | 259000 | 5160000 | 5150000 | 5100000 | 2050000 | 2080000 | 2070000 | 181000 | 178000 | 179000 | 231000 | 232000 | 226000 | 3740000 | 3810000 | 3860000 | 1050000 | 1010000 | 1010000
| HQC 3 | 311000 | 316000 | 317000 | 264000 | 267000 | 265000 | 5110000 | 5190000 | 5240000 | 2090000 | 2090000 | 2030000 | 179000 | 182000 | 184000 | 225000 | 223000 | 231000 | 3710000 | 3710000 | 3820000 | 994000 | 1010000 | 1010000
| HQC 4 | 321000 | 322000 | 320000 | 291000 | 279000 | 278000 | 5230000 | 5240000 | 5320000 | 2090000 | 2120000 | 2160000 | 182000 | 181000 | 185000 | 234000 | 246000 | 249000 | 3900000 | 3860000 | 4000000 | 1040000 | 1060000 | 1070000
| HQC 5 | 320000 | 315000 | 306000 | 270000 | 263000 | 270000 | 5120000 | 4900000 | 5010000 | 2100000 | 2040000 | 2040000 | 183000 | 185000 | 184000 | 232000 | 236000 | 234000 | 3810000 | 3720000 | 3670000 | 1000000 | 992000 | 988000
| HQC 6 | 321000 | 320000 | 313000 | 284000 | 275000 | 275000 | 5280000 | 5160000 | 5060000 | 2130000 | 2070000 | 2030000 | 181000 | 184000 | 181000 | 239000 | 234000 | 234000 | 3890000 | 3820000 | 3840000 | 1020000 | 1010000 | 952000
| Highest Level QC | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| HLQC 1 | 411000 | 411000 | 411000 | 407000 | 414000 | 413000 | 6460000 | 6270000 | 6530000 | 3110000 | 3050000 | 3190000 | 189000 | 188000 | 191000 | 233000 | 238000 | 247000 | 3840000 | 3740000 | 3740000 | 997000 | 995000 | 1040000
| HLQC 2 | 392000 | 403000 | 402000 | 408000 | 418000 | 409000 | 6060000 | 6350000 | 6400000 | 3020000 | 3150000 | 3060000 | 181000 | 184000 | 186000 | 232000 | 235000 | 236000 | 3690000 | 3690000 | 3720000 | 970000 | 1040000 | 1020000
| HLQC 3 | 408000 | 405000 | 412000 | 415000 | 407000 | 424000 | 6350000 | 6400000 | 6340000 | 3100000 | 3060000 | 3080000 | 187000 | 183000 | 186000 | 242000 | 240000 | 241000 | 3770000 | 3670000 | 3700000 | 1020000 | 975000 | 1030000
| HLQC 4 | 418000 | 405000 | 414000 | 460000 | 428000 | 431000 | 6650000 | 6330000 | 6460000 | 3200000 | 3130000 | 3140000 | 188000 | 187000 | 185000 | 256000 | 246000 | 247000 | 3950000 | 3730000 | 3750000 | 1070000 | 1010000 | 1020000
| HLQC 5 | 418000 | 419000 | 425000 | 437000 | 441000 | 431000 | 6430000 | 6440000 | 6470000 | 3210000 | 3080000 | 3230000 | 193000 | 193000 | 190000 | 237000 | 244000 | 246000 | 3840000 | 3840000 | 3740000 | 1040000 | 1000000 | 1020000
| HLQC 6 | 418000 | 421000 | 413000 | 430000 | 424000 | 442000 | 6610000 | 6490000 | 6430000 | 3240000 | 3220000 | 3180000 | 191000 | 193000 | 192000 | 250000 | 250000 | 247000 | 3850000 | 3780000 | 3820000 | 1040000 | 1010000 | 995000
| Blank QC | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| Total Blank 1 | 5130 | 3260 | 5880 | 3080 | 798 | 1210 | 27300 | 9770 | 11400 | 2750 | 952 | 1170 | 834 | 179 | 381 | 385 | 189 | 377 | 705 | 2470 | 3130 | 567 | 453 | 338
| Total Blank 2 | 1900 | 1560 | 3500 | 1110 | 1130 | 1730 | 19400 | 6890 | 11100 | 1020 | 576 | 732 | 206 | 68.9 | 155 | 291 | 156 | 342 | 364 | 1050 | 4000 | 171 | 93 | 540
| Total Blank 3 | 1830 | 1620 | 3860 | 1250 | 917 | 2160 | 12700 | 4990 | 11800 | 720 | 620 | 882 | 247 | 123 | 158 | 276 | 155 | 144 | 1670 | 3230 | 3430 | 121 | 279 | 494
| Matrix Blank 1 | 194000 | 193000 | 190000 | 103000 | 110000 | 112000 | 3310000 | 3130000 | 3120000 | 876000 | 840000 | 820000 | 139 | 128 | 275 | 502 | 145 | 127 | 701 | 495 | 1810 | 165 | 133 | 152
| Matrix Blank 2 | 198000 | 197000 | 195000 | 109000 | 112000 | 115000 | 3050000 | 3070000 | 3120000 | 789000 | 839000 | 796000 | 139 | 305 | 199 | 120 | 217 | 122 | 378 | 2490 | 1650 | 80.8 | 267 | 52.4
| Matrix Blank 3 | 201000 | 202000 | 204000 | 111000 | 110000 | 111000 | 3060000 | 3100000 | 3150000 | 807000 | 809000 | 836000 | 124 | 274 | 352 | 135 | 267 | 855 | 1390 | 4810 | 14000 | 107 | 213 | 600

## Raw Peak Areas
Sample_Name | Cer d18:1/16:0 [M>264] | Cer d18:1/16:0-d7 [M>264] | Cer d18:1/18:0 [M>264] | Cer d18:1/18:0-d7 [M>264] | Cer d18:1/24:0 [M>264] | Cer d18:1/24:0-d7 [M>264] | Cer d18:1/24:1 [M>264] | Cer d18:1/24:1-d7 [M>264]
Total Blank 1
Matrix blank
Calibration line 1 from STD6 to 1 (n=3)
[…]
Pooled QC
Pooled QC
SRM 1950-1
SRM 1950-2
SRM 1950-3
SRM 1950-4
SRM 1950-5
SRM 1950-6
Pooled QC
Pooled QC
T1D-1
T1D-2
T1D-3
T1D-4
T1D-5
T1D-6
Total Blank 2
Pooled QC
Pooled QC
Young AA-1
Young AA-2
Young AA-3
Young AA-4
Young AA-5
Young AA-6
Pooled QC
Pooled QC
hTAG-1
hTAG-2
hTAG-3
hTAG-4
hTAG-5
hTAG-6
Pooled QC
Pooled QC
Calibration line 2 from STD1 to 6 (n=3)
[…]
Matrix blank
Total Blank 3

## QTrap_Final
|  | Cer d18:1/16:0 | D7 Cer d18:1/16:0 | D7 Cer d18:1/16:0 | Cer d18:1/16:0 | calc. With IS | calc. With Cal |  |  | Cer d18:1/18:0 | D7 Cer d18:1/18:0 | D7 Cer d18:1/18:0 | Cer d18:1/18:0 | calc. With IS | calc. With Cal |  |  | Cer d18:1/24:0 | D7 Cer d18:1/24:0 | D7 Cer d18:1/24:0 | Cer d18:1/24:0 | calc. With IS | calc. With Cal |  |  | Cer d18:1/24:1 | D7 Cer d18:1/24:1 | D7 Cer d18:1/24:1 | Cer d18:1/24:1 | calc. With IS | calc. With Cal
|  | Analyte Peak Area (counts) | Analyte Peak Area (counts) | c [pmol/µL] | c [pmol/µL] | c [pmol/µL] | c [pmol/µL] |  |  | Analyte Peak Area (counts) | Analyte Peak Area (counts) | c [pmol/µL] | c [pmol/µL] | c [pmol/µL] | c [pmol/µL] |  |  | Analyte Peak Area (counts) | Analyte Peak Area (counts) | c [pmol/µL] | c [pmol/µL] | c [pmol/µL] | c [pmol/µL] |  |  | Analyte Peak Area (counts) | Analyte Peak Area (counts) | c [pmol/µL] | c [pmol/µL] | c [pmol/µL] | c [pmol/µL]
1 | Cal1 1 | 2800000 | 140000 | 0.125 | 2 | 2.5 | 2.0032 |  | Cal1 1 | 3070000 | 124000 | 0.05 | 2 | 1.2379 | 1.9892 |  | Cal1 1 | 50900000 | 3580000 | 1.5 | 20 | 21.3268 | 20.3974 |  | Cal1 1 | 39700000 | 995000 | 0.5 | 20 | 19.9497 | 20.5152
2 | Cal2 1 | 1390000 | 135000 | 0.125 | 1 | 1.287 | 0.99356 |  | Cal2 1 | 1580000 | 124000 | 0.05 | 1 | 0.6371 | 1.0219 |  | Cal2 1 | 22700000 | 3150000 | 1.5 | 10 | 10.8095 | 9.2109 |  | Cal2 1 | 17100000 | 858000 | 0.5 | 10 | 9.965 | 8.9692
3 | Cal3 1 | 142000 | 134000 | 0.125 | 0.1 | 0.13246 | 0.099919 |  | Cal3 1 | 157000 | 122000 | 0.05 | 0.1 | 0.064344 | 0.098088 |  | Cal3 1 | 1700000 | 2330000 | 1.5 | 1 | 1.0944 | 0.88047 |  | Cal3 1 | 1370000 | 648000 | 0.5 | 1 | 1.0571 | 0.93294
4 | Cal4 1 | 35600 | 138000 | 0.125 | 0.02 | 0.032246 | 0.023731 |  | Cal4 1 | 31800 | 121000 | 0.05 | 0.02 | 0.01314 | 0.016809 |  | Cal4 1 | 366000 | 2410000 | 1.5 | 0.2 | 0.2278 | 0.35129 |  | Cal4 1 | 274000 | 648000 | 0.5 | 0.2 | 0.21142 | 0.37301
5 | Cal5 1 | 15300 | 136000 | 0.125 | 0.01 | 0.014063 | 0.0091953 |  | Cal5 1 | 15900 | 119000 | 0.05 | 0.01 | 0.0066807 | 0.0064863 |  | Cal5 1 | 177000 | 2300000 | 1.5 | 0.1 | 0.11543 | 0.27632 |  | Cal5 1 | 132000 | 657000 | 0.5 | 0.1 | 0.10046 | 0.30046
6 | Cal6 1 | 14200 | 131000 | 0.125 | 0.008 | 0.01355 | 0.0084076 |  | Cal6 1 | 14400 | 121000 | 0.05 | 0.008 | 0.0059504 | 0.0055125 |  | Cal6 1 | 145000 | 2270000 | 1.5 | 0.08 | 0.095815 | 0.26363 |  | Cal6 1 | 110000 | 649000 | 0.5 | 0.08 | 0.084746 | 0.28922
1 | Cal1 2 | 2620000 | 142000 | 0.125 | 2 | 2.3063 | 1.9809 |  | Cal1 2 | 2910000 | 128000 | 0.05 | 2 | 1.1367 | 1.9889 |  | Cal1 2 | 45900000 | 3520000 | 1.5 | 20 | 19.5597 | 20.2627 |  | Cal1 2 | 36100000 | 984000 | 0.5 | 20 | 18.3435 | 20.2523
2 | Cal2 2 | 1380000 | 133000 | 0.125 | 1 | 1.297 | 1.0383 |  | Cal2 2 | 1500000 | 118000 | 0.05 | 1 | 0.63559 | 1.0217 |  | Cal2 2 | 21300000 | 2990000 | 1.5 | 10 | 10.6856 | 9.4777 |  | Cal2 2 | 16800000 | 808000 | 0.5 | 10 | 10.396 | 9.499
3 | Cal3 2 | 149000 | 138000 | 0.125 | 0.1 | 0.13496 | 0.10255 |  | Cal3 2 | 167000 | 123000 | 0.05 | 0.1 | 0.067886 | 0.10731 |  | Cal3 2 | 1800000 | 2440000 | 1.5 | 1 | 1.1066 | 0.92854 |  | Cal3 2 | 1410000 | 664000 | 0.5 | 1 | 1.0617 | 0.92422
4 | Cal4 2 | 31900 | 137000 | 0.125 | 0.02 | 0.029106 | 0.013541 |  | Cal4 2 | 31900 | 120000 | 0.05 | 0.02 | 0.013292 | 0.014635 |  | Cal4 2 | 346000 | 2340000 | 1.5 | 0.2 | 0.22179 | 0.29108 |  | Cal4 2 | 279000 | 641000 | 0.5 | 0.2 | 0.21763 | 0.29407
5 | Cal5 2 | 17500 | 135000 | 0.125 | 0.01 | 0.016204 | 0.0025948 |  | Cal5 2 | 16100 | 121000 | 0.05 | 0.01 | 0.0066529 | 0.0037964 |  | Cal5 2 | 176000 | 2290000 | 1.5 | 0.1 | 0.11528 | 0.21655 |  | Cal5 2 | 131000 | 655000 | 0.5 | 0.1 | 0.1 | 0.21161
6 | Cal6 2 | 14300 | 136000 | 0.125 | 0.008 | 0.013143 | 0.00016232 |  | Cal6 2 | 13000 | 123000 | 0.05 | 0.008 | 0.0052846 | 0.0016699 |  | Cal6 2 | 146000 | 2310000 | 1.5 | 0.08 | 0.094805 | 0.2034 |  | Cal6 2 | 108000 | 641000 | 0.5 | 0.08 | 0.084243 | 0.19879
1 | Cal1 3 | 2800000 | 135000 | 0.125 | 2 | 2.5926 | 2.0004 |  | Cal1 3 | 3130000 | 125000 | 0.05 | 2 | 1.252 | 2.006 |  | Cal1 3 | 50400000 | 3640000 | 1.5 | 20 | 20.7692 | 20.563 |  | Cal1 3 | 39400000 | 946000 | 0.5 | 20 | 20.8245 | 20.5688
2 | Cal2 3 | 1400000 | 136000 | 0.125 | 1 | 1.2868 | 0.99927 |  | Cal2 3 | 1540000 | 120000 | 0.05 | 1 | 0.64167 | 0.98777 |  | Cal2 3 | 21400000 | 3070000 | 1.5 | 10 | 10.456 | 8.8735 |  | Cal2 3 | 16700000 | 825000 | 0.5 | 10 | 10.1212 | 8.8592
3 | Cal3 3 | 143000 | 139000 | 0.125 | 0.1 | 0.1286 | 0.10043 |  | Cal3 3 | 156000 | 125000 | 0.05 | 0.1 | 0.0624 | 0.10146 |  | Cal3 3 | 1690000 | 2400000 | 1.5 | 1 | 1.0562 | 0.92868 |  | Cal3 3 | 1380000 | 672000 | 0.5 | 1 | 1.0268 | 0.9565
4 | Cal4 3 | 28500 | 135000 | 0.125 | 0.02 | 0.026389 | 0.01856 |  | Cal4 3 | 32000 | 118000 | 0.05 | 0.02 | 0.013559 | 0.022046 |  | Cal4 3 | 337000 | 2320000 | 1.5 | 0.2 | 0.21789 | 0.3833 |  | Cal4 3 | 255000 | 640000 | 0.5 | 0.2 | 0.19922 | 0.37618
5 | Cal5 3 | 17200 | 141000 | 0.125 | 0.01 | 0.015248 | 0.01048 |  | Cal5 3 | 15500 | 121000 | 0.05 | 0.01 | 0.006405 | 0.01148 |  | Cal5 3 | 187000 | 2320000 | 1.5 | 0.1 | 0.12091 | 0.32284 |  | Cal5 3 | 140000 | 648000 | 0.5 | 0.1 | 0.10802 | 0.31686
6 | Cal6 3 | 15000 | 133000 | 0.125 | 0.008 | 0.014098 | 0.0089064 |  | Cal6 3 | 12000 | 121000 | 0.05 | 0.008 | 0.0049587 | 0.0092382 |  | Cal6 3 | 152000 | 2330000 | 1.5 | 0.08 | 0.097854 | 0.30873 |  | Cal6 3 | 112000 | 640000 | 0.5 | 0.08 | 0.0875 | 0.30241
|  |  | 2.0746 | [7 empty] | 2.0986 | [7 empty] | 18.3962 | [7 empty] | 17.2835
|  | Cer d18:1/16:0 | D7 Cer d18:1/16:0 | D7 Cer d18:1/16:0 | Cer d18:1/16:0 | calc. With IS | calc. With Cal |  |  | Cer d18:1/18:0 | D7 Cer d18:1/18:0 | D7 Cer d18:1/18:0 | Cer d18:1/18:0 | calc. With IS | calc. With Cal |  |  | Cer d18:1/24:0 | D7 Cer d18:1/24:0 | D7 Cer d18:1/24:0 | Cer d18:1/24:0 | calc. With IS | calc. With Cal |  |  | Cer d18:1/24:1 | D7 Cer d18:1/24:1 | D7 Cer d18:1/24:1 | Cer d18:1/24:1 | calc. With IS | calc. With Cal
|  | Analyte Peak Area (counts) | Analyte Peak Area (counts) | c [pmol/µL] | c [pmol/µL] | c [pmol/µL] | c [pmol/µL] |  |  | Analyte Peak Area (counts) | Analyte Peak Area (counts) | c [pmol/µL] | c [pmol/µL] | c [pmol/µL] | c [pmol/µL] |  |  | Analyte Peak Area (counts) | Analyte Peak Area (counts) | c [pmol/µL] | c [pmol/µL] | c [pmol/µL] | c [pmol/µL] |  |  | Analyte Peak Area (counts) | Analyte Peak Area (counts) | c [pmol/µL] | c [pmol/µL] | c [pmol/µL] | c [pmol/µL]
1 | Cal1 1 2 | 3620000 | 179000 | 0.125 | 2 | 2.5279 | 2.015 |  | Cal1 1 2 | 3840000 | 155000 | 0.05 | 2 | 1.2387 | 2.0022 |  | Cal1 1 2 | 61900000 | 4310000 | 1.5 | 20 | 21.5429 | 20.5979 |  | Cal1 1 2 | 47800000 | 1140000 | 0.5 | 20 | 20.9649 | 20.6758
1 | Cal2 1 2 | 1740000 | 167000 | 0.125 | 1 | 1.3024 | 0.96964 |  | Cal2 1 2 | 1910000 | 151000 | 0.05 | 1 | 0.63245 | 0.99536 |  | Cal2 1 2 | 26000000 | 3590000 | 1.5 | 10 | 10.8635 | 8.8041 |  | Cal2 1 2 | 19600000 | 929000 | 0.5 | 10 | 10.549 | 8.6439
1 | Cal3 1 2 | 181000 | 177000 | 0.125 | 0.1 | 0.12782 | 0.10281 |  | Cal3 1 2 | 197000 | 152000 | 0.05 | 0.1 | 0.064803 | 0.1017 |  | Cal3 1 2 | 2000000 | 2730000 | 1.5 | 1 | 1.0989 | 0.91962 |  | Cal3 1 2 | 1580000 | 762000 | 0.5 | 1 | 1.0367 | 0.95538
1 | Cal4 1 2 | 40300 | 180000 | 0.125 | 0.02 | 0.027986 | 0.024575 |  | Cal4 1 2 | 41600 | 163000 | 0.05 | 0.02 | 0.012761 | 0.020626 |  | Cal4 1 2 | 410000 | 2710000 | 1.5 | 0.2 | 0.22694 | 0.39728 |  | Cal4 1 2 | 313000 | 742000 | 0.5 | 0.2 | 0.21092 | 0.4148
1 | Cal5 1 2 | 21800 | 178000 | 0.125 | 0.01 | 0.015309 | 0.014289 |  | Cal5 1 2 | 22900 | 161000 | 0.05 | 0.01 | 0.0071118 | 0.010871 |  | Cal5 1 2 | 231000 | 2740000 | 1.5 | 0.1 | 0.12646 | 0.33848 |  | Cal5 1 2 | 168000 | 750000 | 0.5 | 0.1 | 0.112 | 0.35293
1 | Cal6 1 2 | 17200 | 182000 | 0.125 | 0.008 | 0.011813 | 0.011731 |  | Cal6 1 2 | 15900 | 160000 | 0.05 | 0.008 | 0.0049688 | 0.0072187 |  | Cal6 1 2 | 183000 | 2780000 | 1.5 | 0.08 | 0.098741 | 0.32271 |  | Cal6 1 2 | 131000 | 747000 | 0.5 | 0.08 | 0.087684 | 0.33715
2 | Cal1 2 2 | 3360000 | 184000 | 0.125 | 2 | 2.2826 | 1.9831 |  | Cal1 2 2 | 3660000 | 159000 | 0.05 | 2 | 1.1509 | 1.9883 |  | Cal1 2 2 | 56300000 | 4330000 | 1.5 | 20 | 19.5035 | 20.3469 |  | Cal1 2 2 | 43700000 | 1150000 | 0.5 | 20 | 19 | 20.4679
2 | Cal2 2 2 | 1760000 | 173000 | 0.125 | 1 | 1.2717 | 1.0337 |  | Cal2 2 2 | 1890000 | 148000 | 0.05 | 1 | 0.63851 | 1.0234 |  | Cal2 2 2 | 25500000 | 3490000 | 1.5 | 10 | 10.9599 | 9.3105 |  | Cal2 2 2 | 19100000 | 952000 | 0.5 | 10 | 10.0315 | 9.0635
2 | Cal3 2 2 | 194000 | 179000 | 0.125 | 0.1 | 0.13547 | 0.1045 |  | Cal3 2 2 | 201000 | 155000 | 0.05 | 0.1 | 0.064839 | 0.10272 |  | Cal3 2 2 | 2040000 | 2740000 | 1.5 | 1 | 1.1168 | 0.90418 |  | Cal3 2 2 | 1580000 | 739000 | 0.5 | 1 | 1.069 | 0.94137
2 | Cal4 2 2 | 41500 | 176000 | 0.125 | 0.02 | 0.029474 | 0.01401 |  | Cal4 2 2 | 40500 | 152000 | 0.05 | 0.02 | 0.013322 | 0.015232 |  | Cal4 2 2 | 421000 | 2760000 | 1.5 | 0.2 | 0.2288 | 0.32405 |  | Cal4 2 2 | 317000 | 747000 | 0.5 | 0.2 | 0.21218 | 0.35585
2 | Cal5 2 2 | 21100 | 183000 | 0.125 | 0.01 | 0.014413 | 0.0019057 |  | Cal5 2 2 | 21400 | 152000 | 0.05 | 0.01 | 0.0070395 | 0.0048202 |  | Cal5 2 2 | 223000 | 2710000 | 1.5 | 0.1 | 0.12343 | 0.2531 |  | Cal5 2 2 | 159000 | 753000 | 0.5 | 0.1 | 0.10558 | 0.28261
2 | Cal6 2 2 | 19300 | 182000 | 0.125 | 0.008 | 0.013255 | 0.0008377 |  | Cal6 2 2 | 19100 | 162000 | 0.05 | 0.008 | 0.0058951 | 0.0035665 |  | Cal6 2 2 | 190000 | 2880000 | 1.5 | 0.08 | 0.098958 | 0.24128 |  | Cal6 2 2 | 129000 | 754000 | 0.5 | 0.08 | 0.085544 | 0.2687
3 | Cal1 3 2 | 3560000 | 175000 | 0.125 | 2 | 2.5429 | 1.9922 |  | Cal1 3 2 | 3920000 | 157000 | 0.05 | 2 | 1.2484 | 1.9907 |  | Cal1 3 2 | 61500000 | 4170000 | 1.5 | 20 | 22.1223 | 20.3511 |  | Cal1 3 2 | 47500000 | 1150000 | 0.5 | 20 | 20.6522 | 20.4911
3 | Cal2 3 2 | 1820000 | 175000 | 0.125 | 1 | 1.3 | 1.0161 |  | Cal2 3 2 | 2010000 | 162000 | 0.05 | 1 | 0.62037 | 1.0191 |  | Cal2 3 2 | 27800000 | 3880000 | 1.5 | 10 | 10.7474 | 9.3049 |  | Cal2 3 2 | 20600000 | 1010000 | 0.5 | 10 | 10.198 | 9.0205
3 | Cal3 3 2 | 178000 | 175000 | 0.125 | 0.1 | 0.12714 | 0.094869 |  | Cal3 3 2 | 195000 | 154000 | 0.05 | 0.1 | 0.063312 | 0.095948 |  | Cal3 3 2 | 2070000 | 2750000 | 1.5 | 1 | 1.1291 | 0.8712 |  | Cal3 3 2 | 1560000 | 758000 | 0.5 | 1 | 1.029 | 0.90162
3 | Cal4 3 2 | 48300 | 179000 | 0.125 | 0.02 | 0.033729 | 0.022105 |  | Cal4 3 2 | 45100 | 163000 | 0.05 | 0.02 | 0.013834 | 0.019702 |  | Cal4 3 2 | 437000 | 2840000 | 1.5 | 0.2 | 0.23081 | 0.33593 |  | Cal4 3 2 | 320000 | 735000 | 0.5 | 0.2 | 0.21769 | 0.37286
3 | Cal5 3 2 | 22100 | 176000 | 0.125 | 0.01 | 0.015696 | 0.0074064 |  | Cal5 3 2 | 19600 | 153000 | 0.05 | 0.01 | 0.0064052 | 0.0067316 |  | Cal5 3 2 | 216000 | 2740000 | 1.5 | 0.1 | 0.11825 | 0.26349 |  | Cal5 3 2 | 159000 | 738000 | 0.5 | 0.1 | 0.10772 | 0.30421
3 | Cal6 3 2 | 18400 | 183000 | 0.125 | 0.008 | 0.012568 | 0.0053307 |  | Cal6 3 2 | 17800 | 157000 | 0.05 | 0.008 | 0.0056688 | 0.005816 |  | Cal6 3 2 | 185000 | 2780000 | 1.5 | 0.08 | 0.09982 | 0.25333 |  | Cal6 3 2 | 125000 | 754000 | 0.5 | 0.08 | 0.082891 | 0.28971
|  |  | 2.3206 | [7 empty] | 2.8998 | [7 empty] | 18.9126 | [7 empty] | 18.2004
|  | Cer d18:1/16:0 | D7 Cer d18:1/16:0 | D7 Cer d18:1/16:0 | Cer d18:1/16:0 |  |  |  |  | Cer d18:1/18:0 | D7 Cer d18:1/18:0 | D7 Cer d18:1/18:0 | Cer d18:1/18:0 |  |  |  |  | Cer d18:1/24:0 | D7 Cer d18:1/24:0 | D7 Cer d18:1/24:0 | Cer d18:1/24:0 |  |  |  |  | Cer d18:1/24:1 | D7 Cer d18:1/24:1 | D7 Cer d18:1/24:1 | Cer d18:1/24:1
|  | Analyte Peak Area (counts) | Analyte Peak Area (counts) | c [pmol/µL] | c [pmol/µL] |  |  |  |  | Analyte Peak Area (counts) | Analyte Peak Area (counts) | c [pmol/µL] | c [pmol/µL] |  |  |  |  | Analyte Peak Area (counts) | Analyte Peak Area (counts) | c [pmol/µL] | c [pmol/µL] |  |  |  |  | Analyte Peak Area (counts) | Analyte Peak Area (counts) | c [pmol/µL] | c [pmol/µL]
1 | SRM 1950-1 1 | 212000 | 157000 | 0.125 | 0.16879 |  |  |  | SRM 1950-1 1 | 95200 | 178000 | 0.05 | 0.026742 |  |  |  | SRM 1950-1 1 | 3240000 | 3620000 | 1.5 | 1.3425 |  |  |  | SRM 1950-1 1 | 866000 | 964000 | 0.5 | 0.44917
1 | SRM 1950-2 1 | 202000 | 159000 | 0.125 | 0.15881 |  |  |  | SRM 1950-2 1 | 89400 | 176000 | 0.05 | 0.025398 |  |  |  | SRM 1950-2 1 | 3060000 | 3520000 | 1.5 | 1.304 |  |  |  | SRM 1950-2 1 | 820000 | 929000 | 0.5 | 0.44133
1 | SRM 1950-3 1 | 201000 | 152000 | 0.125 | 0.1653 |  |  |  | SRM 1950-3 1 | 92100 | 178000 | 0.05 | 0.025871 |  |  |  | SRM 1950-3 1 | 3130000 | 3560000 | 1.5 | 1.3188 |  |  |  | SRM 1950-3 1 | 824000 | 958000 | 0.5 | 0.43006
1 | SRM 1950-4 1 | 204000 | 155000 | 0.125 | 0.16452 |  |  |  | SRM 1950-4 1 | 95300 | 181000 | 0.05 | 0.026326 |  |  |  | SRM 1950-4 1 | 3200000 | 3670000 | 1.5 | 1.3079 |  |  |  | SRM 1950-4 1 | 844000 | 961000 | 0.5 | 0.43913
1 | SRM 1950-5 1 | 210000 | 163000 | 0.125 | 0.16104 |  |  |  | SRM 1950-5 1 | 92300 | 187000 | 0.05 | 0.024679 |  |  |  | SRM 1950-5 1 | 3210000 | 3710000 | 1.5 | 1.2978 |  |  |  | SRM 1950-5 1 | 849000 | 997000 | 0.5 | 0.42578
1 | SRM 1950-6 1 | 206000 | 158000 | 0.125 | 0.16297 |  |  |  | SRM 1950-6 1 | 97200 | 182000 | 0.05 | 0.026703 |  |  |  | SRM 1950-6 1 | 3350000 | 3670000 | 1.5 | 1.3692 |  |  |  | SRM 1950-6 1 | 861000 | 959000 | 0.5 | 0.44891
2 | SRM 1950-1 2 | 213000 | 160000 | 0.125 | 0.16641 |  |  |  | SRM 1950-1 2 | 92700 | 183000 | 0.05 | 0.025328 |  |  |  | SRM 1950-1 2 | 3340000 | 3780000 | 1.5 | 1.3254 |  |  |  | SRM 1950-1 2 | 890000 | 1000000 | 0.5 | 0.445
2 | SRM 1950-2 2 | 202000 | 161000 | 0.125 | 0.15683 |  |  |  | SRM 1950-2 2 | 92600 | 177000 | 0.05 | 0.026158 |  |  |  | SRM 1950-2 2 | 3050000 | 3560000 | 1.5 | 1.2851 |  |  |  | SRM 1950-2 2 | 860000 | 936000 | 0.5 | 0.4594
2 | SRM 1950-3 2 | 197000 | 155000 | 0.125 | 0.15887 |  |  |  | SRM 1950-3 2 | 89800 | 171000 | 0.05 | 0.026257 |  |  |  | SRM 1950-3 2 | 2900000 | 3410000 | 1.5 | 1.2757 |  |  |  | SRM 1950-3 2 | 785000 | 918000 | 0.5 | 0.42756
2 | SRM 1950-4 2 | 198000 | 157000 | 0.125 | 0.15764 |  |  |  | SRM 1950-4 2 | 87300 | 168000 | 0.05 | 0.025982 |  |  |  | SRM 1950-4 2 | 3010000 | 3500000 | 1.5 | 1.29 |  |  |  | SRM 1950-4 2 | 814000 | 933000 | 0.5 | 0.43623
2 | SRM 1950-5 2 | 212000 | 162000 | 0.125 | 0.16358 |  |  |  | SRM 1950-5 2 | 92600 | 179000 | 0.05 | 0.025866 |  |  |  | SRM 1950-5 2 | 3290000 | 3700000 | 1.5 | 1.3338 |  |  |  | SRM 1950-5 2 | 875000 | 982000 | 0.5 | 0.44552
2 | SRM 1950-6 2 | 211000 | 157000 | 0.125 | 0.16799 |  |  |  | SRM 1950-6 2 | 93600 | 181000 | 0.05 | 0.025856 |  |  |  | SRM 1950-6 2 | 3300000 | 3500000 | 1.5 | 1.4143 |  |  |  | SRM 1950-6 2 | 853000 | 947000 | 0.5 | 0.45037
3 | SRM 1950-1 3 | 205000 | 162000 | 0.125 | 0.15818 |  |  |  | SRM 1950-1 3 | 91300 | 174000 | 0.05 | 0.026236 |  |  |  | SRM 1950-1 3 | 3260000 | 3740000 | 1.5 | 1.3075 |  |  |  | SRM 1950-1 3 | 880000 | 999000 | 0.5 | 0.44044
3 | SRM 1950-2 3 | 204000 | 157000 | 0.125 | 0.16242 |  |  |  | SRM 1950-2 3 | 91900 | 178000 | 0.05 | 0.025815 |  |  |  | SRM 1950-2 3 | 3220000 | 3670000 | 1.5 | 1.3161 |  |  |  | SRM 1950-2 3 | 856000 | 971000 | 0.5 | 0.44078
3 | SRM 1950-3 3 | 199000 | 156000 | 0.125 | 0.15946 |  |  |  | SRM 1950-3 3 | 87600 | 170000 | 0.05 | 0.025765 |  |  |  | SRM 1950-3 3 | 2950000 | 3460000 | 1.5 | 1.2789 |  |  |  | SRM 1950-3 3 | 812000 | 909000 | 0.5 | 0.44664
3 | SRM 1950-4 3 | 203000 | 157000 | 0.125 | 0.16162 |  |  |  | SRM 1950-4 3 | 93600 | 176000 | 0.05 | 0.026591 |  |  |  | SRM 1950-4 3 | 3070000 | 3580000 | 1.5 | 1.2863 |  |  |  | SRM 1950-4 3 | 794000 | 957000 | 0.5 | 0.41484
3 | SRM 1950-5 3 | 208000 | 165000 | 0.125 | 0.15758 |  |  |  | SRM 1950-5 3 | 96800 | 182000 | 0.05 | 0.026593 |  |  |  | SRM 1950-5 3 | 3250000 | 3710000 | 1.5 | 1.314 |  |  |  | SRM 1950-5 3 | 846000 | 993000 | 0.5 | 0.42598
3 | SRM 1950-6 3 | 212000 | 158000 | 0.125 | 0.16772 |  |  |  | SRM 1950-6 3 | 98600 | 181000 | 0.05 | 0.027238 |  |  |  | SRM 1950-6 3 | 3230000 | 3490000 | 1.5 | 1.3883 |  |  |  | SRM 1950-6 3 | 851000 | 949000 | 0.5 | 0.44837
| Average | 205500 | 158388.8889 |  | 0.16221 |  |  |  | Average | 92772.2222 | 177888.8889 |  | 0.026078 |  |  |  | Average | 3170000 | 3602777.7778 |  | 1.3198 |  |  |  | Average | 843333.3333 | 959000 |  | 0.43975
| SD | 5090.8414 | 3164.7167 |  | 0.0038124 |  |  |  | SD | 3018.61 | 4736.2537 |  | 0.00058336 |  |  |  | SD | 130298.8019 | 105135.1276 |  | 0.037048 |  |  |  | SD | 28458.5469 | 26995.8845 |  | 0.01085
| RSD | 2.4773 | 1.9981 |  | 2.3503 |  |  |  | RSD | 3.2538 | 2.6625 |  | 2.237 |  |  |  | RSD | 4.1104 | 2.9182 |  | 2.8072 |  |  |  | RSD | 3.3745 | 2.815 |  | 2.4672
T1D-1? |  | Cer d18:1/16:0 | D7 Cer d18:1/16:0 | D7 Cer d18:1/16:0 | Cer d18:1/16:0 |  |  |  |  | Cer d18:1/18:0 | D7 Cer d18:1/18:0 | D7 Cer d18:1/18:0 | Cer d18:1/18:0 |  |  |  |  | Cer d18:1/24:0 | D7 Cer d18:1/24:0 | D7 Cer d18:1/24:0 | Cer d18:1/24:0 |  |  |  |  | Cer d18:1/24:1 | D7 Cer d18:1/24:1 | D7 Cer d18:1/24:1 | Cer d18:1/24:1
|  | Analyte Peak Area (counts) | Analyte Peak Area (counts) | c [pmol/µL] | c [pmol/µL] |  |  |  |  | Analyte Peak Area (counts) | Analyte Peak Area (counts) | c [pmol/µL] | c [pmol/µL] |  |  |  |  | Analyte Peak Area (counts) | Analyte Peak Area (counts) | c [pmol/µL] | c [pmol/µL] |  |  |  |  | Analyte Peak Area (counts) | Analyte Peak Area (counts) | c [pmol/µL] | c [pmol/µL]
1 | T1D-1 1 | 265000 | 172000 | 0.125 | 0.19259 |  |  |  | T1D-1 1 | 136000 | 196000 | 0.05 | 0.034694 |  |  |  | T1D-1 1 | 3600000 | 3960000 | 1.5 | 1.3636 |  |  |  | T1D-1 1 | 980000 | 1050000 | 0.5 | 0.46667
1 | T1D-2 1 | 186000 | 172000 | 0.125 | 0.13517 |  |  |  | T1D-2 1 | 108000 | 185000 | 0.05 | 0.029189 |  |  |  | T1D-2 1 | 3210000 | 3870000 | 1.5 | 1.2442 |  |  |  | T1D-2 1 | 850000 | 1010000 | 0.5 | 0.42079
1 | T1D-3 1 | 187000 | 165000 | 0.125 | 0.14167 |  |  |  | T1D-3 1 | 106000 | 181000 | 0.05 | 0.029282 |  |  |  | T1D-3 1 | 3180000 | 3700000 | 1.5 | 1.2892 |  |  |  | T1D-3 1 | 828000 | 966000 | 0.5 | 0.42857
1 | T1D-4 1 | 189000 | 172000 | 0.125 | 0.13735 |  |  |  | T1D-4 1 | 112000 | 185000 | 0.05 | 0.03027 |  |  |  | T1D-4 1 | 3130000 | 3720000 | 1.5 | 1.2621 |  |  |  | T1D-4 1 | 818000 | 973000 | 0.5 | 0.42035
1 | T1D-5 1 | 181000 | 165000 | 0.125 | 0.13712 |  |  |  | T1D-5 1 | 106000 | 186000 | 0.05 | 0.028495 |  |  |  | T1D-5 1 | 2990000 | 3600000 | 1.5 | 1.2458 |  |  |  | T1D-5 1 | 829000 | 995000 | 0.5 | 0.41658
1 | T1D-6 1 | 184000 | 170000 | 0.125 | 0.13529 |  |  |  | T1D-6 1 | 109000 | 190000 | 0.05 | 0.028684 |  |  |  | T1D-6 1 | 3010000 | 3700000 | 1.5 | 1.2203 |  |  |  | T1D-6 1 | 798000 | 959000 | 0.5 | 0.41606
2 | T1D-1 2 | 874000 | 169000 | 0.125 | 0.64645 |  |  |  | T1D-1 2 | 312000 | 202000 | 0.05 | 0.077228 |  |  |  | T1D-1 2 | 4180000 | 4430000 | 1.5 | 1.4153 |  |  |  | T1D-1 2 | 1180000 | 1170000 | 0.5 | 0.50427
2 | T1D-2 2 | 186000 | 166000 | 0.125 | 0.14006 |  |  |  | T1D-2 2 | 113000 | 188000 | 0.05 | 0.030053 |  |  |  | T1D-2 2 | 3190000 | 3720000 | 1.5 | 1.2863 |  |  |  | T1D-2 2 | 849000 | 1020000 | 0.5 | 0.41618
2 | T1D-3 2 | 183000 | 161000 | 0.125 | 0.14208 |  |  |  | T1D-3 2 | 110000 | 179000 | 0.05 | 0.030726 |  |  |  | T1D-3 2 | 3190000 | 3660000 | 1.5 | 1.3074 |  |  |  | T1D-3 2 | 851000 | 959000 | 0.5 | 0.44369
2 | T1D-4 2 | 181000 | 170000 | 0.125 | 0.13309 |  |  |  | T1D-4 2 | 108000 | 189000 | 0.05 | 0.028571 |  |  |  | T1D-4 2 | 3160000 | 3780000 | 1.5 | 1.254 |  |  |  | T1D-4 2 | 836000 | 955000 | 0.5 | 0.4377
2 | T1D-5 2 | 179000 | 164000 | 0.125 | 0.13643 |  |  |  | T1D-5 2 | 103000 | 180000 | 0.05 | 0.028611 |  |  |  | T1D-5 2 | 3040000 | 3620000 | 1.5 | 1.2597 |  |  |  | T1D-5 2 | 818000 | 972000 | 0.5 | 0.42078
2 | T1D-6 2 | 183000 | 168000 | 0.125 | 0.13616 |  |  |  | T1D-6 2 | 110000 | 186000 | 0.05 | 0.02957 |  |  |  | T1D-6 2 | 3080000 | 3630000 | 1.5 | 1.2727 |  |  |  | T1D-6 2 | 820000 | 973000 | 0.5 | 0.42138
3 | T1D-1 3 | 796000 | 167000 | 0.125 | 0.59581 |  |  |  | T1D-1 3 | 298000 | 198000 | 0.05 | 0.075253 |  |  |  | T1D-1 3 | 4190000 | 4280000 | 1.5 | 1.4685 |  |  |  | T1D-1 3 | 1130000 | 1160000 | 0.5 | 0.48707
3 | T1D-2 3 | 189000 | 167000 | 0.125 | 0.14147 |  |  |  | T1D-2 3 | 107000 | 184000 | 0.05 | 0.029076 |  |  |  | T1D-2 3 | 3080000 | 3780000 | 1.5 | 1.2222 |  |  |  | T1D-2 3 | 830000 | 992000 | 0.5 | 0.41835
3 | T1D-3 3 | 187000 | 163000 | 0.125 | 0.1434 |  |  |  | T1D-3 3 | 106000 | 184000 | 0.05 | 0.028804 |  |  |  | T1D-3 3 | 3150000 | 3640000 | 1.5 | 1.2981 |  |  |  | T1D-3 3 | 843000 | 987000 | 0.5 | 0.42705
3 | T1D-4 3 | 186000 | 172000 | 0.125 | 0.13517 |  |  |  | T1D-4 3 | 106000 | 186000 | 0.05 | 0.028495 |  |  |  | T1D-4 3 | 3190000 | 3780000 | 1.5 | 1.2659 |  |  |  | T1D-4 3 | 826000 | 991000 | 0.5 | 0.41675
3 | T1D-5 3 | 178000 | 166000 | 0.125 | 0.13404 |  |  |  | T1D-5 3 | 106000 | 183000 | 0.05 | 0.028962 |  |  |  | T1D-5 3 | 3060000 | 3590000 | 1.5 | 1.2786 |  |  |  | T1D-5 3 | 817000 | 964000 | 0.5 | 0.42376
3 | T1D-6 3 | 178000 | 165000 | 0.125 | 0.13485 |  |  |  | T1D-6 3 | 107000 | 186000 | 0.05 | 0.028763 |  |  |  | T1D-6 3 | 3010000 | 3700000 | 1.5 | 1.2203 |  |  |  | T1D-6 3 | 814000 | 956000 | 0.5 | 0.42573
| Average | 260666.6667 | 167444.4444 |  | 0.19435 |  |  |  | Average | 131277.7778 | 187111.1111 |  | 0.034707 |  |  |  | Average | 3257777.7778 | 3786666.6667 |  | 1.2874 |  |  |  | Average | 873166.6667 | 1002888.8889 |  | 0.43398
| SD | 204342.0335 | 3286.711 |  | 0.15168 |  |  |  | SD | 61843.7148 | 5943.1463 |  | 0.014755 |  |  |  | SD | 352807.0854 | 222261.1077 |  | 0.064817 |  |  |  | SD | 106647.0034 | 62229.7614 |  | 0.025123
| RSD | 78.3921 | 1.9629 |  | 78.0446 |  |  |  | RSD | 47.1091 | 3.1763 |  | 42.5118 |  |  |  | RSD | 10.8297 | 5.8696 |  | 5.0345 |  |  |  | RSD | 12.2138 | 6.2051 |  | 5.7889
| Extraction | Cer d18:1/16:0 | D7 Cer d18:1/16:0 | D7 Cer d18:1/16:0 | Cer d18:1/16:0 |  |  |  | Extraction | Cer d18:1/18:0 | D7 Cer d18:1/18:0 | D7 Cer d18:1/18:0 | Cer d18:1/18:0 |  |  |  | Extraction | Cer d18:1/24:0 | D7 Cer d18:1/24:0 | D7 Cer d18:1/24:0 | Cer d18:1/24:0 |  |  |  | Extraction | Cer d18:1/24:1 | D7 Cer d18:1/24:1 | D7 Cer d18:1/24:1 | Cer d18:1/24:1
| REPEATED | Analyte Peak Area (counts) | Analyte Peak Area (counts) | c [pmol/µL] | c [pmol/µL] |  |  |  | REPEATED | Analyte Peak Area (counts) | Analyte Peak Area (counts) | c [pmol/µL] | c [pmol/µL] |  |  |  | REPEATED | Analyte Peak Area (counts) | Analyte Peak Area (counts) | c [pmol/µL] | c [pmol/µL] |  |  |  | REPEATED | Analyte Peak Area (counts) | Analyte Peak Area (counts) | c [pmol/µL] | c [pmol/µL]
1 | T1D-1 1 2 | 215000 | 192000 | 0.125 | 0.13997 |  |  |  | T1D-1 1 | 141000 | 245000 | 0.05 | 0.028776 |  |  |  | T1D-1 1 | 3200000 | 3620000 | 1.5 | 1.326 |  |  |  | T1D-1 1 | 854000 | 945000 | 0.5 | 0.45185
1 | T1D-2 1 2 | 209000 | 191000 | 0.125 | 0.13678 |  |  |  | T1D-2 1 | 144000 | 240000 | 0.05 | 0.03 |  |  |  | T1D-2 1 | 2980000 | 3510000 | 1.5 | 1.2735 |  |  |  | T1D-2 1 | 812000 | 924000 | 0.5 | 0.43939
1 | T1D-3 1 2 | 208000 | 196000 | 0.125 | 0.13265 |  |  |  | T1D-3 1 | 138000 | 244000 | 0.05 | 0.028279 |  |  |  | T1D-3 1 | 2930000 | 3410000 | 1.5 | 1.2889 |  |  |  | T1D-3 1 | 773000 | 928000 | 0.5 | 0.41649
1 | T1D-4 1 2 | 195000 | 195000 | 0.125 | 0.125 |  |  |  | T1D-4 1 | 137000 | 242000 | 0.05 | 0.028306 |  |  |  | T1D-4 1 | 2830000 | 3440000 | 1.5 | 1.234 |  |  |  | T1D-4 1 | 743000 | 920000 | 0.5 | 0.4038
1 | T1D-5 1 2 | 216000 | 192000 | 0.125 | 0.14062 |  |  |  | T1D-5 1 | 144000 | 237000 | 0.05 | 0.03038 |  |  |  | T1D-5 1 | 2910000 | 3490000 | 1.5 | 1.2507 |  |  |  | T1D-5 1 | 789000 | 910000 | 0.5 | 0.43352
1 | T1D-6 1 2 | 205000 | 194000 | 0.125 | 0.13209 |  |  |  | T1D-6 1 | 144000 | 247000 | 0.05 | 0.02915 |  |  |  | T1D-6 1 | 2820000 | 3500000 | 1.5 | 1.2086 |  |  |  | T1D-6 1 | 772000 | 928000 | 0.5 | 0.41595
2 | T1D-1 2 2 | 212000 | 196000 | 0.125 | 0.1352 |  |  |  | T1D-1 2 | 147000 | 250000 | 0.05 | 0.0294 |  |  |  | T1D-1 2 | 3040000 | 3590000 | 1.5 | 1.2702 |  |  |  | T1D-1 2 | 826000 | 895000 | 0.5 | 0.46145
2 | T1D-2 2 2 | 207000 | 187000 | 0.125 | 0.13837 |  |  |  | T1D-2 2 | 139000 | 237000 | 0.05 | 0.029325 |  |  |  | T1D-2 2 | 2940000 | 3430000 | 1.5 | 1.2857 |  |  |  | T1D-2 2 | 788000 | 895000 | 0.5 | 0.44022
2 | T1D-3 2 2 | 208000 | 200000 | 0.125 | 0.13 |  |  |  | T1D-3 2 | 134000 | 241000 | 0.05 | 0.027801 |  |  |  | T1D-3 2 | 2920000 | 3590000 | 1.5 | 1.2201 |  |  |  | T1D-3 2 | 755000 | 936000 | 0.5 | 0.40331
2 | T1D-4 2 2 | 208000 | 195000 | 0.125 | 0.13333 |  |  |  | T1D-4 2 | 139000 | 251000 | 0.05 | 0.027689 |  |  |  | T1D-4 2 | 2850000 | 3440000 | 1.5 | 1.2427 |  |  |  | T1D-4 2 | 751000 | 936000 | 0.5 | 0.40118
2 | T1D-5 2 2 | 208000 | 193000 | 0.125 | 0.13472 |  |  |  | T1D-5 2 | 138000 | 238000 | 0.05 | 0.028992 |  |  |  | T1D-5 2 | 2910000 | 3360000 | 1.5 | 1.2991 |  |  |  | T1D-5 2 | 770000 | 887000 | 0.5 | 0.43405
2 | T1D-6 2 2 | 209000 | 198000 | 0.125 | 0.13194 |  |  |  | T1D-6 2 | 142000 | 244000 | 0.05 | 0.029098 |  |  |  | T1D-6 2 | 2940000 | 3470000 | 1.5 | 1.2709 |  |  |  | T1D-6 2 | 775000 | 907000 | 0.5 | 0.42723
3 | T1D-1 3 2 | 217000 | 193000 | 0.125 | 0.14054 |  |  |  | T1D-1 3 | 145000 | 241000 | 0.05 | 0.030083 |  |  |  | T1D-1 3 | 3170000 | 3620000 | 1.5 | 1.3135 |  |  |  | T1D-1 3 | 838000 | 915000 | 0.5 | 0.45792
3 | T1D-2 3 2 | 212000 | 197000 | 0.125 | 0.13452 |  |  |  | T1D-2 3 | 138000 | 244000 | 0.05 | 0.028279 |  |  |  | T1D-2 3 | 2950000 | 3440000 | 1.5 | 1.2863 |  |  |  | T1D-2 3 | 807000 | 926000 | 0.5 | 0.43575
3 | T1D-3 3 2 | 209000 | 196000 | 0.125 | 0.13329 |  |  |  | T1D-3 3 | 139000 | 243000 | 0.05 | 0.028601 |  |  |  | T1D-3 3 | 2860000 | 3540000 | 1.5 | 1.2119 |  |  |  | T1D-3 3 | 744000 | 935000 | 0.5 | 0.39786
3 | T1D-4 3 2 | 212000 | 199000 | 0.125 | 0.13317 |  |  |  | T1D-4 3 | 135000 | 251000 | 0.05 | 0.026892 |  |  |  | T1D-4 3 | 2890000 | 3460000 | 1.5 | 1.2529 |  |  |  | T1D-4 3 | 785000 | 919000 | 0.5 | 0.42709
3 | T1D-5 3 2 | 210000 | 187000 | 0.125 | 0.14037 |  |  |  | T1D-5 3 | 143000 | 241000 | 0.05 | 0.029668 |  |  |  | T1D-5 3 | 2950000 | 3490000 | 1.5 | 1.2679 |  |  |  | T1D-5 3 | 795000 | 888000 | 0.5 | 0.44764
3 | T1D-6 3 2 | 207000 | 193000 | 0.125 | 0.13407 |  |  |  | T1D-6 3 | 138000 | 251000 | 0.05 | 0.02749 |  |  |  | T1D-6 3 | 2980000 | 3490000 | 1.5 | 1.2808 |  |  |  | T1D-6 3 | 794000 | 927000 | 0.5 | 0.42826
| Average | 209277.7778 | 194111.1111 |  | 0.13481 |  |  |  | Average | 140277.7778 | 243722.2222 |  | 0.028789 |  |  |  | Average | 2948333.3333 | 3493888.8889 |  | 1.2658 |  |  |  | Average | 787277.7778 | 917833.3333 |  | 0.42905
| SD | 4711.7531 | 3478.328 |  | 0.0040035 |  |  |  | SD | 3524.6049 | 4555.8943 |  | 0.00093189 |  |  |  | SD | 99568.5135 | 71580.4822 |  | 0.032409 |  |  |  | SD | 30473.6782 | 16938.6147 |  | 0.01894
| RSD | 2.2514 | 1.7919 |  | 2.9697 |  |  |  | RSD | 2.5126 | 1.8693 |  | 3.2369 |  |  |  | RSD | 3.3771 | 2.0487 |  | 2.5604 |  |  |  | RSD | 3.8708 | 1.8455 |  | 4.4145
|  | Cer d18:1/16:0 | D7 Cer d18:1/16:0 | D7 Cer d18:1/16:0 | Cer d18:1/16:0 |  |  |  |  | Cer d18:1/18:0 | D7 Cer d18:1/18:0 | D7 Cer d18:1/18:0 | Cer d18:1/18:0 |  |  |  |  | Cer d18:1/24:0 | D7 Cer d18:1/24:0 | D7 Cer d18:1/24:0 | Cer d18:1/24:0 |  |  |  |  | Cer d18:1/24:1 | D7 Cer d18:1/24:1 | D7 Cer d18:1/24:1 | Cer d18:1/24:1
|  | Analyte Peak Area (counts) | Analyte Peak Area (counts) | c [pmol/µL] | c [pmol/µL] |  |  |  |  | Analyte Peak Area (counts) | Analyte Peak Area (counts) | c [pmol/µL] | c [pmol/µL] |  |  |  |  | Analyte Peak Area (counts) | Analyte Peak Area (counts) | c [pmol/µL] | c [pmol/µL] |  |  |  |  | Analyte Peak Area (counts) | Analyte Peak Area (counts) | c [pmol/µL] | c [pmol/µL]
1 | Young AA-1 1 | 142000 | 183000 | 0.125 | 0.096995 |  |  |  | Young AA-1 1 | 56000 | 186000 | 0.05 | 0.015054 |  |  |  | Young AA-1 1 | 1960000 | 3430000 | 1.5 | 0.85714 |  |  |  | Young AA-1 1 | 567000 | 1050000 | 0.5 | 0.27
1 | Young AA-2 1 | 140000 | 178000 | 0.125 | 0.098315 |  |  |  | Young AA-2 1 | 54100 | 185000 | 0.05 | 0.014622 |  |  |  | Young AA-2 1 | 1970000 | 3280000 | 1.5 | 0.90091 |  |  |  | Young AA-2 1 | 558000 | 1030000 | 0.5 | 0.27087
1 | Young AA-3 1 | 142000 | 187000 | 0.125 | 0.09492 |  |  |  | Young AA-3 1 | 54000 | 189000 | 0.05 | 0.014286 |  |  |  | Young AA-3 1 | 1930000 | 3400000 | 1.5 | 0.85147 |  |  |  | Young AA-3 1 | 549000 | 1030000 | 0.5 | 0.2665
1 | Young AA-4 1 | 144000 | 181000 | 0.125 | 0.099448 |  |  |  | Young AA-4 1 | 54300 | 189000 | 0.05 | 0.014365 |  |  |  | Young AA-4 1 | 1980000 | 3360000 | 1.5 | 0.88393 |  |  |  | Young AA-4 1 | 567000 | 1020000 | 0.5 | 0.27794
1 | Young AA-5 1 | 140000 | 185000 | 0.125 | 0.094595 |  |  |  | Young AA-5 1 | 54000 | 193000 | 0.05 | 0.01399 |  |  |  | Young AA-5 1 | 1950000 | 3360000 | 1.5 | 0.87054 |  |  |  | Young AA-5 1 | 549000 | 1040000 | 0.5 | 0.26394
1 | Young AA-6 1 | 144000 | 186000 | 0.125 | 0.096774 |  |  |  | Young AA-6 1 | 58500 | 198000 | 0.05 | 0.014773 |  |  |  | Young AA-6 1 | 1970000 | 3390000 | 1.5 | 0.87168 |  |  |  | Young AA-6 1 | 544000 | 1040000 | 0.5 | 0.26154
2 | Young AA-1 2 | 136000 | 180000 | 0.125 | 0.094444 |  |  |  | Young AA-1 2 | 56100 | 184000 | 0.05 | 0.015245 |  |  |  | Young AA-1 2 | 1950000 | 3360000 | 1.5 | 0.87054 |  |  |  | Young AA-1 2 | 555000 | 1000000 | 0.5 | 0.2775
2 | Young AA-2 2 | 144000 | 181000 | 0.125 | 0.099448 |  |  |  | Young AA-2 2 | 56000 | 189000 | 0.05 | 0.014815 |  |  |  | Young AA-2 2 | 1950000 | 3390000 | 1.5 | 0.86283 |  |  |  | Young AA-2 2 | 554000 | 1010000 | 0.5 | 0.27426
2 | Young AA-3 2 | 135000 | 177000 | 0.125 | 0.095339 |  |  |  | Young AA-3 2 | 54200 | 196000 | 0.05 | 0.013827 |  |  |  | Young AA-3 2 | 1910000 | 3390000 | 1.5 | 0.84513 |  |  |  | Young AA-3 2 | 543000 | 1090000 | 0.5 | 0.24908
2 | Young AA-4 2 | 135000 | 177000 | 0.125 | 0.095339 |  |  |  | Young AA-4 2 | 54800 | 187000 | 0.05 | 0.014652 |  |  |  | Young AA-4 2 | 1920000 | 3280000 | 1.5 | 0.87805 |  |  |  | Young AA-4 2 | 546000 | 998000 | 0.5 | 0.27355
2 | Young AA-5 2 | 135000 | 181000 | 0.125 | 0.093232 |  |  |  | Young AA-5 2 | 56600 | 183000 | 0.05 | 0.015464 |  |  |  | Young AA-5 2 | 1960000 | 3300000 | 1.5 | 0.89091 |  |  |  | Young AA-5 2 | 546000 | 1010000 | 0.5 | 0.2703
2 | Young AA-6 2 | 149000 | 184000 | 0.125 | 0.10122 |  |  |  | Young AA-6 2 | 57700 | 193000 | 0.05 | 0.014948 |  |  |  | Young AA-6 2 | 1940000 | 3420000 | 1.5 | 0.85088 |  |  |  | Young AA-6 2 | 551000 | 1050000 | 0.5 | 0.26238
3 | Young AA-1 3 | 139000 | 179000 | 0.125 | 0.097067 |  |  |  | Young AA-1 3 | 57400 | 187000 | 0.05 | 0.015348 |  |  |  | Young AA-1 3 | 1960000 | 3320000 | 1.5 | 0.88554 |  |  |  | Young AA-1 3 | 534000 | 1000000 | 0.5 | 0.267
3 | Young AA-2 3 | 139000 | 178000 | 0.125 | 0.097612 |  |  |  | Young AA-2 3 | 55300 | 180000 | 0.05 | 0.015361 |  |  |  | Young AA-2 3 | 1960000 | 3300000 | 1.5 | 0.89091 |  |  |  | Young AA-2 3 | 550000 | 985000 | 0.5 | 0.27919
3 | Young AA-3 3 | 142000 | 186000 | 0.125 | 0.09543 |  |  |  | Young AA-3 3 | 54600 | 197000 | 0.05 | 0.013858 |  |  |  | Young AA-3 3 | 1960000 | 3440000 | 1.5 | 0.85465 |  |  |  | Young AA-3 3 | 561000 | 1050000 | 0.5 | 0.26714
3 | Young AA-4 3 | 143000 | 177000 | 0.125 | 0.10099 |  |  |  | Young AA-4 3 | 55400 | 183000 | 0.05 | 0.015137 |  |  |  | Young AA-4 3 | 1980000 | 3310000 | 1.5 | 0.89728 |  |  |  | Young AA-4 3 | 571000 | 1010000 | 0.5 | 0.28267
3 | Young AA-5 3 | 144000 | 182000 | 0.125 | 0.098901 |  |  |  | Young AA-5 3 | 56400 | 191000 | 0.05 | 0.014764 |  |  |  | Young AA-5 3 | 1950000 | 3340000 | 1.5 | 0.87575 |  |  |  | Young AA-5 3 | 573000 | 1030000 | 0.5 | 0.27816
3 | Young AA-6 3 | 144000 | 186000 | 0.125 | 0.096774 |  |  |  | Young AA-6 3 | 57400 | 195000 | 0.05 | 0.014718 |  |  |  | Young AA-6 3 | 1860000 | 3310000 | 1.5 | 0.8429 |  |  |  | Young AA-6 3 | 531000 | 1030000 | 0.5 | 0.25777
| Average | 140944.4444 | 181555.5556 |  | 0.097047 |  |  |  | Average | 55711.1111 | 189166.6667 |  | 0.014735 |  |  |  | Average | 1947777.7778 | 3354444.4444 |  | 0.87117 |  |  |  | Average | 552722.2222 | 1026277.7778 |  | 0.26943
| SD | 3792.862 | 3386.6113 |  | 0.0022402 |  |  |  | SD | 1375.9396 | 5145.1163 |  | 0.00049357 |  |  |  | SD | 27999.1182 | 50356.752 |  | 0.017639 |  |  |  | SD | 11493.0199 | 24607.5994 |  | 0.0082984
| RSD | 2.691 | 1.8653 |  | 2.3083 |  |  |  | RSD | 2.4698 | 2.7199 |  | 3.3497 |  |  |  | RSD | 1.4375 | 1.5012 |  | 2.0248 |  |  |  | RSD | 2.0793 | 2.3978 |  | 3.0799
|  | Cer d18:1/16:0 | D7 Cer d18:1/16:0 | D7 Cer d18:1/16:0 | Cer d18:1/16:0 |  |  |  |  | Cer d18:1/18:0 | D7 Cer d18:1/18:0 | D7 Cer d18:1/18:0 | Cer d18:1/18:0 |  |  |  |  | Cer d18:1/24:0 | D7 Cer d18:1/24:0 | D7 Cer d18:1/24:0 | Cer d18:1/24:0 |  |  |  |  | Cer d18:1/24:1 | D7 Cer d18:1/24:1 | D7 Cer d18:1/24:1 | Cer d18:1/24:1
|  | Analyte Peak Area (counts) | Analyte Peak Area (counts) | c [pmol/µL] | c [pmol/µL] |  |  |  |  | Analyte Peak Area (counts) | Analyte Peak Area (counts) | c [pmol/µL] | c [pmol/µL] |  |  |  |  | Analyte Peak Area (counts) | Analyte Peak Area (counts) | c [pmol/µL] | c [pmol/µL] |  |  |  |  | Analyte Peak Area (counts) | Analyte Peak Area (counts) | c [pmol/µL] | c [pmol/µL]
1 | hTAG-1 1 | 230000 | 164000 | 0.125 | 0.1753 |  |  |  | hTAG-1 1 | 107000 | 192000 | 0.05 | 0.027865 |  |  |  | hTAG-1 1 | 5180000 | 4120000 | 1.5 | 1.8859 |  |  |  | hTAG-1 1 | 1180000 | 1030000 | 0.5 | 0.57282
1 | hTAG-2 1 | 231000 | 157000 | 0.125 | 0.18392 |  |  |  | hTAG-2 1 | 118000 | 197000 | 0.05 | 0.029949 |  |  |  | hTAG-2 1 | 5370000 | 4160000 | 1.5 | 1.9363 |  |  |  | hTAG-2 1 | 1240000 | 1020000 | 0.5 | 0.60784
1 | hTAG-3 1 | 237000 | 164000 | 0.125 | 0.18064 |  |  |  | hTAG-3 1 | 119000 | 199000 | 0.05 | 0.029899 |  |  |  | hTAG-3 1 | 5350000 | 4120000 | 1.5 | 1.9478 |  |  |  | hTAG-3 1 | 1230000 | 1040000 | 0.5 | 0.59135
1 | hTAG-4 1 | 234000 | 162000 | 0.125 | 0.18056 |  |  |  | hTAG-4 1 | 115000 | 192000 | 0.05 | 0.029948 |  |  |  | hTAG-4 1 | 5330000 | 4030000 | 1.5 | 1.9839 |  |  |  | hTAG-4 1 | 1190000 | 1000000 | 0.5 | 0.595
1 | hTAG-5 1 | 235000 | 166000 | 0.125 | 0.17696 |  |  |  | hTAG-5 1 | 123000 | 200000 | 0.05 | 0.03075 |  |  |  | hTAG-5 1 | 5400000 | 4280000 | 1.5 | 1.8925 |  |  |  | hTAG-5 1 | 1280000 | 1050000 | 0.5 | 0.60952
1 | hTAG-6 1 | 229000 | 160000 | 0.125 | 0.17891 |  |  |  | hTAG-6 1 | 115000 | 197000 | 0.05 | 0.029188 |  |  |  | hTAG-6 1 | 5160000 | 4110000 | 1.5 | 1.8832 |  |  |  | hTAG-6 1 | 1190000 | 984000 | 0.5 | 0.60467
2 | hTAG-1 2 | 228000 | 161000 | 0.125 | 0.17702 |  |  |  | hTAG-1 2 | 111000 | 198000 | 0.05 | 0.02803 |  |  |  | hTAG-1 2 | 5190000 | 4150000 | 1.5 | 1.8759 |  |  |  | hTAG-1 2 | 1180000 | 1010000 | 0.5 | 0.58416
2 | hTAG-2 2 | 232000 | 159000 | 0.125 | 0.18239 |  |  |  | hTAG-2 2 | 118000 | 190000 | 0.05 | 0.031053 |  |  |  | hTAG-2 2 | 5530000 | 4210000 | 1.5 | 1.9703 |  |  |  | hTAG-2 2 | 1220000 | 998000 | 0.5 | 0.61122
2 | hTAG-3 2 | 233000 | 164000 | 0.125 | 0.17759 |  |  |  | hTAG-3 2 | 118000 | 193000 | 0.05 | 0.03057 |  |  |  | hTAG-3 2 | 5400000 | 4080000 | 1.5 | 1.9853 |  |  |  | hTAG-3 2 | 1220000 | 1010000 | 0.5 | 0.60396
2 | hTAG-4 2 | 238000 | 167000 | 0.125 | 0.17814 |  |  |  | hTAG-4 2 | 123000 | 201000 | 0.05 | 0.030597 |  |  |  | hTAG-4 2 | 5430000 | 4130000 | 1.5 | 1.9722 |  |  |  | hTAG-4 2 | 1230000 | 1020000 | 0.5 | 0.60294
2 | hTAG-5 2 | 232000 | 165000 | 0.125 | 0.17576 |  |  |  | hTAG-5 2 | 120000 | 203000 | 0.05 | 0.029557 |  |  |  | hTAG-5 2 | 5230000 | 4160000 | 1.5 | 1.8858 |  |  |  | hTAG-5 2 | 1230000 | 1030000 | 0.5 | 0.59709
2 | hTAG-6 2 | 236000 | 166000 | 0.125 | 0.17771 |  |  |  | hTAG-6 2 | 120000 | 202000 | 0.05 | 0.029703 |  |  |  | hTAG-6 2 | 5320000 | 4230000 | 1.5 | 1.8865 |  |  |  | hTAG-6 2 | 1230000 | 1010000 | 0.5 | 0.60891
3 | hTAG-1 3 | 231000 | 167000 | 0.125 | 0.1729 |  |  |  | hTAG-1 3 | 110000 | 197000 | 0.05 | 0.027919 |  |  |  | hTAG-1 3 | 5330000 | 4290000 | 1.5 | 1.8636 |  |  |  | hTAG-1 3 | 1210000 | 1020000 | 0.5 | 0.59314
3 | hTAG-2 3 | 232000 | 163000 | 0.125 | 0.17791 |  |  |  | hTAG-2 3 | 117000 | 190000 | 0.05 | 0.030789 |  |  |  | hTAG-2 3 | 5430000 | 4150000 | 1.5 | 1.9627 |  |  |  | hTAG-2 3 | 1240000 | 1030000 | 0.5 | 0.60194
3 | hTAG-3 3 | 235000 | 160000 | 0.125 | 0.18359 |  |  |  | hTAG-3 3 | 116000 | 200000 | 0.05 | 0.029 |  |  |  | hTAG-3 3 | 5510000 | 4140000 | 1.5 | 1.9964 |  |  |  | hTAG-3 3 | 1230000 | 976000 | 0.5 | 0.63012
3 | hTAG-4 3 | 234000 | 161000 | 0.125 | 0.18168 |  |  |  | hTAG-4 3 | 116000 | 205000 | 0.05 | 0.028293 |  |  |  | hTAG-4 3 | 5380000 | 4160000 | 1.5 | 1.9399 |  |  |  | hTAG-4 3 | 1230000 | 1000000 | 0.5 | 0.615
3 | hTAG-5 3 | 238000 | 165000 | 0.125 | 0.1803 |  |  |  | hTAG-5 3 | 111000 | 199000 | 0.05 | 0.027889 |  |  |  | hTAG-5 3 | 5310000 | 4190000 | 1.5 | 1.901 |  |  |  | hTAG-5 3 | 1190000 | 998000 | 0.5 | 0.59619
3 | hTAG-6 3 | 226000 | 164000 | 0.125 | 0.17226 |  |  |  | hTAG-6 3 | 117000 | 202000 | 0.05 | 0.02896 |  |  |  | hTAG-6 3 | 5160000 | 4100000 | 1.5 | 1.8878 |  |  |  | hTAG-6 3 | 1200000 | 1010000 | 0.5 | 0.59406
| Average | 232833.3333 | 163055.5556 |  | 0.17853 |  |  |  | Average | 116333.3333 | 197611.1111 |  | 0.029442 |  |  |  | Average | 5333888.8889 | 4156111.1111 |  | 1.9254 |  |  |  | Average | 1217777.7778 | 1013111.1111 |  | 0.60111
| SD | 3287.1805 | 2798.2578 |  | 0.0032253 |  |  |  | SD | 4216.3702 | 4411.335 |  | 0.0010686 |  |  |  | SD | 109754.2147 | 63869.5623 |  | 0.043614 |  |  |  | SD | 25067.8093 | 18477.8813 |  | 0.012277
| RSD | 1.4118 | 1.7161 |  | 1.8066 |  |  |  | RSD | 3.6244 | 2.2323 |  | 3.6296 |  |  |  | RSD | 2.0577 | 1.5368 |  | 2.2652 |  |  |  | RSD | 2.0585 | 1.8239 |  | 2.0424
|  | Cer d18:1/16:0 | D7 Cer d18:1/16:0 | D7 Cer d18:1/16:0 | Cer d18:1/16:0 |  |  |  |  | Cer d18:1/18:0 | D7 Cer d18:1/18:0 | D7 Cer d18:1/18:0 | Cer d18:1/18:0 |  |  |  |  | Cer d18:1/24:0 | D7 Cer d18:1/24:0 | D7 Cer d18:1/24:0 | Cer d18:1/24:0 |  |  |  |  | Cer d18:1/24:1 | D7 Cer d18:1/24:1 | D7 Cer d18:1/24:1 | Cer d18:1/24:1
|  | Analyte Peak Area (counts) | Analyte Peak Area (counts) | c [pmol/µL] | c [pmol/µL] |  |  |  |  | Analyte Peak Area (counts) | Analyte Peak Area (counts) | c [pmol/µL] | c [pmol/µL] |  |  |  |  | Analyte Peak Area (counts) | Analyte Peak Area (counts) | c [pmol/µL] | c [pmol/µL] |  |  |  |  | Analyte Peak Area (counts) | Analyte Peak Area (counts) | c [pmol/µL] | c [pmol/µL]
| Total Blank 1 1 | 3120 | 145 | - | - |  |  |  | Total Blank 1 1 | 1630 | 152 | - | - |  |  |  | Total Blank 1 1 | 27500 | 112 | - | - |  |  |  | Total Blank 1 1 | 1260 | 210 | - | -
| Total Blank 1 2 | 1460 | 135 | - | - |  |  |  | Total Blank 1 2 | 667 | 164 | - | - |  |  |  | Total Blank 1 2 | 18600 | 574 | - | - |  |  |  | Total Blank 1 2 | 1270 | 135 | - | -
| Total Blank 1 3 | 1430 | 120 | - | - |  |  |  | Total Blank 1 3 | 698 | 180 | - | - |  |  |  | Total Blank 1 3 | 18800 | 624 | - | - |  |  |  | Total Blank 1 3 | 1530 | 101 | - | -
| Total Blank 2 1 | 1260 | 229 | - | - |  |  |  | Total Blank 2 1 | 841 | 168 | - | - |  |  |  | Total Blank 2 1 | 23500 | 5550 | - | - |  |  |  | Total Blank 2 1 | 2540 | 502 | - | -
| Total Blank 2 2 | 1420 | 199 | - | - |  |  |  | Total Blank 2 2 | 768 | 299 | - | - |  |  |  | Total Blank 2 2 | 15900 | 4340 | - | - |  |  |  | Total Blank 2 2 | 1110 | 567 | - | -
| Total Blank 2 3 | 1310 | 82.9 | - | - |  |  |  | Total Blank 2 3 | 413 | 152 | - | - |  |  |  | Total Blank 2 3 | 17200 | 3100 | - | - |  |  |  | Total Blank 2 3 | 737 | 358 | - | -
| Total Blank 3 1 | 3210 | 223 | - | - |  |  |  | Total Blank 3 1 | 1570 | 157 | - | - |  |  |  | Total Blank 3 1 | 17200 | 1070 | - | - |  |  |  | Total Blank 3 1 | 946 | 153 | - | -
| Total Blank 3 2 | 1610 | 187 | - | - |  |  |  | Total Blank 3 2 | 873 | 257 | - | - |  |  |  | Total Blank 3 2 | 20600 | 3360 | - | - |  |  |  | Total Blank 3 2 | 1340 | 421 | - | -
| Total Blank 3 3 | 1880 | 142 | - | - |  |  |  | Total Blank 3 3 | 709 | 162 | - | - |  |  |  | Total Blank 3 3 | 18200 | 3280 | - | - |  |  |  | Total Blank 3 3 | 733 | 203 | - | -
| Total Blank 4 1 | 5130 | 834 | - | - |  |  |  | Total Blank 4 1 | 3080 | 385 | - | - |  |  |  | Total Blank 4 1 | 27300 | 705 | - | - |  |  |  | Total Blank 4 1 | 2750 | 567 | - | -
| Total Blank 4 2 | 1900 | 206 | - | - |  |  |  | Total Blank 4 2 | 1110 | 291 | - | - |  |  |  | Total Blank 4 2 | 19400 | 364 | - | - |  |  |  | Total Blank 4 2 | 1020 | 171 | - | -
| Total Blank 4 3 | 1830 | 247 | - | - |  |  |  | Total Blank 4 3 | 1250 | 276 | - | - |  |  |  | Total Blank 4 3 | 12700 | 1670 | - | - |  |  |  | Total Blank 4 3 | 720 | 121 | - | -
| Total Blank 5 1 | 3260 | 179 | - | - |  |  |  | Total Blank 5 1 | 798 | 189 | - | - |  |  |  | Total Blank 5 1 | 9770 | 2470 | - | - |  |  |  | Total Blank 5 1 | 952 | 453 | - | -
| Total Blank 5 2 | 1560 | 68.9 | - | - |  |  |  | Total Blank 5 2 | 1130 | 156 | - | - |  |  |  | Total Blank 5 2 | 6890 | 1050 | - | - |  |  |  | Total Blank 5 2 | 576 | 93 | - | -
| Total Blank 5 3 | 1620 | 123 | - | - |  |  |  | Total Blank 5 3 | 917 | 155 | - | - |  |  |  | Total Blank 5 3 | 4990 | 3230 | - | - |  |  |  | Total Blank 5 3 | 620 | 279 | - | -
| Total Blank 6 1 | 5880 | 381 | - | - |  |  |  | Total Blank 6 1 | 1210 | 377 | - | - |  |  |  | Total Blank 6 1 | 11400 | 3130 | - | - |  |  |  | Total Blank 6 1 | 1170 | 338 | - | -
| Total Blank 6 2 | 3500 | 155 | - | - |  |  |  | Total Blank 6 2 | 1730 | 342 | - | - |  |  |  | Total Blank 6 2 | 11100 | 4000 | - | - |  |  |  | Total Blank 6 2 | 732 | 540 | - | -
| Total Blank 6 3 | 3860 | 158 | - | - |  |  |  | Total Blank 6 3 | 2160 | 144 | - | - |  |  |  | Total Blank 6 3 | 11800 | 3430 | - | - |  |  |  | Total Blank 6 3 | 882 | 494 | - | -
| Average | 2513.3333 | 211.9333 | [5 empty] | Average | 1197.4444 | 222.5556 | [5 empty] | Average | 16269.4444 | 2336.6111 | [5 empty] | Average | 1160.4444 | 317
| SD | 1346.679 | 165.6655 | [5 empty] | SD | 629.5877 | 82.1606 | [5 empty] | SD | 6176.3811 | 1554.8024 | [5 empty] | SD | 586.0953 | 169.0151
| RSD | 53.5814 | 78.1687 | [5 empty] | RSD | 52.5776 | 36.9169 | [5 empty] | RSD | 37.9631 | 66.5409 | [5 empty] | RSD | 50.5061 | 53.3171
|  | Cer d18:1/16:0 | D7 Cer d18:1/16:0 | D7 Cer d18:1/16:0 | Cer d18:1/16:0 |  |  |  |  | Cer d18:1/18:0 | D7 Cer d18:1/18:0 | D7 Cer d18:1/18:0 | Cer d18:1/18:0 |  |  |  |  | Cer d18:1/24:0 | D7 Cer d18:1/24:0 | D7 Cer d18:1/24:0 | Cer d18:1/24:0 |  |  |  |  | Cer d18:1/24:1 | D7 Cer d18:1/24:1 | D7 Cer d18:1/24:1 | Cer d18:1/24:1
|  | Analyte Peak Area (counts) | Analyte Peak Area (counts) | c [pmol/µL] | c [pmol/µL] |  |  |  |  | Analyte Peak Area (counts) | Analyte Peak Area (counts) | c [pmol/µL] | c [pmol/µL] |  |  |  |  | Analyte Peak Area (counts) | Analyte Peak Area (counts) | c [pmol/µL] | c [pmol/µL] |  |  |  |  | Analyte Peak Area (counts) | Analyte Peak Area (counts) | c [pmol/µL] | c [pmol/µL]
1 | Matrix blank 1 1 | 170000 | 113 | - | - |  |  |  | Matrix blank 1 1 | 81300 | 200 | - | - |  |  |  | Matrix blank 1 1 | 2940000 | 551 | - | - |  |  |  | Matrix blank 1 1 | 766000 | 97.9 | - | -
1 | Matrix blank 2 1 | 198000 | 125 | - | - |  |  |  | Matrix blank 2 1 | 97000 | 626 | - | - |  |  |  | Matrix blank 2 1 | 3340000 | 6060 | - | - |  |  |  | Matrix blank 2 1 | 872000 | 1110 | - | -
2 | Matrix blank 1 2 | 177000 | 406 | - | - |  |  |  | Matrix blank 1 2 | 80200 | 322 | - | - |  |  |  | Matrix blank 1 2 | 3010000 | 194 | - | - |  |  |  | Matrix blank 1 2 | 792000 | 257 | - | -
2 | Matrix blank 2 2 | 205000 | 181 | - | - |  |  |  | Matrix blank 2 2 | 99100 | 160 | - | - |  |  |  | Matrix blank 2 2 | 3340000 | 5680 | - | - |  |  |  | Matrix blank 2 2 | 883000 | 654 | - | -
3 | Matrix blank 1 3 | 184000 | 126 | - | - |  |  |  | Matrix blank 1 3 | 84100 | 281 | - | - |  |  |  | Matrix blank 1 3 | 3050000 | 975 | - | - |  |  |  | Matrix blank 1 3 | 820000 | 141 | - | -
3 | Matrix blank 2 3 | 199000 | 191 | - | - |  |  |  | Matrix blank 2 3 | 98800 | 225 | - | - |  |  |  | Matrix blank 2 3 | 3420000 | 5370 | - | - |  |  |  | Matrix blank 2 3 | 890000 | 1090 | - | -
| Average | 188833.3333 | 190.3333 | [5 empty] | Average | 90083.3333 | 302.3333 | [5 empty] | Average | 3183333.3333 | 3138.3333 | [5 empty] | Average | 837166.6667 | 558.3167
| SD | 12694.0493 | 100.8194 | [5 empty] | SD | 8324.1449 | 154.0094 | [5 empty] | SD | 188030.7303 | 2582.6326 | [5 empty] | SD | 47442.6554 | 422.9116
| RSD | 6.7224 | 52.9699 | [5 empty] | RSD | 9.2405 | 50.9403 | [5 empty] | RSD | 5.9067 | 82.2931 | [5 empty] | RSD | 5.6671 | 75.7476
|  | Cer d18:1/16:0 | D7 Cer d18:1/16:0 | D7 Cer d18:1/16:0 | Cer d18:1/16:0 |  |  |  |  | Cer d18:1/18:0 | D7 Cer d18:1/18:0 | D7 Cer d18:1/18:0 | Cer d18:1/18:0 |  |  |  |  | Cer d18:1/24:0 | D7 Cer d18:1/24:0 | D7 Cer d18:1/24:0 | Cer d18:1/24:0 |  |  |  |  | Cer d18:1/24:1 | D7 Cer d18:1/24:1 | D7 Cer d18:1/24:1 | Cer d18:1/24:1
| REPEATED | Analyte Peak Area (counts) | Analyte Peak Area (counts) | c [pmol/µL] | c [pmol/µL] |  |  |  | REPEATED | Analyte Peak Area (counts) | Analyte Peak Area (counts) | c [pmol/µL] | c [pmol/µL] |  |  |  | REPEATED | Analyte Peak Area (counts) | Analyte Peak Area (counts) | c [pmol/µL] | c [pmol/µL] |  |  |  | REPEATED | Analyte Peak Area (counts) | Analyte Peak Area (counts) | c [pmol/µL] | c [pmol/µL]
1 | Matrix blank QC 1 1 | 194000 | 139 | - | - |  |  |  | Matrix blank QC 1 1 | 103000 | 502 | - | - |  |  |  | Matrix blank QC 1 1 | 3310000 | 701 | - | - |  |  |  | Matrix blank QC 1 1 | 876000 | 165 | - | -
1 | Matrix blank QC 2 1 | 198000 | 139 | - | - |  |  |  | Matrix blank QC 2 1 | 109000 | 120 | - | - |  |  |  | Matrix blank QC 2 1 | 3050000 | 378 | - | - |  |  |  | Matrix blank QC 2 1 | 789000 | 80.8 | - | -
1 | Matrix blank QC 3 1 | 201000 | 124 | - | - |  |  |  | Matrix blank QC 3 1 | 111000 | 135 | - | - |  |  |  | Matrix blank QC 3 1 | 3060000 | 1390 | - | - |  |  |  | Matrix blank QC 3 1 | 807000 | 107 | - | -
2 | Matrix blank QC 1 2 | 193000 | 128 | - | - |  |  |  | Matrix blank QC 1 2 | 110000 | 145 | - | - |  |  |  | Matrix blank QC 1 2 | 3130000 | 495 | - | - |  |  |  | Matrix blank QC 1 2 | 840000 | 133 | - | -
2 | Matrix blank QC 2 2 | 197000 | 305 | - | - |  |  |  | Matrix blank QC 2 2 | 112000 | 217 | - | - |  |  |  | Matrix blank QC 2 2 | 3070000 | 2490 | - | - |  |  |  | Matrix blank QC 2 2 | 839000 | 267 | - | -
2 | Matrix blank QC 3 2 | 202000 | 274 | - | - |  |  |  | Matrix blank QC 3 2 | 110000 | 267 | - | - |  |  |  | Matrix blank QC 3 2 | 3100000 | 4810 | - | - |  |  |  | Matrix blank QC 3 2 | 809000 | 213 | - | -
3 | Matrix blank QC 1 3 | 190000 | 275 | - | - |  |  |  | Matrix blank QC 1 3 | 112000 | 127 | - | - |  |  |  | Matrix blank QC 1 3 | 3120000 | 1810 | - | - |  |  |  | Matrix blank QC 1 3 | 820000 | 152 | - | -
3 | Matrix blank QC 2 3 | 195000 | 199 | - | - |  |  |  | Matrix blank QC 2 3 | 115000 | 122 | - | - |  |  |  | Matrix blank QC 2 3 | 3120000 | 1650 | - | - |  |  |  | Matrix blank QC 2 3 | 796000 | 52.4 | - | -
3 | Matrix blank QC 3 3 | 204000 | 352 | - | - |  |  |  | Matrix blank QC 3 3 | 111000 | 855 | - | - |  |  |  | Matrix blank QC 3 3 | 3150000 | 14000 | - | - |  |  |  | Matrix blank QC 3 3 | 836000 | 600 | - | -
| Average | 197111.1111 | 215 | [5 empty] | Average | 110333.3333 | 276.6667 | [5 empty] | Average | 3123333.3333 | 3080.4444 | [5 empty] | Average | 823555.5556 | 196.6889
| SD | 4331.9086 | 82.7365 | [5 empty] | SD | 3055.0505 | 235.0343 | [5 empty] | SD | 73333.3333 | 4065.3545 | [5 empty] | SD | 25530.4224 | 155.3565
| RSD | 2.1977 | 38.4821 | [5 empty] | RSD | 2.7689 | 84.9521 | [5 empty] | RSD | 2.3479 | 131.973 | [5 empty] | RSD | 3.1 | 78.9859
|  | Cer d18:1/16:0 | D7 Cer d18:1/16:0 | D7 Cer d18:1/16:0 | Cer d18:1/16:0 |  |  |  |  | Cer d18:1/18:0 | D7 Cer d18:1/18:0 | D7 Cer d18:1/18:0 | Cer d18:1/18:0 |  |  |  |  | Cer d18:1/24:0 | D7 Cer d18:1/24:0 | D7 Cer d18:1/24:0 | Cer d18:1/24:0 |  |  |  |  | Cer d18:1/24:1 | D7 Cer d18:1/24:1 | D7 Cer d18:1/24:1 | Cer d18:1/24:1
|  | Analyte Peak Area (counts) | Analyte Peak Area (counts) | c [pmol/µL] | c [pmol/µL] |  |  |  |  | Analyte Peak Area (counts) | Analyte Peak Area (counts) | c [pmol/µL] | c [pmol/µL] |  |  |  |  | Analyte Peak Area (counts) | Analyte Peak Area (counts) | c [pmol/µL] | c [pmol/µL] |  |  |  |  | Analyte Peak Area (counts) | Analyte Peak Area (counts) | c [pmol/µL] | c [pmol/µL]
| Pooled QC 1 | 195000 | 157000 | 0.125 | 0.15525 |  |  |  | Pooled QC 1 | 93800 | 171000 | 0.05 | 0.027427 |  |  |  | Pooled QC 1 | 3490000 | 3560000 | 1.5 | 1.4705 |  |  |  | Pooled QC 1 | 918000 | 957000 | 0.5 | 0.47962
| Pooled QC 2 | 192000 | 155000 | 0.125 | 0.15484 |  |  |  | Pooled QC 2 | 93600 | 171000 | 0.05 | 0.027368 |  |  |  | Pooled QC 2 | 3480000 | 3550000 | 1.5 | 1.4704 |  |  |  | Pooled QC 2 | 884000 | 965000 | 0.5 | 0.45803
| Pooled QC 3 | 196000 | 161000 | 0.125 | 0.15217 |  |  |  | Pooled QC 3 | 89800 | 180000 | 0.05 | 0.024944 |  |  |  | Pooled QC 3 | 3480000 | 3600000 | 1.5 | 1.45 |  |  |  | Pooled QC 3 | 915000 | 977000 | 0.5 | 0.46827
| Pooled QC 4 | 194000 | 158000 | 0.125 | 0.15348 |  |  |  | Pooled QC 4 | 94900 | 177000 | 0.05 | 0.026808 |  |  |  | Pooled QC 4 | 3500000 | 3730000 | 1.5 | 1.4075 |  |  |  | Pooled QC 4 | 906000 | 992000 | 0.5 | 0.45665
| Pooled QC 5 | 195000 | 161000 | 0.125 | 0.1514 |  |  |  | Pooled QC 5 | 99100 | 180000 | 0.05 | 0.027528 |  |  |  | Pooled QC 5 | 3450000 | 3530000 | 1.5 | 1.466 |  |  |  | Pooled QC 5 | 899000 | 964000 | 0.5 | 0.46629
| Pooled QC 6 | 204000 | 164000 | 0.125 | 0.15549 |  |  |  | Pooled QC 6 | 99400 | 191000 | 0.05 | 0.026021 |  |  |  | Pooled QC 6 | 3580000 | 3740000 | 1.5 | 1.4358 |  |  |  | Pooled QC 6 | 936000 | 1010000 | 0.5 | 0.46337
| Pooled QC 7 | 203000 | 163000 | 0.125 | 0.15567 |  |  |  | Pooled QC 7 | 106000 | 193000 | 0.05 | 0.027461 |  |  |  | Pooled QC 7 | 3690000 | 3820000 | 1.5 | 1.449 |  |  |  | Pooled QC 7 | 956000 | 1010000 | 0.5 | 0.47327
| Pooled QC 8 | 206000 | 168000 | 0.125 | 0.15327 |  |  |  | Pooled QC 8 | 103000 | 195000 | 0.05 | 0.02641 |  |  |  | Pooled QC 8 | 3770000 | 3950000 | 1.5 | 1.4316 |  |  |  | Pooled QC 8 | 999000 | 1030000 | 0.5 | 0.48495
| Pooled QC 9 | 213000 | 170000 | 0.125 | 0.15662 |  |  |  | Pooled QC 9 | 109000 | 204000 | 0.05 | 0.026716 |  |  |  | Pooled QC 9 | 3710000 | 3860000 | 1.5 | 1.4417 |  |  |  | Pooled QC 9 | 972000 | 1050000 | 0.5 | 0.46286
| Pooled QC 10 | 207000 | 175000 | 0.125 | 0.14786 |  |  |  | Pooled QC 10 | 108000 | 201000 | 0.05 | 0.026866 |  |  |  | Pooled QC 10 | 3800000 | 3850000 | 1.5 | 1.4805 |  |  |  | Pooled QC 10 | 951000 | 1000000 | 0.5 | 0.4755
| Pooled QC 11 | 217000 | 168000 | 0.125 | 0.16146 |  |  |  | Pooled QC 11 | 114000 | 205000 | 0.05 | 0.027805 |  |  |  | Pooled QC 11 | 4120000 | 4290000 | 1.5 | 1.4406 |  |  |  | Pooled QC 11 | 1030000 | 1090000 | 0.5 | 0.47248
| Pooled QC 12 | 211000 | 172000 | 0.125 | 0.15334 |  |  |  | Pooled QC 12 | 114000 | 209000 | 0.05 | 0.027273 |  |  |  | Pooled QC 12 | 4060000 | 4160000 | 1.5 | 1.4639 |  |  |  | Pooled QC 12 | 1040000 | 1110000 | 0.5 | 0.46847
| Pooled QC 13 | 222000 | 179000 | 0.125 | 0.15503 |  |  |  | Pooled QC 13 | 117000 | 215000 | 0.05 | 0.027209 |  |  |  | Pooled QC 13 | 4060000 | 4250000 | 1.5 | 1.4329 |  |  |  | Pooled QC 13 | 1090000 | 1160000 | 0.5 | 0.46983
| Pooled QC 14 | 226000 | 182000 | 0.125 | 0.15522 |  |  |  | Pooled QC 14 | 117000 | 224000 | 0.05 | 0.026116 |  |  |  | Pooled QC 14 | 4120000 | 4370000 | 1.5 | 1.4142 |  |  |  | Pooled QC 14 | 1070000 | 1150000 | 0.5 | 0.46522
| Pooled QC 15 | 224000 | 186000 | 0.125 | 0.15054 |  |  |  | Pooled QC 15 | 121000 | 232000 | 0.05 | 0.026078 |  |  |  | Pooled QC 15 | 4200000 | 4320000 | 1.5 | 1.4583 |  |  |  | Pooled QC 15 | 1060000 | 1140000 | 0.5 | 0.46491
| Pooled QC 16 | 229000 | 186000 | 0.125 | 0.1539 |  |  |  | Pooled QC 16 | 124000 | 236000 | 0.05 | 0.026271 |  |  |  | Pooled QC 16 | 4260000 | 4450000 | 1.5 | 1.436 |  |  |  | Pooled QC 16 | 1120000 | 1190000 | 0.5 | 0.47059
| Pooled QC 17 | 237000 | 192000 | 0.125 | 0.1543 |  |  |  | Pooled QC 17 | 128000 | 244000 | 0.05 | 0.02623 |  |  |  | Pooled QC 17 | 4450000 | 4580000 | 1.5 | 1.4574 |  |  |  | Pooled QC 17 | 1120000 | 1190000 | 0.5 | 0.47059
| Pooled QC 18 | 237000 | 194000 | 0.125 | 0.15271 |  |  |  | Pooled QC 18 | 127000 | 245000 | 0.05 | 0.025918 |  |  |  | Pooled QC 18 | 4390000 | 4590000 | 1.5 | 1.4346 |  |  |  | Pooled QC 18 | 1140000 | 1200000 | 0.5 | 0.475
| Pooled QC 19 | 236000 | 191000 | 0.125 | 0.15445 |  |  |  | Pooled QC 19 | 132000 | 252000 | 0.05 | 0.02619 |  |  |  | Pooled QC 19 | 4350000 | 4610000 | 1.5 | 1.4154 |  |  |  | Pooled QC 19 | 1160000 | 1230000 | 0.5 | 0.47154
| Pooled QC 20 | 243000 | 196000 | 0.125 | 0.15497 |  |  |  | Pooled QC 20 | 138000 | 254000 | 0.05 | 0.027165 |  |  |  | Pooled QC 20 | 4420000 | 4530000 | 1.5 | 1.4636 |  |  |  | Pooled QC 20 | 1170000 | 1230000 | 0.5 | 0.47561
| Pooled QC 21 | 248000 | 201000 | 0.125 | 0.15423 |  |  |  | Pooled QC 21 | 138000 | 277000 | 0.05 | 0.02491 |  |  |  | Pooled QC 21 | 4760000 | 4860000 | 1.5 | 1.4691 |  |  |  | Pooled QC 21 | 1220000 | 1270000 | 0.5 | 0.48031
| Pooled QC 22 | 247000 | 210000 | 0.125 | 0.14702 |  |  |  | Pooled QC 22 | 138000 | 273000 | 0.05 | 0.025275 |  |  |  | Pooled QC 22 | 4660000 | 4910000 | 1.5 | 1.4236 |  |  |  | Pooled QC 22 | 1220000 | 1290000 | 0.5 | 0.47287
REPEATED QC | Pooled QC 11 2 | 193000 | 154000 | 0.125 | 0.15666 |  |  | REPEATED QC | Pooled QC 11 2 | 110000 | 204000 | 0.05 | 0.026961 |  |  | REPEATED QC | Pooled QC 11 2 | 3610000 | 3780000 | 1.5 | 1.4325 |  |  | REPEATED QC | Pooled QC 11 2 | 929000 | 960000 | 0.5 | 0.48385
| Pooled QC 12 2 | 208000 | 162000 | 0.125 | 0.16049 |  |  |  | Pooled QC 12 2 | 113000 | 212000 | 0.05 | 0.026651 |  |  |  | Pooled QC 12 2 | 3720000 | 3860000 | 1.5 | 1.4456 |  |  |  | Pooled QC 12 2 | 945000 | 1040000 | 0.5 | 0.45433
| Pooled QC 13 2 | 213000 | 175000 | 0.125 | 0.15214 |  |  |  | Pooled QC 13 2 | 116000 | 223000 | 0.05 | 0.026009 |  |  |  | Pooled QC 13 2 | 3890000 | 4080000 | 1.5 | 1.4301 |  |  |  | Pooled QC 13 2 | 998000 | 1050000 | 0.5 | 0.47524
| Pooled QC 14 2 | 218000 | 172000 | 0.125 | 0.15843 |  |  |  | Pooled QC 14 2 | 118000 | 227000 | 0.05 | 0.025991 |  |  |  | Pooled QC 14 2 | 3750000 | 3950000 | 1.5 | 1.4241 |  |  |  | Pooled QC 14 2 | 1010000 | 1060000 | 0.5 | 0.47642
| Pooled QC 15 2 | 217000 | 175000 | 0.125 | 0.155 |  |  |  | Pooled QC 15 2 | 124000 | 232000 | 0.05 | 0.026724 |  |  |  | Pooled QC 15 2 | 3880000 | 4020000 | 1.5 | 1.4478 |  |  |  | Pooled QC 15 2 | 979000 | 1060000 | 0.5 | 0.46179
| Pooled QC 16 2 | 214000 | 177000 | 0.125 | 0.15113 |  |  |  | Pooled QC 16 2 | 121000 | 221000 | 0.05 | 0.027376 |  |  |  | Pooled QC 16 2 | 3750000 | 4000000 | 1.5 | 1.4062 |  |  |  | Pooled QC 16 2 | 988000 | 1070000 | 0.5 | 0.46168
| Pooled QC 17 2 | 218000 | 183000 | 0.125 | 0.14891 |  |  |  | Pooled QC 17 2 | 130000 | 241000 | 0.05 | 0.026971 |  |  |  | Pooled QC 17 2 | 3820000 | 3990000 | 1.5 | 1.4361 |  |  |  | Pooled QC 17 2 | 1010000 | 1040000 | 0.5 | 0.48558
| Pooled QC 18 2 | 218000 | 186000 | 0.125 | 0.14651 |  |  |  | Pooled QC 18 2 | 126000 | 240000 | 0.05 | 0.02625 |  |  |  | Pooled QC 18 2 | 3830000 | 4040000 | 1.5 | 1.422 |  |  |  | Pooled QC 18 2 | 1010000 | 1090000 | 0.5 | 0.4633
| Pooled QC 19 2 | 224000 | 190000 | 0.125 | 0.14737 |  |  |  | Pooled QC 19 2 | 137000 | 254000 | 0.05 | 0.026969 |  |  |  | Pooled QC 19 2 | 3850000 | 3910000 | 1.5 | 1.477 |  |  |  | Pooled QC 19 2 | 1010000 | 1060000 | 0.5 | 0.47642
| Pooled QC 20 2 | 226000 | 189000 | 0.125 | 0.14947 |  |  |  | Pooled QC 20 2 | 132000 | 249000 | 0.05 | 0.026506 |  |  |  | Pooled QC 20 2 | 3880000 | 3960000 | 1.5 | 1.4697 |  |  |  | Pooled QC 20 2 | 1020000 | 1080000 | 0.5 | 0.47222
| Pooled QC 21 2 | 235000 | 193000 | 0.125 | 0.1522 |  |  |  | Pooled QC 21 2 | 138000 | 268000 | 0.05 | 0.025746 |  |  |  | Pooled QC 21 2 | 3940000 | 4070000 | 1.5 | 1.4521 |  |  |  | Pooled QC 21 2 | 1050000 | 1080000 | 0.5 | 0.48611
| Pooled QC 22 2 | 238000 | 186000 | 0.125 | 0.15995 |  |  |  | Pooled QC 22 2 | 137000 | 256000 | 0.05 | 0.026758 |  |  |  | Pooled QC 22 2 | 3870000 | 3990000 | 1.5 | 1.4549 |  |  |  | Pooled QC 22 2 | 1020000 | 1040000 | 0.5 | 0.49038
| Average | 217764.7059 | 177382.3529 |  | 0.15357 |  |  |  | Average | 117841.1765 | 222235.2941 |  | 0.026556 |  |  |  | Average | 3929117.6471 | 4081176.4706 |  | 1.4444 |  |  |  | Average | 1021911.7647 | 1083382.3529 |  | 0.47157
| SD | 16049.6634 | 14181.6259 |  | 0.00362 |  |  |  | SD | 14432.5636 | 29545.6365 |  | 0.00071272 |  |  |  | SD | 340547.7219 | 359057.3009 |  | 0.020056 |  |  |  | SD | 89194.2255 | 90841.4957 |  | 0.0087259
| RSD | 7.3702 | 7.9949 |  | 2.3572 |  |  |  | RSD | 12.2475 | 13.2948 |  | 2.6839 |  |  |  | RSD | 8.6673 | 8.7979 |  | 1.3885 |  |  |  | RSD | 8.7282 | 8.385 |  | 1.8504
|  | Average Cer 34:1 | SD | [6 empty] | Average Cer 36:1 | SD | [6 empty] | Average Cer 42:1 | SD | [6 empty] | Average Cer 42:2 | SD
| SRM 1950 | 0.16221 | 0.0038124 | [5 empty] | SRM 1950 | 0.026078 | 0.00058336 | [5 empty] | SRM 1950 | 1.3198 | 0.037048 | [5 empty] | SRM 1950 | 0.43975 | 0.01085
| T1D | 0.13455 | 0.0041008 | [5 empty] | T1D | 0.02875 | 0.00095353 | [5 empty] | T1D | 1.2648 | 0.029756 | [5 empty] | T1D | 0.42876 | 0.017521
| young AA | 0.097047 | 0.0022402 | [5 empty] | young AA | 0.014735 | 0.00049357 | [5 empty] | young AA | 0.87117 | 0.017639 | [5 empty] | young AA | 0.26943 | 0.0082984
| high TAG | 0.17853 | 0.0032253 | [5 empty] | high TAG | 0.029442 | 0.0010686 | [5 empty] | high TAG | 1.9254 | 0.043614 | [5 empty] | high TAG | 0.60111 | 0.012277
| QC | 0.15357 | 0.00362 | [5 empty] | QC | 0.026556 | 0.00071272 | [5 empty] | QC | 1.4444 | 0.020056 | [5 empty] | QC | 0.47157 | 0.0087259
| Calculated Average NIST | 0.14308 | [6 empty] | Calculated Average NIST | 0.024751 | [6 empty] | Calculated Average NIST | 1.3453 | [6 empty] | Calculated Average NIST | 0.43476
| MQC 1 | 0.1461 | 0.0031088 | [5 empty] | MQC 1 | 0.025671 | 0.00061593 | [5 empty] | MQC 1 | 1.4235 | 0.025523 | [5 empty] | MQC 1 | 0.45142 | 0.008118
| MQC 2 | 0.14459 | 0.005181 | [5 empty] | MQC 2 | 0.024999 | 0.00090517 | [5 empty] | MQC 2 | 1.3727 | 0.055757 | [5 empty] | MQC 2 | 0.44817 | 0.013191
WRONG HQC/HLQC | NEW IS VIAL | [7 empty] | NEW IS VIAL | [7 empty] | NEW IS VIAL | [7 empty] | NEW IS VIAL
|  | Cer d18:1/16:0 | D7 Cer d18:1/16:0 | D7 Cer d18:1/16:0 | Cer d18:1/16:0 |  |  |  |  | Cer d18:1/18:0 | D7 Cer d18:1/18:0 | D7 Cer d18:1/18:0 | Cer d18:1/18:0 |  |  |  |  | Cer d18:1/24:0 | D7 Cer d18:1/24:0 | D7 Cer d18:1/24:0 | Cer d18:1/24:0 |  |  |  |  | Cer d18:1/24:1 | D7 Cer d18:1/24:1 | D7 Cer d18:1/24:1 | Cer d18:1/24:1
| Sample Name | Analyte Peak Area (counts) | Analyte Peak Area (counts) | c [pmol/µL] | c [pmol/µL] |  |  |  | Sample Name | Analyte Peak Area (counts) | Analyte Peak Area (counts) | c [pmol/µL] | c [pmol/µL] |  |  |  | Sample Name | Analyte Peak Area (counts) | Analyte Peak Area (counts) | c [pmol/µL] | c [pmol/µL] |  |  |  | Sample Name | Analyte Peak Area (counts) | Analyte Peak Area (counts) | c [pmol/µL] | c [pmol/µL]
1 | LLQC1 1 | 80400 | 171000 | 0.125 | 0.058772 |  |  |  | LLQC1 1 | 29600 | 156000 | 0.05 | 0.0094872 |  |  |  | LLQC1 1 | 976000 | 2730000 | 1.5 | 0.53626 |  |  |  | LLQC1 1 | 257000 | 739000 | 0.5 | 0.17388
1 | LLQC2 1 | 75400 | 173000 | 0.125 | 0.05448 |  |  |  | LLQC2 1 | 28200 | 150000 | 0.05 | 0.0094 |  |  |  | LLQC2 1 | 882000 | 2470000 | 1.5 | 0.53563 |  |  |  | LLQC2 1 | 236000 | 705000 | 0.5 | 0.16738
1 | LLQC3 1 | 76600 | 174000 | 0.125 | 0.055029 |  |  |  | LLQC3 1 | 29300 | 157000 | 0.05 | 0.0093312 |  |  |  | LLQC3 1 | 899000 | 2680000 | 1.5 | 0.50317 |  |  |  | LLQC3 1 | 238000 | 714000 | 0.5 | 0.16667
1 | LLQC4 1 | 77100 | 172000 | 0.125 | 0.056032 |  |  |  | LLQC4 1 | 28800 | 152000 | 0.05 | 0.0094737 |  |  |  | LLQC4 1 | 902000 | 2530000 | 1.5 | 0.53478 |  |  |  | LLQC4 1 | 233000 | 730000 | 0.5 | 0.15959
1 | LLQC5 1 | 72500 | 172000 | 0.125 | 0.052689 |  |  |  | LLQC5 1 | 30400 | 152000 | 0.05 | 0.01 |  |  |  | LLQC5 1 | 909000 | 2520000 | 1.5 | 0.54107 |  |  |  | LLQC5 1 | 229000 | 715000 | 0.5 | 0.16014
1 | LLQC6 1 | 73200 | 177000 | 0.125 | 0.051695 |  |  |  | LLQC6 1 | 29800 | 160000 | 0.05 | 0.0093125 |  |  |  | LLQC6 1 | 875000 | 2660000 | 1.5 | 0.49342 |  |  |  | LLQC6 1 | 226000 | 742000 | 0.5 | 0.15229
2 | LLQC1 2 | 81500 | 174000 | 0.125 | 0.058549 |  |  |  | LLQC1 2 | 29400 | 159000 | 0.05 | 0.0092453 |  |  |  | LLQC1 2 | 954000 | 2630000 | 1.5 | 0.54411 |  |  |  | LLQC1 2 | 257000 | 741000 | 0.5 | 0.17341
2 | LLQC2 2 | 74000 | 169000 | 0.125 | 0.054734 |  |  |  | LLQC2 2 | 27700 | 146000 | 0.05 | 0.0094863 |  |  |  | LLQC2 2 | 883000 | 2570000 | 1.5 | 0.51537 |  |  |  | LLQC2 2 | 238000 | 707000 | 0.5 | 0.16832
2 | LLQC3 2 | 77900 | 170000 | 0.125 | 0.057279 |  |  |  | LLQC3 2 | 29000 | 152000 | 0.05 | 0.0095395 |  |  |  | LLQC3 2 | 914000 | 2630000 | 1.5 | 0.52129 |  |  |  | LLQC3 2 | 228000 | 725000 | 0.5 | 0.15724
2 | LLQC4 2 | 78100 | 171000 | 0.125 | 0.057091 |  |  |  | LLQC4 2 | 30600 | 147000 | 0.05 | 0.010408 |  |  |  | LLQC4 2 | 875000 | 2530000 | 1.5 | 0.51877 |  |  |  | LLQC4 2 | 223000 | 715000 | 0.5 | 0.15594
2 | LLQC5 2 | 76200 | 170000 | 0.125 | 0.056029 |  |  |  | LLQC5 2 | 31800 | 154000 | 0.05 | 0.010325 |  |  |  | LLQC5 2 | 896000 | 2570000 | 1.5 | 0.52296 |  |  |  | LLQC5 2 | 234000 | 727000 | 0.5 | 0.16094
2 | LLQC6 2 | 72600 | 184000 | 0.125 | 0.049321 |  |  |  | LLQC6 2 | 29500 | 163000 | 0.05 | 0.0090491 |  |  |  | LLQC6 2 | 883000 | 2710000 | 1.5 | 0.48875 |  |  |  | LLQC6 2 | 237000 | 735000 | 0.5 | 0.16122
3 | LLQC1 3 | 77500 | 174000 | 0.125 | 0.055675 |  |  |  | LLQC1 3 | 30200 | 152000 | 0.05 | 0.0099342 |  |  |  | LLQC1 3 | 964000 | 2580000 | 1.5 | 0.56047 |  |  |  | LLQC1 3 | 253000 | 752000 | 0.5 | 0.16822
3 | LLQC2 3 | 76900 | 173000 | 0.125 | 0.055564 |  |  |  | LLQC2 3 | 31100 | 156000 | 0.05 | 0.0099679 |  |  |  | LLQC2 3 | 894000 | 2550000 | 1.5 | 0.52588 |  |  |  | LLQC2 3 | 238000 | 705000 | 0.5 | 0.16879
3 | LLQC3 3 | 79300 | 173000 | 0.125 | 0.057298 |  |  |  | LLQC3 3 | 29800 | 153000 | 0.05 | 0.0097386 |  |  |  | LLQC3 3 | 930000 | 2700000 | 1.5 | 0.51667 |  |  |  | LLQC3 3 | 237000 | 752000 | 0.5 | 0.15758
3 | LLQC4 3 | 74200 | 176000 | 0.125 | 0.052699 |  |  |  | LLQC4 3 | 31400 | 158000 | 0.05 | 0.0099367 |  |  |  | LLQC4 3 | 919000 | 2590000 | 1.5 | 0.53224 |  |  |  | LLQC4 3 | 235000 | 728000 | 0.5 | 0.1614
3 | LLQC5 3 | 76300 | 171000 | 0.125 | 0.055775 |  |  |  | LLQC5 3 | 30700 | 147000 | 0.05 | 0.010442 |  |  |  | LLQC5 3 | 883000 | 2610000 | 1.5 | 0.50747 |  |  |  | LLQC5 3 | 233000 | 713000 | 0.5 | 0.16339
3 | LLQC6 3 | 73300 | 175000 | 0.125 | 0.052357 |  |  |  | LLQC6 3 | 29500 | 162000 | 0.05 | 0.0091049 |  |  |  | LLQC6 3 | 882000 | 2660000 | 1.5 | 0.49737 |  |  |  | LLQC6 3 | 234000 | 757000 | 0.5 | 0.15456
| Average | 76277.7778 | 173277.7778 |  | 0.055059 |  |  |  | Average | 29822.2222 | 154222.2222 |  | 0.0096768 |  |  |  | Average | 906666.6667 | 2606666.6667 |  | 0.52198 |  |  |  | Average | 237000 | 727888.8889 |  | 0.16283
| SD | 2565.0548 | 3330.091 |  | 0.0024295 |  |  |  | SD | 1039.5274 | 4905.2756 |  | 0.00042398 |  |  |  | SD | 30136.7255 | 71336.4485 |  | 0.018535 |  |  |  | SD | 9357.1126 | 16329.5536 |  | 0.0061296
| RSD | 3.3628 | 1.9218 |  | 4.4125 |  |  |  | RSD | 3.4857 | 3.1807 |  | 4.3814 |  |  |  | RSD | 3.3239 | 2.7367 |  | 3.5508 |  |  |  | RSD | 3.9481 | 2.2434 |  | 3.7644
|  | Cer d18:1/16:0 | D7 Cer d18:1/16:0 | D7 Cer d18:1/16:0 | Cer d18:1/16:0 |  |  |  |  | Cer d18:1/18:0 | D7 Cer d18:1/18:0 | D7 Cer d18:1/18:0 | Cer d18:1/18:0 |  |  |  |  | Cer d18:1/24:0 | D7 Cer d18:1/24:0 | D7 Cer d18:1/24:0 | Cer d18:1/24:0 |  |  |  |  | Cer d18:1/24:1 | D7 Cer d18:1/24:1 | D7 Cer d18:1/24:1 | Cer d18:1/24:1
| Sample Name | Analyte Peak Area (counts) | Analyte Peak Area (counts) | c [pmol/µL] | c [pmol/µL] |  |  |  |  | Analyte Peak Area (counts) | Analyte Peak Area (counts) | c [pmol/µL] | c [pmol/µL] |  |  |  | Sample Name | Analyte Peak Area (counts) | Analyte Peak Area (counts) | c [pmol/µL] | c [pmol/µL] |  |  |  | Sample Name | Analyte Peak Area (counts) | Analyte Peak Area (counts) | c [pmol/µL] | c [pmol/µL]
1 | LQC1 1 | 118000 | 177000 | 0.125 | 0.083333 |  |  |  | LQC1 1 | 50000 | 175000 | 0.05 | 0.014286 |  |  |  | LQC1 1 | 1510000 | 2960000 | 1.5 | 0.7652 |  |  |  | LQC1 1 | 378000 | 802000 | 0.5 | 0.23566
1 | LQC2 1 | 112000 | 184000 | 0.125 | 0.076087 |  |  |  | LQC2 1 | 44600 | 172000 | 0.05 | 0.012965 |  |  |  | LQC2 1 | 1430000 | 2840000 | 1.5 | 0.75528 |  |  |  | LQC2 1 | 365000 | 783000 | 0.5 | 0.23308
1 | LQC3 1 | 111000 | 179000 | 0.125 | 0.077514 |  |  |  | LQC3 1 | 45200 | 169000 | 0.05 | 0.013373 |  |  |  | LQC3 1 | 1400000 | 2880000 | 1.5 | 0.72917 |  |  |  | LQC3 1 | 357000 | 805000 | 0.5 | 0.22174
1 | LQC4 1 | 112000 | 184000 | 0.125 | 0.076087 |  |  |  | LQC4 1 | 44400 | 176000 | 0.05 | 0.012614 |  |  |  | LQC4 1 | 1420000 | 2900000 | 1.5 | 0.73448 |  |  |  | LQC4 1 | 363000 | 816000 | 0.5 | 0.22243
1 | LQC5 1 | 108000 | 182000 | 0.125 | 0.074176 |  |  |  | LQC5 1 | 45100 | 169000 | 0.05 | 0.013343 |  |  |  | LQC5 1 | 1380000 | 2840000 | 1.5 | 0.72887 |  |  |  | LQC5 1 | 352000 | 795000 | 0.5 | 0.22138
1 | LQC6 1 | 114000 | 175000 | 0.125 | 0.081429 |  |  |  | LQC6 1 | 46900 | 171000 | 0.05 | 0.013713 |  |  |  | LQC6 1 | 1410000 | 2850000 | 1.5 | 0.74211 |  |  |  | LQC6 1 | 380000 | 813000 | 0.5 | 0.2337
2 | LQC1 2 | 114000 | 182000 | 0.125 | 0.078297 |  |  |  | LQC1 2 | 49800 | 168000 | 0.05 | 0.014821 |  |  |  | LQC1 2 | 1430000 | 2940000 | 1.5 | 0.72959 |  |  |  | LQC1 2 | 380000 | 797000 | 0.5 | 0.23839
2 | LQC2 2 | 107000 | 181000 | 0.125 | 0.073895 |  |  |  | LQC2 2 | 45200 | 170000 | 0.05 | 0.013294 |  |  |  | LQC2 2 | 1410000 | 2840000 | 1.5 | 0.74472 |  |  |  | LQC2 2 | 359000 | 793000 | 0.5 | 0.22636
2 | LQC3 2 | 110000 | 182000 | 0.125 | 0.075549 |  |  |  | LQC3 2 | 46400 | 166000 | 0.05 | 0.013976 |  |  |  | LQC3 2 | 1400000 | 2940000 | 1.5 | 0.71429 |  |  |  | LQC3 2 | 363000 | 796000 | 0.5 | 0.22802
2 | LQC4 2 | 111000 | 179000 | 0.125 | 0.077514 |  |  |  | LQC4 2 | 46100 | 169000 | 0.05 | 0.013639 |  |  |  | LQC4 2 | 1400000 | 2950000 | 1.5 | 0.71186 |  |  |  | LQC4 2 | 360000 | 810000 | 0.5 | 0.22222
2 | LQC5 2 | 109000 | 182000 | 0.125 | 0.074863 |  |  |  | LQC5 2 | 44600 | 172000 | 0.05 | 0.012965 |  |  |  | LQC5 2 | 1400000 | 2970000 | 1.5 | 0.70707 |  |  |  | LQC5 2 | 359000 | 799000 | 0.5 | 0.22466
2 | LQC6 2 | 110000 | 186000 | 0.125 | 0.073925 |  |  |  | LQC6 2 | 46300 | 173000 | 0.05 | 0.013382 |  |  |  | LQC6 2 | 1440000 | 2860000 | 1.5 | 0.75524 |  |  |  | LQC6 2 | 365000 | 796000 | 0.5 | 0.22927
3 | LQC1 3 | 115000 | 182000 | 0.125 | 0.078984 |  |  |  | LQC1 3 | 48000 | 176000 | 0.05 | 0.013636 |  |  |  | LQC1 3 | 1490000 | 2860000 | 1.5 | 0.78147 |  |  |  | LQC1 3 | 388000 | 788000 | 0.5 | 0.24619
3 | LQC2 3 | 110000 | 179000 | 0.125 | 0.076816 |  |  |  | LQC2 3 | 45400 | 169000 | 0.05 | 0.013432 |  |  |  | LQC2 3 | 1420000 | 2820000 | 1.5 | 0.75532 |  |  |  | LQC2 3 | 361000 | 783000 | 0.5 | 0.23052
3 | LQC3 3 | 114000 | 189000 | 0.125 | 0.075397 |  |  |  | LQC3 3 | 47100 | 174000 | 0.05 | 0.013534 |  |  |  | LQC3 3 | 1420000 | 2820000 | 1.5 | 0.75532 |  |  |  | LQC3 3 | 372000 | 822000 | 0.5 | 0.22628
3 | LQC4 3 | 106000 | 180000 | 0.125 | 0.073611 |  |  |  | LQC4 3 | 44000 | 171000 | 0.05 | 0.012865 |  |  |  | LQC4 3 | 1380000 | 2830000 | 1.5 | 0.73145 |  |  |  | LQC4 3 | 356000 | 807000 | 0.5 | 0.22057
3 | LQC5 3 | 109000 | 184000 | 0.125 | 0.074049 |  |  |  | LQC5 3 | 45100 | 170000 | 0.05 | 0.013265 |  |  |  | LQC5 3 | 1330000 | 2880000 | 1.5 | 0.69271 |  |  |  | LQC5 3 | 351000 | 811000 | 0.5 | 0.2164
3 | LQC6 3 | 111000 | 181000 | 0.125 | 0.076657 |  |  |  | LQC6 3 | 44900 | 174000 | 0.05 | 0.012902 |  |  |  | LQC6 3 | 1410000 | 2870000 | 1.5 | 0.73693 |  |  |  | LQC6 3 | 365000 | 782000 | 0.5 | 0.23338
| Average | 111166.6667 | 181555.5556 |  | 0.076566 |  |  |  | Average | 46061.1111 | 171333.3333 |  | 0.013445 |  |  |  | Average | 1415555.5556 | 2880555.5556 |  | 0.73728 |  |  |  | Average | 365222.2222 | 799888.8889 |  | 0.22835
| SD | 2929.7326 | 3166.1793 |  | 0.0025906 |  |  |  | SD | 1693.1689 | 2788.8668 |  | 0.00052147 |  |  |  | SD | 38329.3074 | 49043.3167 |  | 0.021628 |  |  |  | SD | 10075.0272 | 11474.0717 |  | 0.0072448
| RSD | 2.6354 | 1.7439 |  | 3.3835 |  |  |  | RSD | 3.6759 | 1.6277 |  | 3.8786 |  |  |  | RSD | 2.7077 | 1.7026 |  | 2.9335 |  |  |  | RSD | 2.7586 | 1.4345 |  | 3.1727
|  | Cer d18:1/16:0 | D7 Cer d18:1/16:0 | D7 Cer d18:1/16:0 | Cer d18:1/16:0 |  |  |  |  | Cer d18:1/18:0 | D7 Cer d18:1/18:0 | D7 Cer d18:1/18:0 | Cer d18:1/18:0 |  |  |  |  | Cer d18:1/24:0 | D7 Cer d18:1/24:0 | D7 Cer d18:1/24:0 | Cer d18:1/24:0 |  |  |  |  | Cer d18:1/24:1 | D7 Cer d18:1/24:1 | D7 Cer d18:1/24:1 | Cer d18:1/24:1
| Sample Name | Analyte Peak Area (counts) | Analyte Peak Area (counts) | c [pmol/µL] | c [pmol/µL] |  |  |  |  | Analyte Peak Area (counts) | Analyte Peak Area (counts) | c [pmol/µL] | c [pmol/µL] |  |  |  | Sample Name | Analyte Peak Area (counts) | Analyte Peak Area (counts) | c [pmol/µL] | c [pmol/µL] |  |  |  | Sample Name | Analyte Peak Area (counts) | Analyte Peak Area (counts) | c [pmol/µL] | c [pmol/µL]
1 | MQC1 1 | 203000 | 174000 | 0.125 | 0.14583 |  |  |  | MQC1 1 | 104000 | 200000 | 0.05 | 0.026 |  |  |  | MQC1 1 | 3230000 | 3500000 | 1.5 | 1.3843 |  |  |  | MQC1 1 | 861000 | 980000 | 0.5 | 0.43929
1 | MQC2 1 | 196000 | 169000 | 0.125 | 0.14497 |  |  |  | MQC2 1 | 96300 | 189000 | 0.05 | 0.025476 |  |  |  | MQC2 1 | 3150000 | 3270000 | 1.5 | 1.445 |  |  |  | MQC2 1 | 825000 | 910000 | 0.5 | 0.4533
1 | MQC3 1 | 200000 | 165000 | 0.125 | 0.15152 |  |  |  | MQC3 1 | 98700 | 187000 | 0.05 | 0.02639 |  |  |  | MQC3 1 | 3280000 | 3470000 | 1.5 | 1.4179 |  |  |  | MQC3 1 | 832000 | 939000 | 0.5 | 0.44302
1 | MQC4 1 | 200000 | 174000 | 0.125 | 0.14368 |  |  |  | MQC4 1 | 101000 | 196000 | 0.05 | 0.025765 |  |  |  | MQC4 1 | 3270000 | 3460000 | 1.5 | 1.4176 |  |  |  | MQC4 1 | 836000 | 939000 | 0.5 | 0.44515
1 | MQC5 1 | 203000 | 173000 | 0.125 | 0.14668 |  |  |  | MQC5 1 | 99500 | 199000 | 0.05 | 0.025 |  |  |  | MQC5 1 | 3190000 | 3410000 | 1.5 | 1.4032 |  |  |  | MQC5 1 | 830000 | 925000 | 0.5 | 0.44865
1 | MQC6 1 | 205000 | 168000 | 0.125 | 0.15253 |  |  |  | MQC6 1 | 103000 | 197000 | 0.05 | 0.026142 |  |  |  | MQC6 1 | 3150000 | 3260000 | 1.5 | 1.4494 |  |  |  | MQC6 1 | 837000 | 918000 | 0.5 | 0.45588
2 | MQC1 2 | 198000 | 167000 | 0.125 | 0.1482 |  |  |  | MQC1 2 | 101000 | 192000 | 0.05 | 0.026302 |  |  |  | MQC1 2 | 3260000 | 3430000 | 1.5 | 1.4257 |  |  |  | MQC1 2 | 860000 | 950000 | 0.5 | 0.45263
2 | MQC2 2 | 198000 | 167000 | 0.125 | 0.1482 |  |  |  | MQC2 2 | 98000 | 187000 | 0.05 | 0.026203 |  |  |  | MQC2 2 | 3220000 | 3310000 | 1.5 | 1.4592 |  |  |  | MQC2 2 | 815000 | 912000 | 0.5 | 0.44682
2 | MQC3 2 | 190000 | 166000 | 0.125 | 0.14307 |  |  |  | MQC3 2 | 98300 | 192000 | 0.05 | 0.025599 |  |  |  | MQC3 2 | 3260000 | 3360000 | 1.5 | 1.4554 |  |  |  | MQC3 2 | 834000 | 918000 | 0.5 | 0.45425
2 | MQC4 2 | 200000 | 170000 | 0.125 | 0.14706 |  |  |  | MQC4 2 | 103000 | 197000 | 0.05 | 0.026142 |  |  |  | MQC4 2 | 3200000 | 3330000 | 1.5 | 1.4414 |  |  |  | MQC4 2 | 866000 | 936000 | 0.5 | 0.46261
2 | MQC5 2 | 193000 | 171000 | 0.125 | 0.14108 |  |  |  | MQC5 2 | 98400 | 202000 | 0.05 | 0.024356 |  |  |  | MQC5 2 | 3100000 | 3360000 | 1.5 | 1.3839 |  |  |  | MQC5 2 | 813000 | 920000 | 0.5 | 0.44185
2 | MQC6 2 | 200000 | 174000 | 0.125 | 0.14368 |  |  |  | MQC6 2 | 98900 | 195000 | 0.05 | 0.025359 |  |  |  | MQC6 2 | 3200000 | 3420000 | 1.5 | 1.4035 |  |  |  | MQC6 2 | 859000 | 935000 | 0.5 | 0.45936
3 | MQC1 3 | 193000 | 168000 | 0.125 | 0.1436 |  |  |  | MQC1 3 | 102000 | 193000 | 0.05 | 0.026425 |  |  |  | MQC1 3 | 3190000 | 3330000 | 1.5 | 1.4369 |  |  |  | MQC1 3 | 863000 | 951000 | 0.5 | 0.45373
3 | MQC2 3 | 200000 | 167000 | 0.125 | 0.1497 |  |  |  | MQC2 3 | 96300 | 195000 | 0.05 | 0.024692 |  |  |  | MQC2 3 | 3210000 | 3380000 | 1.5 | 1.4246 |  |  |  | MQC2 3 | 827000 | 923000 | 0.5 | 0.448
3 | MQC3 3 | 193000 | 170000 | 0.125 | 0.14191 |  |  |  | MQC3 3 | 97300 | 190000 | 0.05 | 0.025605 |  |  |  | MQC3 3 | 3180000 | 3360000 | 1.5 | 1.4196 |  |  |  | MQC3 3 | 847000 | 915000 | 0.5 | 0.46284
3 | MQC4 3 | 196000 | 165000 | 0.125 | 0.14848 |  |  |  | MQC4 3 | 96200 | 193000 | 0.05 | 0.024922 |  |  |  | MQC4 3 | 3210000 | 3410000 | 1.5 | 1.412 |  |  |  | MQC4 3 | 865000 | 931000 | 0.5 | 0.46455
3 | MQC5 3 | 198000 | 171000 | 0.125 | 0.14474 |  |  |  | MQC5 3 | 99200 | 196000 | 0.05 | 0.025306 |  |  |  | MQC5 3 | 3180000 | 3460000 | 1.5 | 1.3786 |  |  |  | MQC5 3 | 814000 | 933000 | 0.5 | 0.43623
3 | MQC6 3 | 204000 | 176000 | 0.125 | 0.14489 |  |  |  | MQC6 3 | 104000 | 197000 | 0.05 | 0.026396 |  |  |  | MQC6 3 | 3270000 | 3350000 | 1.5 | 1.4642 |  |  |  | MQC6 3 | 861000 | 941000 | 0.5 | 0.45749
| Average | 198333.3333 | 169722.2222 |  | 0.1461 |  |  |  | Average | 99727.7778 | 194277.7778 |  | 0.025671 |  |  |  | Average | 3208333.3333 | 3381666.6667 |  | 1.4235 |  |  |  | Average | 841388.8889 | 932000 |  | 0.45142
| SD | 4082.4829 | 3296.5564 |  | 0.0031088 |  |  |  | SD | 2549.6853 | 4147.3627 |  | 0.00061593 |  |  |  | SD | 46933.7595 | 66520.6735 |  | 0.025523 |  |  |  | SD | 18463.5945 | 16723.2373 |  | 0.008118
| RSD | 2.0584 | 1.9423 |  | 2.1278 |  |  |  | RSD | 2.5566 | 2.1348 |  | 2.3993 |  |  |  | RSD | 1.4629 | 1.9671 |  | 1.793 |  |  |  | RSD | 2.1944 | 1.7943 |  | 1.7983
|  | Cer d18:1/16:0 | D7 Cer d18:1/16:0 | D7 Cer d18:1/16:0 | Cer d18:1/16:0 |  |  |  |  | Cer d18:1/18:0 | D7 Cer d18:1/18:0 | D7 Cer d18:1/18:0 | Cer d18:1/18:0 |  |  |  |  | Cer d18:1/24:0 | D7 Cer d18:1/24:0 | D7 Cer d18:1/24:0 | Cer d18:1/24:0 |  |  |  |  | Cer d18:1/24:1 | D7 Cer d18:1/24:1 | D7 Cer d18:1/24:1 | Cer d18:1/24:1
| Sample Name | Analyte Peak Area (counts) | Analyte Peak Area (counts) | c [pmol/µL] | c [pmol/µL] |  |  |  |  | Analyte Peak Area (counts) | Analyte Peak Area (counts) | c [pmol/µL] | c [pmol/µL] |  |  |  | Sample Name | Analyte Peak Area (counts) | Analyte Peak Area (counts) | c [pmol/µL] | c [pmol/µL] |  |  |  | Sample Name | Analyte Peak Area (counts) | Analyte Peak Area (counts) | c [pmol/µL] | c [pmol/µL]
1 | HQC1 1 | 194000 | 189000 | 0.125 | 0.12831 |  |  |  | HQC1 1 | 98900 | 209000 | 0.05 | 0.02366 |  |  |  | HQC1 1 | 2750000 | 3390000 | 1.5 | 1.2168 |  |  |  | HQC1 1 | 802000 | 945000 | 0.5 | 0.42434
1 | HQC2 1 | 206000 | 182000 | 0.125 | 0.14148 |  |  |  | HQC2 1 | 112000 | 199000 | 0.05 | 0.028141 |  |  |  | HQC2 1 | 3180000 | 3480000 | 1.5 | 1.3707 |  |  |  | HQC2 1 | 897000 | 971000 | 0.5 | 0.46189
1 | HQC3 1 | 182000 | 176000 | 0.125 | 0.12926 |  |  |  | HQC3 1 | 97600 | 194000 | 0.05 | 0.025155 |  |  |  | HQC3 1 | 2650000 | 3230000 | 1.5 | 1.2307 |  |  |  | HQC3 1 | 767000 | 890000 | 0.5 | 0.4309
1 | HQC4 1 | 184000 | 182000 | 0.125 | 0.12637 |  |  |  | HQC4 1 | 94300 | 193000 | 0.05 | 0.02443 |  |  |  | HQC4 1 | 2600000 | 3170000 | 1.5 | 1.2303 |  |  |  | HQC4 1 | 764000 | 864000 | 0.5 | 0.44213
1 | HQC5 1 | 188000 | 183000 | 0.125 | 0.12842 |  |  |  | HQC5 1 | 92400 | 192000 | 0.05 | 0.024063 |  |  |  | HQC5 1 | 2640000 | 3190000 | 1.5 | 1.2414 |  |  |  | HQC5 1 | 783000 | 874000 | 0.5 | 0.44794
1 | HQC6 1 | 179000 | 179000 | 0.125 | 0.125 |  |  |  | HQC6 1 | 95900 | 189000 | 0.05 | 0.02537 |  |  |  | HQC6 1 | 2590000 | 3180000 | 1.5 | 1.2217 |  |  |  | HQC6 1 | 751000 | 872000 | 0.5 | 0.43062
2 | HQC1 2 | 193000 | 185000 | 0.125 | 0.13041 |  |  |  | HQC1 2 | 99800 | 204000 | 0.05 | 0.024461 |  |  |  | HQC1 2 | 2720000 | 3370000 | 1.5 | 1.2107 |  |  |  | HQC1 2 | 782000 | 938000 | 0.5 | 0.41684
2 | HQC2 2 | 202000 | 175000 | 0.125 | 0.14429 |  |  |  | HQC2 2 | 107000 | 198000 | 0.05 | 0.02702 |  |  |  | HQC2 2 | 3110000 | 3450000 | 1.5 | 1.3522 |  |  |  | HQC2 2 | 916000 | 925000 | 0.5 | 0.49514
2 | HQC3 2 | 186000 | 179000 | 0.125 | 0.12989 |  |  |  | HQC3 2 | 96600 | 194000 | 0.05 | 0.024897 |  |  |  | HQC3 2 | 2630000 | 3240000 | 1.5 | 1.2176 |  |  |  | HQC3 2 | 759000 | 886000 | 0.5 | 0.42833
2 | HQC4 2 | 185000 | 173000 | 0.125 | 0.13367 |  |  |  | HQC4 2 | 95200 | 195000 | 0.05 | 0.02441 |  |  |  | HQC4 2 | 2670000 | 3100000 | 1.5 | 1.2919 |  |  |  | HQC4 2 | 760000 | 888000 | 0.5 | 0.42793
2 | HQC5 2 | 184000 | 184000 | 0.125 | 0.125 |  |  |  | HQC5 2 | 95300 | 198000 | 0.05 | 0.024066 |  |  |  | HQC5 2 | 2550000 | 3150000 | 1.5 | 1.2143 |  |  |  | HQC5 2 | 746000 | 880000 | 0.5 | 0.42386
2 | HQC6 2 | 181000 | 180000 | 0.125 | 0.12569 |  |  |  | HQC6 2 | 94300 | 192000 | 0.05 | 0.024557 |  |  |  | HQC6 2 | 2560000 | 3190000 | 1.5 | 1.2038 |  |  |  | HQC6 2 | 740000 | 857000 | 0.5 | 0.43174
3 | HQC1 3 | 190000 | 186000 | 0.125 | 0.12769 |  |  |  | HQC1 3 | 95200 | 202000 | 0.05 | 0.023564 |  |  |  | HQC1 3 | 2700000 | 3350000 | 1.5 | 1.209 |  |  |  | HQC1 3 | 786000 | 934000 | 0.5 | 0.42077
3 | HQC2 3 | 197000 | 177000 | 0.125 | 0.13912 |  |  |  | HQC2 3 | 105000 | 202000 | 0.05 | 0.02599 |  |  |  | HQC2 3 | 3050000 | 3350000 | 1.5 | 1.3657 |  |  |  | HQC2 3 | 868000 | 924000 | 0.5 | 0.4697
3 | HQC3 3 | 186000 | 177000 | 0.125 | 0.13136 |  |  |  | HQC3 3 | 95200 | 201000 | 0.05 | 0.023682 |  |  |  | HQC3 3 | 2610000 | 3260000 | 1.5 | 1.2009 |  |  |  | HQC3 3 | 766000 | 884000 | 0.5 | 0.43326
3 | HQC4 3 | 182000 | 177000 | 0.125 | 0.12853 |  |  |  | HQC4 3 | 97300 | 191000 | 0.05 | 0.025471 |  |  |  | HQC4 3 | 2590000 | 3090000 | 1.5 | 1.2573 |  |  |  | HQC4 3 | 763000 | 884000 | 0.5 | 0.43156
3 | HQC5 3 | 180000 | 178000 | 0.125 | 0.1264 |  |  |  | HQC5 3 | 96700 | 197000 | 0.05 | 0.024543 |  |  |  | HQC5 3 | 2660000 | 3190000 | 1.5 | 1.2508 |  |  |  | HQC5 3 | 750000 | 873000 | 0.5 | 0.42955
3 | HQC6 3 | 181000 | 176000 | 0.125 | 0.12855 |  |  |  | HQC6 3 | 95200 | 192000 | 0.05 | 0.024792 |  |  |  | HQC6 3 | 2510000 | 3150000 | 1.5 | 1.1952 |  |  |  | HQC6 3 | 753000 | 872000 | 0.5 | 0.43177
| Average | 187777.7778 | 179888.8889 |  | 0.13052 |  |  |  | Average | 97994.4444 | 196777.7778 |  | 0.024904 |  |  |  | Average | 2709444.4444 | 3251666.6667 |  | 1.2489 |  |  |  | Average | 786277.7778 | 897833.3333 |  | 0.43768
| SD | 7561.2723 | 4175.1765 |  | 0.0054766 |  |  |  | SD | 4922.4511 | 5148.4146 |  | 0.0011526 |  |  |  | SD | 190947.1196 | 114903.341 |  | 0.055787 |  |  |  | SD | 50951.8352 | 31843.8028 |  | 0.019146
| RSD | 4.0267 | 2.321 |  | 4.1959 |  |  |  | RSD | 5.0232 | 2.6164 |  | 4.628 |  |  |  | RSD | 7.0475 | 3.5337 |  | 4.4668 |  |  |  | RSD | 6.4801 | 3.5467 |  | 4.3744
|  | Cer d18:1/16:0 | D7 Cer d18:1/16:0 | D7 Cer d18:1/16:0 | Cer d18:1/16:0 |  |  |  |  | Cer d18:1/18:0 | D7 Cer d18:1/18:0 | D7 Cer d18:1/18:0 | Cer d18:1/18:0 |  |  |  |  | Cer d18:1/24:0 | D7 Cer d18:1/24:0 | D7 Cer d18:1/24:0 | Cer d18:1/24:0 |  |  |  |  | Cer d18:1/24:1 | D7 Cer d18:1/24:1 | D7 Cer d18:1/24:1 | Cer d18:1/24:1
| Sample Name | Analyte Peak Area (counts) | Analyte Peak Area (counts) | c [pmol/µL] | c [pmol/µL] |  |  |  |  | Analyte Peak Area (counts) | Analyte Peak Area (counts) | c [pmol/µL] | c [pmol/µL] |  |  |  | Sample Name | Analyte Peak Area (counts) | Analyte Peak Area (counts) | c [pmol/µL] | c [pmol/µL] |  |  |  | Sample Name | Analyte Peak Area (counts) | Analyte Peak Area (counts) | c [pmol/µL] | c [pmol/µL]
1 | HLQC1 1 | 159000 | 185000 | 0.125 | 0.10743 |  |  |  | HLQC1 1 | 82500 | 190000 | 0.05 | 0.021711 |  |  |  | HLQC1 1 | 2110000 | 3010000 | 1.5 | 1.0515 |  |  |  | HLQC1 1 | 645000 | 828000 | 0.5 | 0.38949
1 | HLQC2 1 | 168000 | 188000 | 0.125 | 0.1117 |  |  |  | HLQC2 1 | 88800 | 197000 | 0.05 | 0.022538 |  |  |  | HLQC2 1 | 2190000 | 3120000 | 1.5 | 1.0529 |  |  |  | HLQC2 1 | 699000 | 841000 | 0.5 | 0.41558
1 | HLQC3 1 | 161000 | 190000 | 0.125 | 0.10592 |  |  |  | HLQC3 1 | 81300 | 191000 | 0.05 | 0.021283 |  |  |  | HLQC3 1 | 1990000 | 2980000 | 1.5 | 1.0017 |  |  |  | HLQC3 1 | 655000 | 839000 | 0.5 | 0.39035
1 | HLQC4 1 | 157000 | 188000 | 0.125 | 0.10439 |  |  |  | HLQC4 1 | 82700 | 187000 | 0.05 | 0.022112 |  |  |  | HLQC4 1 | 2080000 | 3040000 | 1.5 | 1.0263 |  |  |  | HLQC4 1 | 639000 | 858000 | 0.5 | 0.37238
1 | HLQC5 1 | 159000 | 182000 | 0.125 | 0.1092 |  |  |  | HLQC5 1 | 82100 | 188000 | 0.05 | 0.021835 |  |  |  | HLQC5 1 | 2070000 | 3090000 | 1.5 | 1.0049 |  |  |  | HLQC5 1 | 655000 | 853000 | 0.5 | 0.38394
1 | HLQC6 1 | 159000 | 192000 | 0.125 | 0.10352 |  |  |  | HLQC6 1 | 87800 | 197000 | 0.05 | 0.022284 |  |  |  | HLQC6 1 | 2070000 | 3110000 | 1.5 | 0.99839 |  |  |  | HLQC6 1 | 661000 | 852000 | 0.5 | 0.38791
2 | HLQC1 2 | 156000 | 186000 | 0.125 | 0.10484 |  |  |  | HLQC1 2 | 87400 | 194000 | 0.05 | 0.022526 |  |  |  | HLQC1 2 | 2070000 | 3000000 | 1.5 | 1.035 |  |  |  | HLQC1 2 | 668000 | 846000 | 0.5 | 0.3948
2 | HLQC2 2 | 165000 | 184000 | 0.125 | 0.11209 |  |  |  | HLQC2 2 | 85400 | 186000 | 0.05 | 0.022957 |  |  |  | HLQC2 2 | 2190000 | 3150000 | 1.5 | 1.0429 |  |  |  | HLQC2 2 | 683000 | 852000 | 0.5 | 0.40082
2 | HLQC3 2 | 159000 | 190000 | 0.125 | 0.10461 |  |  |  | HLQC3 2 | 87500 | 188000 | 0.05 | 0.023271 |  |  |  | HLQC3 2 | 2090000 | 3040000 | 1.5 | 1.0312 |  |  |  | HLQC3 2 | 642000 | 841000 | 0.5 | 0.38169
2 | HLQC4 2 | 158000 | 187000 | 0.125 | 0.10561 |  |  |  | HLQC4 2 | 79800 | 191000 | 0.05 | 0.02089 |  |  |  | HLQC4 2 | 2070000 | 3130000 | 1.5 | 0.99201 |  |  |  | HLQC4 2 | 636000 | 848000 | 0.5 | 0.375
2 | HLQC5 2 | 158000 | 187000 | 0.125 | 0.10561 |  |  |  | HLQC5 2 | 86000 | 196000 | 0.05 | 0.021939 |  |  |  | HLQC5 2 | 2110000 | 3030000 | 1.5 | 1.0446 |  |  |  | HLQC5 2 | 655000 | 857000 | 0.5 | 0.38215
2 | HLQC6 2 | 155000 | 191000 | 0.125 | 0.10144 |  |  |  | HLQC6 2 | 84600 | 189000 | 0.05 | 0.022381 |  |  |  | HLQC6 2 | 2100000 | 3110000 | 1.5 | 1.0129 |  |  |  | HLQC6 2 | 632000 | 852000 | 0.5 | 0.37089
3 | HLQC1 3 | 153000 | 183000 | 0.125 | 0.10451 |  |  |  | HLQC1 3 | 85000 | 192000 | 0.05 | 0.022135 |  |  |  | HLQC1 3 | 2120000 | 3010000 | 1.5 | 1.0565 |  |  |  | HLQC1 3 | 664000 | 853000 | 0.5 | 0.38921
3 | HLQC2 3 | 165000 | 187000 | 0.125 | 0.11029 |  |  |  | HLQC2 3 | 90600 | 193000 | 0.05 | 0.023472 |  |  |  | HLQC2 3 | 2210000 | 3000000 | 1.5 | 1.105 |  |  |  | HLQC2 3 | 691000 | 857000 | 0.5 | 0.40315
3 | HLQC3 3 | 158000 | 191000 | 0.125 | 0.1034 |  |  |  | HLQC3 3 | 86500 | 193000 | 0.05 | 0.022409 |  |  |  | HLQC3 3 | 2060000 | 3020000 | 1.5 | 1.0232 |  |  |  | HLQC3 3 | 662000 | 847000 | 0.5 | 0.39079
3 | HLQC4 3 | 160000 | 195000 | 0.125 | 0.10256 |  |  |  | HLQC4 3 | 86300 | 195000 | 0.05 | 0.022128 |  |  |  | HLQC4 3 | 2120000 | 3140000 | 1.5 | 1.0127 |  |  |  | HLQC4 3 | 665000 | 870000 | 0.5 | 0.38218
3 | HLQC5 3 | 154000 | 190000 | 0.125 | 0.10132 |  |  |  | HLQC5 3 | 83200 | 191000 | 0.05 | 0.02178 |  |  |  | HLQC5 3 | 2110000 | 3100000 | 1.5 | 1.021 |  |  |  | HLQC5 3 | 651000 | 856000 | 0.5 | 0.38026
3 | HLQC6 3 | 157000 | 193000 | 0.125 | 0.10168 |  |  |  | HLQC6 3 | 83900 | 194000 | 0.05 | 0.021624 |  |  |  | HLQC6 3 | 2080000 | 3080000 | 1.5 | 1.013 |  |  |  | HLQC6 3 | 646000 | 861000 | 0.5 | 0.37515
| Average | 158944.4444 | 188277.7778 |  | 0.10556 |  |  |  | Average | 85077.7778 | 191777.7778 |  | 0.022182 |  |  |  | Average | 2102222.2222 | 3064444.4444 |  | 1.0292 |  |  |  | Average | 658277.7778 | 850611.1111 |  | 0.38699
| SD | 3763.4532 | 3428.7276 |  | 0.0032614 |  |  |  | SD | 2748.0408 | 3275.4229 |  | 0.00063067 |  |  |  | SD | 51376.1245 | 53977.5902 |  | 0.026461 |  |  |  | SD | 17903.462 | 9244.4511 |  | 0.011191
| RSD | 2.3678 | 1.8211 |  | 3.0895 |  |  |  | RSD | 3.23 | 1.7079 |  | 2.8432 |  |  |  | RSD | 2.4439 | 1.7614 |  | 2.571 |  |  |  | RSD | 2.7197 | 1.0868 |  | 2.8919
| NEW IS VIAL | [7 empty] | NEW IS VIAL | [7 empty] | NEW IS VIAL | [7 empty] | NEW IS VIAL
| Extraction | [7 empty] | Extraction | [7 empty] | Extraction | [7 empty] | Extraction
| REPEATED | [7 empty] | REPEATED | [7 empty] | REPEATED | [7 empty] | REPEATED
|  | Cer d18:1/16:0 | D7 Cer d18:1/16:0 | D7 Cer d18:1/16:0 | Cer d18:1/16:0 |  |  |  |  | Cer d18:1/18:0 | D7 Cer d18:1/18:0 | D7 Cer d18:1/18:0 | Cer d18:1/18:0 |  |  |  |  | Cer d18:1/24:0 | D7 Cer d18:1/24:0 | D7 Cer d18:1/24:0 | Cer d18:1/24:0 |  |  |  |  | Cer d18:1/24:1 | D7 Cer d18:1/24:1 | D7 Cer d18:1/24:1 | Cer d18:1/24:1
| Sample Name | Analyte Peak Area (counts) | Analyte Peak Area (counts) | c [pmol/µL] | c [pmol/µL] |  |  |  | Sample Name | Analyte Peak Area (counts) | Analyte Peak Area (counts) | c [pmol/µL] | c [pmol/µL] |  |  |  | Sample Name | Analyte Peak Area (counts) | Analyte Peak Area (counts) | c [pmol/µL] | c [pmol/µL] |  |  |  | Sample Name | Analyte Peak Area (counts) | Analyte Peak Area (counts) | c [pmol/µL] | c [pmol/µL]
1 | LLQC1 1 | 74100 | 194000 | 0.125 | 0.047745 |  |  |  | LLQC1 1 | 32000 | 177000 | 0.05 | 0.0090395 |  |  |  | LLQC1 1 | 957000 | 2910000 | 1.5 | 0.4933 |  |  |  | LLQC1 1 | 247000 | 799000 | 0.5 | 0.15457
1 | LLQC2 1 | 80400 | 195000 | 0.125 | 0.051538 |  |  |  | LLQC2 1 | 33400 | 188000 | 0.05 | 0.008883 |  |  |  | LLQC2 1 | 930000 | 2960000 | 1.5 | 0.47128 |  |  |  | LLQC2 1 | 250000 | 814000 | 0.5 | 0.15356
1 | LLQC3 1 | 82400 | 196000 | 0.125 | 0.052551 |  |  |  | LLQC3 1 | 34000 | 182000 | 0.05 | 0.0093407 |  |  |  | LLQC3 1 | 926000 | 2960000 | 1.5 | 0.46926 |  |  |  | LLQC3 1 | 251000 | 807000 | 0.5 | 0.15551
1 | LLQC4 1 | 81200 | 200000 | 0.125 | 0.05075 |  |  |  | LLQC4 1 | 31600 | 181000 | 0.05 | 0.0087293 |  |  |  | LLQC4 1 | 883000 | 2950000 | 1.5 | 0.44898 |  |  |  | LLQC4 1 | 230000 | 805000 | 0.5 | 0.14286
1 | LLQC5 1 | 81600 | 199000 | 0.125 | 0.051256 |  |  |  | LLQC5 1 | 32000 | 184000 | 0.05 | 0.0086957 |  |  |  | LLQC5 1 | 928000 | 2970000 | 1.5 | 0.46869 |  |  |  | LLQC5 1 | 240000 | 786000 | 0.5 | 0.15267
1 | LLQC6 1 | 80900 | 201000 | 0.125 | 0.050311 |  |  |  | LLQC6 1 | 32300 | 186000 | 0.05 | 0.0086828 |  |  |  | LLQC6 1 | 886000 | 2810000 | 1.5 | 0.47295 |  |  |  | LLQC6 1 | 235000 | 786000 | 0.5 | 0.14949
2 | LLQC1 2 | 78500 | 193000 | 0.125 | 0.050842 |  |  |  | LLQC1 2 | 32500 | 174000 | 0.05 | 0.0093391 |  |  |  | LLQC1 2 | 932000 | 2890000 | 1.5 | 0.48374 |  |  |  | LLQC1 2 | 233000 | 794000 | 0.5 | 0.14673
2 | LLQC2 2 | 81200 | 197000 | 0.125 | 0.051523 |  |  |  | LLQC2 2 | 33200 | 182000 | 0.05 | 0.0091209 |  |  |  | LLQC2 2 | 927000 | 3040000 | 1.5 | 0.4574 |  |  |  | LLQC2 2 | 243000 | 806000 | 0.5 | 0.15074
2 | LLQC3 2 | 82100 | 195000 | 0.125 | 0.052628 |  |  |  | LLQC3 2 | 32900 | 183000 | 0.05 | 0.0089891 |  |  |  | LLQC3 2 | 946000 | 3020000 | 1.5 | 0.46987 |  |  |  | LLQC3 2 | 238000 | 799000 | 0.5 | 0.14894
2 | LLQC4 2 | 75900 | 199000 | 0.125 | 0.047676 |  |  |  | LLQC4 2 | 29800 | 175000 | 0.05 | 0.0085143 |  |  |  | LLQC4 2 | 899000 | 2920000 | 1.5 | 0.46182 |  |  |  | LLQC4 2 | 228000 | 776000 | 0.5 | 0.14691
2 | LLQC5 2 | 80700 | 195000 | 0.125 | 0.051731 |  |  |  | LLQC5 2 | 32900 | 187000 | 0.05 | 0.0087968 |  |  |  | LLQC5 2 | 927000 | 2950000 | 1.5 | 0.47136 |  |  |  | LLQC5 2 | 238000 | 765000 | 0.5 | 0.15556
2 | LLQC6 2 | 78900 | 203000 | 0.125 | 0.048584 |  |  |  | LLQC6 2 | 30300 | 184000 | 0.05 | 0.0082337 |  |  |  | LLQC6 2 | 907000 | 2940000 | 1.5 | 0.46276 |  |  |  | LLQC6 2 | 233000 | 775000 | 0.5 | 0.15032
3 | LLQC1 3 | 79100 | 191000 | 0.125 | 0.051767 |  |  |  | LLQC1 3 | 33200 | 182000 | 0.05 | 0.0091209 |  |  |  | LLQC1 3 | 946000 | 2960000 | 1.5 | 0.47939 |  |  |  | LLQC1 3 | 253000 | 806000 | 0.5 | 0.15695
3 | LLQC2 3 | 80200 | 195000 | 0.125 | 0.05141 |  |  |  | LLQC2 3 | 33300 | 183000 | 0.05 | 0.0090984 |  |  |  | LLQC2 3 | 925000 | 2990000 | 1.5 | 0.46405 |  |  |  | LLQC2 3 | 243000 | 780000 | 0.5 | 0.15577
3 | LLQC3 3 | 78300 | 198000 | 0.125 | 0.049432 |  |  |  | LLQC3 3 | 31800 | 180000 | 0.05 | 0.0088333 |  |  |  | LLQC3 3 | 937000 | 2970000 | 1.5 | 0.47323 |  |  |  | LLQC3 3 | 242000 | 812000 | 0.5 | 0.14901
3 | LLQC4 3 | 78600 | 196000 | 0.125 | 0.050128 |  |  |  | LLQC4 3 | 32400 | 178000 | 0.05 | 0.0091011 |  |  |  | LLQC4 3 | 912000 | 2950000 | 1.5 | 0.46373 |  |  |  | LLQC4 3 | 240000 | 805000 | 0.5 | 0.14907
3 | LLQC5 3 | 79600 | 200000 | 0.125 | 0.04975 |  |  |  | LLQC5 3 | 31500 | 186000 | 0.05 | 0.0084677 |  |  |  | LLQC5 3 | 929000 | 2940000 | 1.5 | 0.47398 |  |  |  | LLQC5 3 | 241000 | 785000 | 0.5 | 0.1535
3 | LLQC6 3 | 81400 | 200000 | 0.125 | 0.050875 |  |  |  | LLQC6 3 | 32100 | 179000 | 0.05 | 0.0089665 |  |  |  | LLQC6 3 | 913000 | 2930000 | 1.5 | 0.46741 |  |  |  | LLQC6 3 | 243000 | 784000 | 0.5 | 0.15497
| Average | 79727.7778 | 197055.5556 |  | 0.050583 |  |  |  | Average | 32288.8889 | 181722.2222 |  | 0.0088863 |  |  |  | Average | 922777.7778 | 2947777.7778 |  | 0.46962 |  |  |  | Average | 240444.4444 | 793555.5556 |  | 0.15151
| SD | 2090.137 | 3045.4378 |  | 0.0014278 |  |  |  | SD | 1036.4946 | 3870.1932 |  | 0.00028923 |  |  |  | SD | 19245.6505 | 48253.4461 |  | 0.0096014 |  |  |  | SD | 6841.233 | 13945.2191 |  | 0.0037627
| RSD | 2.6216 | 1.5455 |  | 2.8226 |  |  |  | RSD | 3.2101 | 2.1297 |  | 3.2549 |  |  |  | RSD | 2.0856 | 1.6369 |  | 2.0445 |  |  |  | RSD | 2.8452 | 1.7573 |  | 2.4835
|  | Cer d18:1/16:0 | D7 Cer d18:1/16:0 | D7 Cer d18:1/16:0 | Cer d18:1/16:0 |  |  |  |  | Cer d18:1/18:0 | D7 Cer d18:1/18:0 | D7 Cer d18:1/18:0 | Cer d18:1/18:0 |  |  |  |  | Cer d18:1/24:0 | D7 Cer d18:1/24:0 | D7 Cer d18:1/24:0 | Cer d18:1/24:0 |  |  |  |  | Cer d18:1/24:1 | D7 Cer d18:1/24:1 | D7 Cer d18:1/24:1 | Cer d18:1/24:1
| Sample Name | Analyte Peak Area (counts) | Analyte Peak Area (counts) | c [pmol/µL] | c [pmol/µL] |  |  |  |  | Analyte Peak Area (counts) | Analyte Peak Area (counts) | c [pmol/µL] | c [pmol/µL] |  |  |  | Sample Name | Analyte Peak Area (counts) | Analyte Peak Area (counts) | c [pmol/µL] | c [pmol/µL] |  |  |  | Sample Name | Analyte Peak Area (counts) | Analyte Peak Area (counts) | c [pmol/µL] | c [pmol/µL]
1 | LQC1 1 | 124000 | 193000 | 0.125 | 0.080311 |  |  |  | LQC1 1 | 56000 | 196000 | 0.05 | 0.014286 |  |  |  | LQC1 1 | 1630000 | 3260000 | 1.5 | 0.75 |  |  |  | LQC1 1 | 416000 | 860000 | 0.5 | 0.24186
1 | LQC2 1 | 125000 | 200000 | 0.125 | 0.078125 |  |  |  | LQC2 1 | 55500 | 202000 | 0.05 | 0.013738 |  |  |  | LQC2 1 | 1720000 | 3300000 | 1.5 | 0.78182 |  |  |  | LQC2 1 | 428000 | 879000 | 0.5 | 0.24346
1 | LQC3 1 | 119000 | 193000 | 0.125 | 0.077073 |  |  |  | LQC3 1 | 54000 | 194000 | 0.05 | 0.013918 |  |  |  | LQC3 1 | 1470000 | 3070000 | 1.5 | 0.71824 |  |  |  | LQC3 1 | 379000 | 803000 | 0.5 | 0.23599
1 | LQC4 1 | 120000 | 201000 | 0.125 | 0.074627 |  |  |  | LQC4 1 | 52700 | 195000 | 0.05 | 0.013513 |  |  |  | LQC4 1 | 1540000 | 3150000 | 1.5 | 0.73333 |  |  |  | LQC4 1 | 389000 | 828000 | 0.5 | 0.2349
1 | LQC5 1 | 116000 | 194000 | 0.125 | 0.074742 |  |  |  | LQC5 1 | 53600 | 190000 | 0.05 | 0.014105 |  |  |  | LQC5 1 | 1460000 | 3060000 | 1.5 | 0.71569 |  |  |  | LQC5 1 | 369000 | 822000 | 0.5 | 0.22445
1 | LQC6 1 | 123000 | 203000 | 0.125 | 0.075739 |  |  |  | LQC6 1 | 52500 | 204000 | 0.05 | 0.012868 |  |  |  | LQC6 1 | 1500000 | 3270000 | 1.5 | 0.68807 |  |  |  | LQC6 1 | 390000 | 860000 | 0.5 | 0.22674
2 | LQC1 2 | 123000 | 194000 | 0.125 | 0.079253 |  |  |  | LQC1 2 | 53500 | 196000 | 0.05 | 0.013648 |  |  |  | LQC1 2 | 1630000 | 3170000 | 1.5 | 0.77129 |  |  |  | LQC1 2 | 407000 | 852000 | 0.5 | 0.23885
2 | LQC2 2 | 132000 | 198000 | 0.125 | 0.083333 |  |  |  | LQC2 2 | 56500 | 204000 | 0.05 | 0.013848 |  |  |  | LQC2 2 | 1710000 | 3340000 | 1.5 | 0.76796 |  |  |  | LQC2 2 | 432000 | 879000 | 0.5 | 0.24573
2 | LQC3 2 | 123000 | 198000 | 0.125 | 0.077652 |  |  |  | LQC3 2 | 52100 | 193000 | 0.05 | 0.013497 |  |  |  | LQC3 2 | 1530000 | 3080000 | 1.5 | 0.74513 |  |  |  | LQC3 2 | 396000 | 807000 | 0.5 | 0.24535
2 | LQC4 2 | 120000 | 199000 | 0.125 | 0.075377 |  |  |  | LQC4 2 | 51800 | 197000 | 0.05 | 0.013147 |  |  |  | LQC4 2 | 1520000 | 3220000 | 1.5 | 0.70807 |  |  |  | LQC4 2 | 381000 | 829000 | 0.5 | 0.22979
2 | LQC5 2 | 119000 | 199000 | 0.125 | 0.074749 |  |  |  | LQC5 2 | 51600 | 199000 | 0.05 | 0.012965 |  |  |  | LQC5 2 | 1500000 | 3140000 | 1.5 | 0.71656 |  |  |  | LQC5 2 | 397000 | 859000 | 0.5 | 0.23108
2 | LQC6 2 | 123000 | 206000 | 0.125 | 0.074636 |  |  |  | LQC6 2 | 54800 | 201000 | 0.05 | 0.013632 |  |  |  | LQC6 2 | 1500000 | 3190000 | 1.5 | 0.70533 |  |  |  | LQC6 2 | 380000 | 861000 | 0.5 | 0.22067
3 | LQC1 3 | 126000 | 190000 | 0.125 | 0.082895 |  |  |  | LQC1 3 | 55500 | 193000 | 0.05 | 0.014378 |  |  |  | LQC1 3 | 1590000 | 3140000 | 1.5 | 0.75955 |  |  |  | LQC1 3 | 421000 | 846000 | 0.5 | 0.24882
3 | LQC2 3 | 129000 | 191000 | 0.125 | 0.084424 |  |  |  | LQC2 3 | 53900 | 199000 | 0.05 | 0.013543 |  |  |  | LQC2 3 | 1680000 | 3300000 | 1.5 | 0.76364 |  |  |  | LQC2 3 | 427000 | 857000 | 0.5 | 0.24912
3 | LQC3 3 | 118000 | 186000 | 0.125 | 0.079301 |  |  |  | LQC3 3 | 48800 | 190000 | 0.05 | 0.012842 |  |  |  | LQC3 3 | 1500000 | 3080000 | 1.5 | 0.73052 |  |  |  | LQC3 3 | 385000 | 831000 | 0.5 | 0.23165
3 | LQC4 3 | 124000 | 198000 | 0.125 | 0.078283 |  |  |  | LQC4 3 | 51800 | 196000 | 0.05 | 0.013214 |  |  |  | LQC4 3 | 1470000 | 3180000 | 1.5 | 0.6934 |  |  |  | LQC4 3 | 386000 | 829000 | 0.5 | 0.23281
3 | LQC5 3 | 122000 | 201000 | 0.125 | 0.075871 |  |  |  | LQC5 3 | 51400 | 201000 | 0.05 | 0.012786 |  |  |  | LQC5 3 | 1460000 | 3160000 | 1.5 | 0.69304 |  |  |  | LQC5 3 | 388000 | 837000 | 0.5 | 0.23178
3 | LQC6 3 | 117000 | 210000 | 0.125 | 0.069643 |  |  |  | LQC6 3 | 52700 | 202000 | 0.05 | 0.013045 |  |  |  | LQC6 3 | 1480000 | 3220000 | 1.5 | 0.68944 |  |  |  | LQC6 3 | 386000 | 853000 | 0.5 | 0.22626
| Average | 122388.8889 | 197444.4444 |  | 0.077557 |  |  |  | Average | 53261.1111 | 197333.3333 |  | 0.013498 |  |  |  | Average | 1549444.4444 | 3185000 |  | 0.7295 |  |  |  | Average | 397611.1111 | 844000 |  | 0.23552
| SD | 3988.0222 | 5727.344 |  | 0.0035752 |  |  |  | SD | 1892.4232 | 4294.6996 |  | 0.0004825 |  |  |  | SD | 85340.6392 | 82613.5582 |  | 0.029995 |  |  |  | SD | 18782.7903 | 21411.3158 |  | 0.0084073
| RSD | 3.2585 | 2.9007 |  | 4.6097 |  |  |  | RSD | 3.5531 | 2.1764 |  | 3.5745 |  |  |  | RSD | 5.5078 | 2.5938 |  | 4.1117 |  |  |  | RSD | 4.7239 | 2.5369 |  | 3.5697
|  | Cer d18:1/16:0 | D7 Cer d18:1/16:0 | D7 Cer d18:1/16:0 | Cer d18:1/16:0 |  |  |  |  | Cer d18:1/18:0 | D7 Cer d18:1/18:0 | D7 Cer d18:1/18:0 | Cer d18:1/18:0 |  |  |  |  | Cer d18:1/24:0 | D7 Cer d18:1/24:0 | D7 Cer d18:1/24:0 | Cer d18:1/24:0 |  |  |  |  | Cer d18:1/24:1 | D7 Cer d18:1/24:1 | D7 Cer d18:1/24:1 | Cer d18:1/24:1
| Sample Name | Analyte Peak Area (counts) | Analyte Peak Area (counts) | c [pmol/µL] | c [pmol/µL] |  |  |  |  | Analyte Peak Area (counts) | Analyte Peak Area (counts) | c [pmol/µL] | c [pmol/µL] |  |  |  | Sample Name | Analyte Peak Area (counts) | Analyte Peak Area (counts) | c [pmol/µL] | c [pmol/µL] |  |  |  | Sample Name | Analyte Peak Area (counts) | Analyte Peak Area (counts) | c [pmol/µL] | c [pmol/µL]
1 | MQC1 1 | 207000 | 177000 | 0.125 | 0.14619 |  |  |  | MQC1 1 | 119000 | 224000 | 0.05 | 0.026562 |  |  |  | MQC1 1 | 3580000 | 3820000 | 1.5 | 1.4058 |  |  |  | MQC1 1 | 914000 | 1020000 | 0.5 | 0.44804
1 | MQC2 1 | 207000 | 183000 | 0.125 | 0.14139 |  |  |  | MQC2 1 | 114000 | 229000 | 0.05 | 0.024891 |  |  |  | MQC2 1 | 3620000 | 3780000 | 1.5 | 1.4365 |  |  |  | MQC2 1 | 961000 | 1040000 | 0.5 | 0.46202
1 | MQC3 1 | 204000 | 176000 | 0.125 | 0.14489 |  |  |  | MQC3 1 | 106000 | 215000 | 0.05 | 0.024651 |  |  |  | MQC3 1 | 3210000 | 3620000 | 1.5 | 1.3301 |  |  |  | MQC3 1 | 848000 | 957000 | 0.5 | 0.44305
1 | MQC4 1 | 203000 | 186000 | 0.125 | 0.13642 |  |  |  | MQC4 1 | 111000 | 226000 | 0.05 | 0.024558 |  |  |  | MQC4 1 | 3390000 | 3890000 | 1.5 | 1.3072 |  |  |  | MQC4 1 | 878000 | 1020000 | 0.5 | 0.43039
1 | MQC5 1 | 218000 | 178000 | 0.125 | 0.15309 |  |  |  | MQC5 1 | 121000 | 231000 | 0.05 | 0.02619 |  |  |  | MQC5 1 | 3750000 | 3940000 | 1.5 | 1.4277 |  |  |  | MQC5 1 | 976000 | 1040000 | 0.5 | 0.46923
1 | MQC6 1 | 212000 | 184000 | 0.125 | 0.14402 |  |  |  | MQC6 1 | 110000 | 224000 | 0.05 | 0.024554 |  |  |  | MQC6 1 | 3330000 | 3730000 | 1.5 | 1.3391 |  |  |  | MQC6 1 | 854000 | 1000000 | 0.5 | 0.427
2 | MQC1 2 | 214000 | 182000 | 0.125 | 0.14698 |  |  |  | MQC1 2 | 115000 | 222000 | 0.05 | 0.025901 |  |  |  | MQC1 2 | 3550000 | 3840000 | 1.5 | 1.3867 |  |  |  | MQC1 2 | 933000 | 1040000 | 0.5 | 0.44856
2 | MQC2 2 | 211000 | 177000 | 0.125 | 0.14901 |  |  |  | MQC2 2 | 112000 | 229000 | 0.05 | 0.024454 |  |  |  | MQC2 2 | 3550000 | 3840000 | 1.5 | 1.3867 |  |  |  | MQC2 2 | 941000 | 1030000 | 0.5 | 0.4568
2 | MQC3 2 | 202000 | 185000 | 0.125 | 0.13649 |  |  |  | MQC3 2 | 109000 | 219000 | 0.05 | 0.024886 |  |  |  | MQC3 2 | 3250000 | 3640000 | 1.5 | 1.3393 |  |  |  | MQC3 2 | 855000 | 964000 | 0.5 | 0.44346
2 | MQC4 2 | 207000 | 186000 | 0.125 | 0.13911 |  |  |  | MQC4 2 | 110000 | 234000 | 0.05 | 0.023504 |  |  |  | MQC4 2 | 3490000 | 3910000 | 1.5 | 1.3389 |  |  |  | MQC4 2 | 925000 | 1030000 | 0.5 | 0.44903
2 | MQC5 2 | 215000 | 178000 | 0.125 | 0.15098 |  |  |  | MQC5 2 | 123000 | 226000 | 0.05 | 0.027212 |  |  |  | MQC5 2 | 3720000 | 3820000 | 1.5 | 1.4607 |  |  |  | MQC5 2 | 965000 | 1020000 | 0.5 | 0.47304
2 | MQC6 2 | 203000 | 181000 | 0.125 | 0.14019 |  |  |  | MQC6 2 | 112000 | 226000 | 0.05 | 0.024779 |  |  |  | MQC6 2 | 3310000 | 3730000 | 1.5 | 1.3311 |  |  |  | MQC6 2 | 878000 | 980000 | 0.5 | 0.44796
3 | MQC1 3 | 206000 | 181000 | 0.125 | 0.14227 |  |  |  | MQC1 3 | 113000 | 224000 | 0.05 | 0.025223 |  |  |  | MQC1 3 | 3430000 | 3780000 | 1.5 | 1.3611 |  |  |  | MQC1 3 | 931000 | 999000 | 0.5 | 0.46597
3 | MQC2 3 | 214000 | 182000 | 0.125 | 0.14698 |  |  |  | MQC2 3 | 114000 | 229000 | 0.05 | 0.024891 |  |  |  | MQC2 3 | 3600000 | 3870000 | 1.5 | 1.3953 |  |  |  | MQC2 3 | 935000 | 1030000 | 0.5 | 0.45388
3 | MQC3 3 | 202000 | 176000 | 0.125 | 0.14347 |  |  |  | MQC3 3 | 107000 | 217000 | 0.05 | 0.024654 |  |  |  | MQC3 3 | 3270000 | 3770000 | 1.5 | 1.3011 |  |  |  | MQC3 3 | 859000 | 997000 | 0.5 | 0.43079
3 | MQC4 3 | 213000 | 188000 | 0.125 | 0.14162 |  |  |  | MQC4 3 | 111000 | 234000 | 0.05 | 0.023718 |  |  |  | MQC4 3 | 3430000 | 3870000 | 1.5 | 1.3295 |  |  |  | MQC4 3 | 930000 | 1060000 | 0.5 | 0.43868
3 | MQC5 3 | 222000 | 178000 | 0.125 | 0.1559 |  |  |  | MQC5 3 | 119000 | 240000 | 0.05 | 0.024792 |  |  |  | MQC5 3 | 3790000 | 3770000 | 1.5 | 1.508 |  |  |  | MQC5 3 | 947000 | 1060000 | 0.5 | 0.4467
3 | MQC6 3 | 208000 | 181000 | 0.125 | 0.14365 |  |  |  | MQC6 3 | 113000 | 230000 | 0.05 | 0.024565 |  |  |  | MQC6 3 | 3290000 | 3730000 | 1.5 | 1.3231 |  |  |  | MQC6 3 | 858000 | 992000 | 0.5 | 0.43246
| Average | 209333.3333 | 181055.5556 |  | 0.14459 |  |  |  | Average | 113277.7778 | 226611.1111 |  | 0.024999 |  |  |  | Average | 3475555.5556 | 3797222.2222 |  | 1.3727 |  |  |  | Average | 910444.4444 | 1015500 |  | 0.44817
| SD | 5676.4621 | 3643.445 |  | 0.005181 |  |  |  | SD | 4531.4402 | 6074.8927 |  | 0.00090517 |  |  |  | SD | 174903.8537 | 84577.7457 |  | 0.055757 |  |  |  | SD | 42099.3827 | 29054.0684 |  | 0.013191
| RSD | 2.7117 | 2.0123 |  | 3.5832 |  |  |  | RSD | 4.0003 | 2.6808 |  | 3.6208 |  |  |  | RSD | 5.0324 | 2.2274 |  | 4.062 |  |  |  | RSD | 4.624 | 2.8611 |  | 2.9433
|  | Cer d18:1/16:0 | D7 Cer d18:1/16:0 | D7 Cer d18:1/16:0 | Cer d18:1/16:0 |  |  |  |  | Cer d18:1/18:0 | D7 Cer d18:1/18:0 | D7 Cer d18:1/18:0 | Cer d18:1/18:0 |  |  |  |  | Cer d18:1/24:0 | D7 Cer d18:1/24:0 | D7 Cer d18:1/24:0 | Cer d18:1/24:0 |  |  |  |  | Cer d18:1/24:1 | D7 Cer d18:1/24:1 | D7 Cer d18:1/24:1 | Cer d18:1/24:1
| Sample Name | Analyte Peak Area (counts) | Analyte Peak Area (counts) | c [pmol/µL] | c [pmol/µL] |  |  |  |  | Analyte Peak Area (counts) | Analyte Peak Area (counts) | c [pmol/µL] | c [pmol/µL] |  |  |  | Sample Name | Analyte Peak Area (counts) | Analyte Peak Area (counts) | c [pmol/µL] | c [pmol/µL] |  |  |  | Sample Name | Analyte Peak Area (counts) | Analyte Peak Area (counts) | c [pmol/µL] | c [pmol/µL]
1 | HQC1 1 | 308000 | 177000 | 0.125 | 0.21751 |  |  |  | HQC1 1 | 260000 | 234000 | 0.05 | 0.055556 |  |  |  | HQC1 1 | 5100000 | 3930000 | 1.5 | 1.9466 |  |  |  | HQC1 1 | 2040000 | 1030000 | 0.5 | 0.99029
1 | HQC2 1 | 313000 | 181000 | 0.125 | 0.21616 |  |  |  | HQC2 1 | 267000 | 231000 | 0.05 | 0.057792 |  |  |  | HQC2 1 | 5160000 | 3740000 | 1.5 | 2.0695 |  |  |  | HQC2 1 | 2050000 | 1050000 | 0.5 | 0.97619
1 | HQC3 1 | 311000 | 179000 | 0.125 | 0.21718 |  |  |  | HQC3 1 | 264000 | 225000 | 0.05 | 0.058667 |  |  |  | HQC3 1 | 5110000 | 3710000 | 1.5 | 2.066 |  |  |  | HQC3 1 | 2090000 | 994000 | 0.5 | 1.0513
1 | HQC4 1 | 321000 | 182000 | 0.125 | 0.22047 |  |  |  | HQC4 1 | 291000 | 234000 | 0.05 | 0.062179 |  |  |  | HQC4 1 | 5230000 | 3900000 | 1.5 | 2.0115 |  |  |  | HQC4 1 | 2090000 | 1040000 | 0.5 | 1.0048
1 | HQC5 1 | 320000 | 183000 | 0.125 | 0.21858 |  |  |  | HQC5 1 | 270000 | 232000 | 0.05 | 0.05819 |  |  |  | HQC5 1 | 5120000 | 3810000 | 1.5 | 2.0157 |  |  |  | HQC5 1 | 2100000 | 1000000 | 0.5 | 1.05
1 | HQC6 1 | 321000 | 181000 | 0.125 | 0.22169 |  |  |  | HQC6 1 | 284000 | 239000 | 0.05 | 0.059414 |  |  |  | HQC6 1 | 5280000 | 3890000 | 1.5 | 2.036 |  |  |  | HQC6 1 | 2130000 | 1020000 | 0.5 | 1.0441
2 | HQC1 2 | 306000 | 182000 | 0.125 | 0.21016 |  |  |  | HQC1 2 | 262000 | 234000 | 0.05 | 0.055983 |  |  |  | HQC1 2 | 5100000 | 3850000 | 1.5 | 1.987 |  |  |  | HQC1 2 | 2040000 | 1050000 | 0.5 | 0.97143
2 | HQC2 2 | 324000 | 178000 | 0.125 | 0.22753 |  |  |  | HQC2 2 | 264000 | 232000 | 0.05 | 0.056897 |  |  |  | HQC2 2 | 5150000 | 3810000 | 1.5 | 2.0276 |  |  |  | HQC2 2 | 2080000 | 1010000 | 0.5 | 1.0297
2 | HQC3 2 | 316000 | 182000 | 0.125 | 0.21703 |  |  |  | HQC3 2 | 267000 | 223000 | 0.05 | 0.059865 |  |  |  | HQC3 2 | 5190000 | 3710000 | 1.5 | 2.0984 |  |  |  | HQC3 2 | 2090000 | 1010000 | 0.5 | 1.0347
2 | HQC4 2 | 322000 | 181000 | 0.125 | 0.22238 |  |  |  | HQC4 2 | 279000 | 246000 | 0.05 | 0.056707 |  |  |  | HQC4 2 | 5240000 | 3860000 | 1.5 | 2.0363 |  |  |  | HQC4 2 | 2120000 | 1060000 | 0.5 | 1
2 | HQC5 2 | 315000 | 185000 | 0.125 | 0.21284 |  |  |  | HQC5 2 | 263000 | 236000 | 0.05 | 0.05572 |  |  |  | HQC5 2 | 4900000 | 3720000 | 1.5 | 1.9758 |  |  |  | HQC5 2 | 2040000 | 992000 | 0.5 | 1.0282
2 | HQC6 2 | 320000 | 184000 | 0.125 | 0.21739 |  |  |  | HQC6 2 | 275000 | 234000 | 0.05 | 0.058761 |  |  |  | HQC6 2 | 5160000 | 3820000 | 1.5 | 2.0262 |  |  |  | HQC6 2 | 2070000 | 1010000 | 0.5 | 1.0248
3 | HQC1 3 | 304000 | 184000 | 0.125 | 0.20652 |  |  |  | HQC1 3 | 269000 | 233000 | 0.05 | 0.057725 |  |  |  | HQC1 3 | 5160000 | 3770000 | 1.5 | 2.0531 |  |  |  | HQC1 3 | 2070000 | 1010000 | 0.5 | 1.0248
3 | HQC2 3 | 310000 | 179000 | 0.125 | 0.21648 |  |  |  | HQC2 3 | 259000 | 226000 | 0.05 | 0.057301 |  |  |  | HQC2 3 | 5100000 | 3860000 | 1.5 | 1.9819 |  |  |  | HQC2 3 | 2070000 | 1010000 | 0.5 | 1.0248
3 | HQC3 3 | 317000 | 184000 | 0.125 | 0.21535 |  |  |  | HQC3 3 | 265000 | 231000 | 0.05 | 0.057359 |  |  |  | HQC3 3 | 5240000 | 3820000 | 1.5 | 2.0576 |  |  |  | HQC3 3 | 2030000 | 1010000 | 0.5 | 1.005
3 | HQC4 3 | 320000 | 185000 | 0.125 | 0.21622 |  |  |  | HQC4 3 | 278000 | 249000 | 0.05 | 0.055823 |  |  |  | HQC4 3 | 5320000 | 4000000 | 1.5 | 1.995 |  |  |  | HQC4 3 | 2160000 | 1070000 | 0.5 | 1.0093
3 | HQC5 3 | 306000 | 184000 | 0.125 | 0.20788 |  |  |  | HQC5 3 | 270000 | 234000 | 0.05 | 0.057692 |  |  |  | HQC5 3 | 5010000 | 3670000 | 1.5 | 2.0477 |  |  |  | HQC5 3 | 2040000 | 988000 | 0.5 | 1.0324
3 | HQC6 3 | 313000 | 181000 | 0.125 | 0.21616 |  |  |  | HQC6 3 | 275000 | 234000 | 0.05 | 0.058761 |  |  |  | HQC6 3 | 5060000 | 3840000 | 1.5 | 1.9766 |  |  |  | HQC6 3 | 2030000 | 952000 | 0.5 | 1.0662
| Average | 314833.3333 | 181777.7778 |  | 0.21653 |  |  |  | Average | 270111.1111 | 233722.2222 |  | 0.0578 |  |  |  | Average | 5146111.1111 | 3817222.2222 |  | 2.0227 |  |  |  | Average | 2074444.4444 | 1017000 |  | 1.0204
| SD | 6066.7582 | 2322.7272 |  | 0.0049141 |  |  |  | SD | 8438.5945 | 6189.6469 |  | 0.0016303 |  |  |  | SD | 97247.6157 | 83653.0902 |  | 0.038703 |  |  |  | SD | 35935.4703 | 28231.1884 |  | 0.025078
| RSD | 1.927 | 1.2778 |  | 2.2695 |  |  |  | RSD | 3.1241 | 2.6483 |  | 2.8206 |  |  |  | RSD | 1.8897 | 2.1915 |  | 1.9134 |  |  |  | RSD | 1.7323 | 2.7759 |  | 2.4576
|  | Cer d18:1/16:0 | D7 Cer d18:1/16:0 | D7 Cer d18:1/16:0 | Cer d18:1/16:0 |  |  |  |  | Cer d18:1/18:0 | D7 Cer d18:1/18:0 | D7 Cer d18:1/18:0 | Cer d18:1/18:0 |  |  |  |  | Cer d18:1/24:0 | D7 Cer d18:1/24:0 | D7 Cer d18:1/24:0 | Cer d18:1/24:0 |  |  |  |  | Cer d18:1/24:1 | D7 Cer d18:1/24:1 | D7 Cer d18:1/24:1 | Cer d18:1/24:1
| Sample Name | Analyte Peak Area (counts) | Analyte Peak Area (counts) | c [pmol/µL] | c [pmol/µL] |  |  |  |  | Analyte Peak Area (counts) | Analyte Peak Area (counts) | c [pmol/µL] | c [pmol/µL] |  |  |  | Sample Name | Analyte Peak Area (counts) | Analyte Peak Area (counts) | c [pmol/µL] | c [pmol/µL] |  |  |  | Sample Name | Analyte Peak Area (counts) | Analyte Peak Area (counts) | c [pmol/µL] | c [pmol/µL]
1 | HLQC1 1 | 411000 | 189000 | 0.125 | 0.27183 |  |  |  | HLQC1 1 | 407000 | 233000 | 0.05 | 0.087339 |  |  |  | HLQC1 1 | 6460000 | 3840000 | 1.5 | 2.5234 |  |  |  | HLQC1 1 | 3110000 | 997000 | 0.5 | 1.5597
1 | HLQC2 1 | 392000 | 181000 | 0.125 | 0.27072 |  |  |  | HLQC2 1 | 408000 | 232000 | 0.05 | 0.087931 |  |  |  | HLQC2 1 | 6060000 | 3690000 | 1.5 | 2.4634 |  |  |  | HLQC2 1 | 3020000 | 970000 | 0.5 | 1.5567
1 | HLQC3 1 | 408000 | 187000 | 0.125 | 0.27273 |  |  |  | HLQC3 1 | 415000 | 242000 | 0.05 | 0.085744 |  |  |  | HLQC3 1 | 6350000 | 3770000 | 1.5 | 2.5265 |  |  |  | HLQC3 1 | 3100000 | 1020000 | 0.5 | 1.5196
1 | HLQC4 1 | 418000 | 188000 | 0.125 | 0.27793 |  |  |  | HLQC4 1 | 460000 | 256000 | 0.05 | 0.089844 |  |  |  | HLQC4 1 | 6650000 | 3950000 | 1.5 | 2.5253 |  |  |  | HLQC4 1 | 3200000 | 1070000 | 0.5 | 1.4953
1 | HLQC5 1 | 418000 | 193000 | 0.125 | 0.27073 |  |  |  | HLQC5 1 | 437000 | 237000 | 0.05 | 0.092194 |  |  |  | HLQC5 1 | 6430000 | 3840000 | 1.5 | 2.5117 |  |  |  | HLQC5 1 | 3210000 | 1040000 | 0.5 | 1.5433
1 | HLQC6 1 | 418000 | 191000 | 0.125 | 0.27356 |  |  |  | HLQC6 1 | 430000 | 250000 | 0.05 | 0.086 |  |  |  | HLQC6 1 | 6610000 | 3850000 | 1.5 | 2.5753 |  |  |  | HLQC6 1 | 3240000 | 1040000 | 0.5 | 1.5577
2 | HLQC1 2 | 411000 | 188000 | 0.125 | 0.27327 |  |  |  | HLQC1 2 | 414000 | 238000 | 0.05 | 0.086975 |  |  |  | HLQC1 2 | 6270000 | 3740000 | 1.5 | 2.5147 |  |  |  | HLQC1 2 | 3050000 | 995000 | 0.5 | 1.5327
2 | HLQC2 2 | 403000 | 184000 | 0.125 | 0.27378 |  |  |  | HLQC2 2 | 418000 | 235000 | 0.05 | 0.088936 |  |  |  | HLQC2 2 | 6350000 | 3690000 | 1.5 | 2.5813 |  |  |  | HLQC2 2 | 3150000 | 1040000 | 0.5 | 1.5144
2 | HLQC3 2 | 405000 | 183000 | 0.125 | 0.27664 |  |  |  | HLQC3 2 | 407000 | 240000 | 0.05 | 0.084792 |  |  |  | HLQC3 2 | 6400000 | 3670000 | 1.5 | 2.6158 |  |  |  | HLQC3 2 | 3060000 | 975000 | 0.5 | 1.5692
2 | HLQC4 2 | 405000 | 187000 | 0.125 | 0.27072 |  |  |  | HLQC4 2 | 428000 | 246000 | 0.05 | 0.086992 |  |  |  | HLQC4 2 | 6330000 | 3730000 | 1.5 | 2.5456 |  |  |  | HLQC4 2 | 3130000 | 1010000 | 0.5 | 1.5495
2 | HLQC5 2 | 419000 | 193000 | 0.125 | 0.27137 |  |  |  | HLQC5 2 | 441000 | 244000 | 0.05 | 0.090369 |  |  |  | HLQC5 2 | 6440000 | 3840000 | 1.5 | 2.5156 |  |  |  | HLQC5 2 | 3080000 | 1000000 | 0.5 | 1.54
2 | HLQC6 2 | 421000 | 193000 | 0.125 | 0.27267 |  |  |  | HLQC6 2 | 424000 | 250000 | 0.05 | 0.0848 |  |  |  | HLQC6 2 | 6490000 | 3780000 | 1.5 | 2.5754 |  |  |  | HLQC6 2 | 3220000 | 1010000 | 0.5 | 1.5941
3 | HLQC1 3 | 411000 | 191000 | 0.125 | 0.26898 |  |  |  | HLQC1 3 | 413000 | 247000 | 0.05 | 0.083603 |  |  |  | HLQC1 3 | 6530000 | 3740000 | 1.5 | 2.619 |  |  |  | HLQC1 3 | 3190000 | 1040000 | 0.5 | 1.5337
3 | HLQC2 3 | 402000 | 186000 | 0.125 | 0.27016 |  |  |  | HLQC2 3 | 409000 | 236000 | 0.05 | 0.086653 |  |  |  | HLQC2 3 | 6400000 | 3720000 | 1.5 | 2.5806 |  |  |  | HLQC2 3 | 3060000 | 1020000 | 0.5 | 1.5
3 | HLQC3 3 | 412000 | 186000 | 0.125 | 0.27688 |  |  |  | HLQC3 3 | 424000 | 241000 | 0.05 | 0.087967 |  |  |  | HLQC3 3 | 6340000 | 3700000 | 1.5 | 2.5703 |  |  |  | HLQC3 3 | 3080000 | 1030000 | 0.5 | 1.4951
3 | HLQC4 3 | 414000 | 185000 | 0.125 | 0.27973 |  |  |  | HLQC4 3 | 431000 | 247000 | 0.05 | 0.087247 |  |  |  | HLQC4 3 | 6460000 | 3750000 | 1.5 | 2.584 |  |  |  | HLQC4 3 | 3140000 | 1020000 | 0.5 | 1.5392
3 | HLQC5 3 | 425000 | 190000 | 0.125 | 0.27961 |  |  |  | HLQC5 3 | 431000 | 246000 | 0.05 | 0.087602 |  |  |  | HLQC5 3 | 6470000 | 3740000 | 1.5 | 2.5949 |  |  |  | HLQC5 3 | 3230000 | 1020000 | 0.5 | 1.5833
3 | HLQC6 3 | 413000 | 192000 | 0.125 | 0.26888 |  |  |  | HLQC6 3 | 442000 | 247000 | 0.05 | 0.089474 |  |  |  | HLQC6 3 | 6430000 | 3820000 | 1.5 | 2.5249 |  |  |  | HLQC6 3 | 3180000 | 995000 | 0.5 | 1.598
| Average | 411444.4444 | 188166.6667 |  | 0.27334 |  |  |  | Average | 424388.8889 | 242611.1111 |  | 0.08747 |  |  |  | Average | 6415000 | 3770000 |  | 2.5527 |  |  |  | Average | 3136111.1111 | 1016222.2222 |  | 1.5434
| SD | 7833.1363 | 3515.8372 |  | 0.0033445 |  |  |  | SD | 14185.8258 | 6412.998 |  | 0.0020897 |  |  |  | SD | 126721.4794 | 71102.43 |  | 0.040592 |  |  |  | SD | 67506.2869 | 24548.2644 |  | 0.030423
| RSD | 1.9038 | 1.8685 |  | 1.2236 |  |  |  | RSD | 3.3426 | 2.6433 |  | 2.389 |  |  |  | RSD | 1.9754 | 1.886 |  | 1.5902 |  |  |  | RSD | 2.1525 | 2.4156 |  | 1.9711
| LLQC | 0.050583 | [6 empty] | LLQC | 0.0088863 | [6 empty] | LLQC | 0.46962 | [6 empty] | LLQC | 0.15151
| LQC | 0.077557 | [6 empty] | LQC | 0.013498 | [6 empty] | LQC | 0.7295 | [6 empty] | LQC | 0.23552
| MQC | 0.14459 | [6 empty] | MQC | 0.024999 | [6 empty] | MQC | 1.3727 | [6 empty] | MQC | 0.44817
| HQC | 0.21653 | [6 empty] | HQC | 0.0578 | [6 empty] | HQC | 2.0227 | [6 empty] | HQC | 1.0204
| HLQC | 0.27334 | [6 empty] | HLQC | 0.08747 | [6 empty] | HLQC | 2.5527 | [6 empty] | HLQC | 1.5434

## Samples_Extr.plan
Pooled sample (6 samples in the end):
Prepare mixture of:
· 100 µL NIST SRM 1950
· 100 µL NIST high TAG
· 100 µL NIST T1D
· 100 µL NIST young AA
Prepare 6 times (QC 1, QC 2, QC 3, QC 4, QC 5, QC 6 - For measurements, I prepared a pooled sample of QC 1-6 and used this.)
· 10 µL Pooled sample + 20 µL IS d7 + 570 µL 2:8 EtOAc/IPA
Note: I would suggest using 120 µL of each NIST plasma in order to have some reserve.
Calibration curve (3 x each level extracted= 18 samples):
Stocks:
Std 1 (provided non-labeled Ceramide Mix)
Std 2: 250 µL Std 1 + 250 µL 2:8 EtOAc/IPA |  | End volume: 450 µL
Std 3: 50 µL Std 2 + 450 µL 2:8 EtOAc/IPA |  | End volume: 400 µL
Std 4: 100 µL Std 3 + 400 µL 2:8 EtOAc/IPA |  | End volume: 250 µL
Std 5: 250 µL Std 4 + 250 µL 2:8 EtOAc/IPA |  | End volume: 100 µL
Std 6: 400 µL Std 5 + 100 µL 2:8 EtOAc/IPA |  | End volume: 500 µL
Prepare 3 calibration lines (3 x 6):
· 10 µL 5% BSA + 20 µL Std(1-6) + 20 µL IS d7 + 550 µL 2:8 EtOAc/IPA
Matrix blank (3 samples):
Prepare 3 times:
· 10 µL Pooled sample + 590 µL 2:8 EtOAc/IPA
Total blank ( 3 samples):
Prepare 3 times:
· 10 µL 5% BSA + 590 µL 2:8 EtOAc/IPA
Samples (24 samples):
6 x NIST SRM 1950 (on the vial normal plasma is written)
6 x NIST high TAG
6 x NIST T1D
6 x NIST young AA
· 10 µL sample + 20 µL IS d7 + 570 µL 2:8 EtOAc/IPA
Intra QC assay (30 samples):
MQC = Pooled sample (100 µL 1950, 100 µL TAG, 100 µL T1D, 100 µLyoung AA prepared at the beginning)
Stocks:
LLQC: 30 µL pooled sample + 60 µL H2O
LQC: 40 µL pooled sample + 40 µL H2O
Prepare 6 times for each level:
· LLQC: | 10 µL LLQC + 20 µL IS d7 + 570 µL 2:8 EtOAc/IPA
· LQC: | 10 µL LQC + 20 µL IS d7 + 570 µL 2:8 EtOAc/IPA
· MQC: | 10 µL MQC + 20 µL IS d7 + 570 µL 2:8 EtOAc/IPA
· HQC: | 10 µL MQC + 10 µL Std3 + 20 µL IS d7 + 560 µL 2:8 EtOAc/IPA
· HLQC: | 10 µL MQC + 20 µL Std3 + 20 µL IS d7 + 550 µL 2:8 EtOAc/IPA