## Participant Information
Participant Information
Participant Information |  | Group Name
|  | Lab code | assigned by coordinators
| Institution | Institution name
Street
City
Postal code
Country
| Contact person | Name
Address
Email
Phone

## Summary
Summary
Summary | This protocol describes the extraction and LC-MRM procedure for absolute quantification of Cer d18:1/16:0, Cer d18:1/18:0, Cer d18:1/24:0 and Cer d18:1/24:1 in human plasma according to Kauhanen et al, Anal Bioanal Chem (2016) 408:3475–3483.
A calibration curve is built by injecting several dilutions of the non-labelled standards mixed with a fixed concentration of labelled standards. Each calibration point is prepared and analysed in triplicates (n=3).
To establish the endogenous concentration of the four selected ceramides in the study samples (NIST SRM1950, NIST high TAG, NIST T1D, and NIST young A-A) each sample is extracted and analyzed in six replicates (n =6).
Prepare the pooled QC by mixing 100 mL each of NIST SRM1950, NIST high TAG, NIST T1D, and NIST young AA (total 400 mL).
We suggest proceeding first to measure the ceramides concentration in the four study samples and then continue with the intra-assay validation to avoid wasting material.

## Extraction - Overview
Extraction Overview
Extraction Overview | Table 1 - Extraction - Materials
Chemicals, solvents and materials
Ethyl acetate (EtAc)
2-propanol (IPA)
5% Bovine Serum Albumin (BSA)
96-well plates/tubes
Sealing foil
2 mL DeepWell plates/tubes/LC vials
Table 2 - Extraction - Sample Types
| Sample type | Components
| Pooled Quality Control samples (QC) | 10 µL pooled QC |  |  | 20 µL Labelled IS |  | 570 µL EtAc:IPA 2:8 (v/v)
| Matrix Blank (MB) | 10 µL Matrix (pooled QC) only |  |  | - |  | 590 µL EtAc:IPA 2:8
| Study Samples (S) | 10 µL Human plasma (NIST SRM1950, high TAG, T1D, young AA) |  |  | 20 µL Labelled IS |  | 570 µL EtAc:IPA 2:8
| Total Blank (TB) | 10 µL 5% BSA |  |  | - |  | 590 µL EtAc:IPA 2:8
| Calibration line standards (STD) | 10 µL 5% BSA + 20 µL non labelled STD mixture (Table 5) |  |  | 20 µL Labelled IS |  | 550 µL EtAc:IPA 2:8
Table 3 - Extraction - Labelled Internal Standard Mixture Solution (provided)
| Labelled IS mixture in EtAc:IPA 2:8 (v/v) | Concentration |  |  |  |  | Unit | Vendor
| D7-Cer d18:1/16:0 | 0.125 |  |  |  |  | pmol/µL [µM] | Avanti Polar Lipids
| D7-Cer d18:1/18:0 | 0.050
| D7-Cer d18:1/24:0 | 1500
| D7-Cer d18:1/24:1 | 0.500
Table 4 - Extraction - Non-Labelled Standards - for STD1 Mixture Solution (provided)
| Non-labelled standards in EtAc:IPA 2:8 (v/v) | Concentration |  |  |  |  | Unit | Vendor
| Cer d18:1/16:0 | 2.0 |  |  |  |  | pmol/µL [µM] | Avanti Polar Lipids
| Cer d18:1/18:0 | 2.0
| Cer d18:1/24:0 | 20.0
| Cer d18:1/24:1 | 20.0
Table 5 - Serial dilutions of Non-Labelled Standards from STD1 Mixture Solution for calibration line
| Add 20 µL of | STD1 Conc. | STD2 Conc. | STD3 Conc. | STD4 Conc. | STD5 Conc. | STD6 Conc. | Unit
| Cer d18:1/16:0 | 2.0 | 1.0 | 0.10 | 0.02 | 0.01 | 0.008 | pmol/µL [µM]
| Cer d18:1/18:0 | 2.0 | 1.0 | 0.10 | 0.02 | 0.01 | 0.008
| Cer d18:1/24:0 | 20.0 | 10.0 | 1.0 | 0.2 | 0.1 | 0.08
| Cer d18:1/24:1 | 20.0 | 10.0 | 1.0 | 0.2 | 0.1 | 0.08

## Extraction - Procedure
Extraction Procedure
Extraction Procedure | Extraction Procedure
| Step | Description | Reference Table
| 1 | Prepare different dilutions of STD1 by diluting the provided mixture (Extraction - Overview, Table 4) with EtAc:IPA 2:8 as suggested in table | Extraction - Overview'!B34
| 2 | Thaw plasma on ice (+4°C) and bring to room temperature prior to extraction. Mix by pipetting before use.
| 3 | Prepare samples in plates/tubes manually or with a robot, according to specifications in column C of table | Extraction - Overview'!C14
| 4 | Add 20 µL of labelled internal standard mixture to samples according to column F of table | Extraction - Overview'!F14
| 5 | For the calibration line, add 20 µL of different dilutions of the non-labelled mixture (STD1-6).
| 6 | Add 590 µL of ethyl acetate:isopropanol (2:8, vol/vol) to Total Blank (TB) and Matrix Blank (MB); 550 µL to calibration line (STD); 570 µL to study samples and QCs according to column H of table | Extraction - Overview'!H14
| 7 | Mix samples 10 min by pipetting or vortexing
| 8 | Centrifuge samples for 10 min at 3000 x gat room temperature
| 9 | Transfer 50 µL of the clear supernatant to MS-vials or MS plates, close/seal, store samples at -20 °C prior to LC-MRM analysis.

## LC-MRM - Overview
LC-MRM Overview
LC-MRM Overview | Table 6 Chemicals and solvents
Ammonium formate
Formic acid
Ultra-pure water
Acetonitrile
Table 7 LC-MRM LC-MS Details
| Property | Specifications
| Mass spectrometry | 8060 triple quadrupole mass spectrometer Shimadzu
| LC | Nexera X2 UHPLC system
| Injection volume (indicative) | 5 µL
| Weak wash | Water
| Strong wash | Acetonitrile
| Pre-column | Waters Acquity BEH C18, 1.7 µm VanGuard Pre-Column
| Column | Waters Acquity BEH C18, 2.1 × 50 mm id. 1.7 µm
| Column temperature | 60 °C
| Mobile phase A | 10 mM ammonium acetate in water with 0.1% formic acid
| Mobile phase B | 10 mM ammonium acetate in acetonitrile:2-propanol (4:3, vol/vol) with 0.1% formic acid
Table 8 LC-MRM LC gradient
| Time (minutes) | A% | B% |  | Flow (µL/min)
| 0.00 | 15 | 85 |  | 500
| 0.50 | 15 | 85 |  | 500
| 1.50 | 0 | 100 |  | 500
| 4.00 | 0 | 100 |  | 500
| 4.10 | 15 | 85 |  | 500
| 5.00 | 15 | 85 |  | 500
Table 9 LC-MRM MS conditions (LCMS-8060, Shimadzu Co.)
| Polarity mode | Positive
| Electrospray voltage | 4000 V
| Desolvation temperature | 250 °C
| Heat block temperature | 400 °C
| Nebulizing gas (N2) flow rate | 3.0 L/min
| Drying gas (N2) flow rate | 10.0 L/min
| Heating gas (Air) flow rate | 10.0 L/min
| Target Q1 pre bias | -20 V
| Target Q3 pre bias | -12 V
| Target collision energy (CE) | -25 eV
Table 10 LC-MRM Mass transitions used for absolute quantification. On different systems, RT and CE might be different!
| Analyte | Mass transitions | IS | Mass transitions | RT (min) | CE
| Cer d18:1/16:0 | 538.5 – 264.3 | D7- Cer d18:1/16:0 | 545.6 - 271.3 | 1.9-2.0 | 40
C34H68NO3 – C18H34N
| Cer d18:1/18:0 | 566.6 – 264.3 | D7- Cer d18:1/18:0 | 573.6 – 271.3 | 2.1-2.2 | 40
C36H72NO3 – C18H34N
| Cer d18:1/24:0 | 650.6 – 264.3 | D7- Cer d18:1/24:0 | 657.7 – 271.3 | 2.4-2.5 | 40
C42H84NO3 – C18H34N
| Cer d18:1/24:1 | 648.6 – 264.3 | D7- Cer d18:1/24:1 | 655.7 – 271.3 | 2.3-2.4 | 40
C42H82NO3 – C18H34N
Table 11 Data processing
| Processing software | Shimadzu LabSolutions 5.91

## LC-MRM - Sample Sequence
Sample Sequence (Example)
Sample Sequence (Example) | Sample Sequence 1
| Number | Sample Name | Remarks
| 1 | Total Blank 1
| 2 | Matrix blank
| 3 | Calibration line 1 from STD6 to 1 (n=3)
| 4 | […]
| 5 | Pooled QC
| 6 | Pooled QC
| 7 | SRM 1950-1
| 8 | SRM 1950-2
| 9 | SRM 1950-3
| 10 | SRM 1950-4
| 11 | SRM 1950-5
| 12 | SRM 1950-6
| 13 | Pooled QC
| 14 | Pooled QC
| 15 | T1D-1
| 16 | T1D-2
| 17 | T1D-3
| 18 | T1D-4
| 19 | T1D-5
| 20 | T1D-6
| 21 | Total Blank 2
| 22 | Pooled QC
| 23 | Pooled QC
| 24 | Young AA-1
| 25 | Young AA-2
| 26 | Young AA-3
| 27 | Young AA-4
| 28 | Young AA-5
| 29 | Young AA-6
| 30 | Pooled QC
| 31 | Pooled QC
| 32 | hTAG-1
| 33 | hTAG-2
| 34 | hTAG-3
| 35 | hTAG-4
| 36 | hTAG-5
| 37 | hTAG-6
| 38 | Pooled QC
| 39 | Pooled QC
| 40 | Calibration line 2 from STD1 to 6 (n=3)
| 41 | Intra-assay QCs: LLQC, LQC, MQC, HQC
| 42 | Matrix blank
| 43 | Total Blank 3

## LC-MRM - Intra Assay QC Proc
Intra Assay Procedure
Intra Assay Procedure | Quality Controls for Intra-Assay Variation
| Step | Description | Reference Table
| 1 | Prepare the Pooled NIST QC sample containing equal volumes of the four NIST reference plasma you received.
| 2 | The pooled QC representing middle quality control (MQC), low QC (LQC), and lower limit of quantification QC (LLOQ) is prepared by diluting MQC (e.g. two- and three-fold) with water.
| 3 | High QC (HQC) and upper limit of quantification (ULOQ) QC are prepared by spiking in 10 µL and 20 µL respectively of STD3 endogenous standards from table | Extraction - Overview'!B34
| 4 | Internal standards are added to all QC samples.
Intra-Assay Variation
| 6 | The precision and accuracy of the assay is determined for each of the four ceramides measured in LLOQ, LQC, MQC, HQC, and HLOQ.
| 7 | Intra-assay variation, precision, and accuracy is calculated for each of the ceramide in replicates of six (n = 6) at each QC concentration.
| 8 | The intra-assay precision (percentage coefficient variance, %CV) and accuracy (percentage accuracy, %Accuracy) are calculated from the nominal concentrations according to the formulas reported in corresponding sheets and tables of this document.
| 9 | Recovery estimation is not necessary for the published method but only for different ones. Stability estimation is not necessary for the published method but only for different ones. Calibration line is prepared in 5% BSA.

## LC-MRM - Intra Assay QC Res
INTRA ASSAY - QC RESULTS
INTRA ASSAY - QC RESULTS | Table 12 LC-MRM Calibration Line 1 Data (Beginning). Please report the values of each of the three replicates separately for each compound and standard. Each STD prepared in triplicate, injected once from each vial
| Calibration Line 1 | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| STD 1 | 9437114 | 9399450 | 9428055 | 7937255 | 7951827 | 7946600 | 46180037 | 45388775 | 46063100 | 40055782 | 39323814 | 39530253 | 466626 | 484274 | 485573 | 288654 | 291335 | 293397 | 3195799 | 3198492 | 3183887 | 926746 | 947023 | 914438
| STD 2 | 4274713 | 4278345 | 4271519 | 3742802 | 3723732 | 3711129 | 20078640 | 20620161 | 20438955 | 18871767 | 18943100 | 18787439 | 456173 | 454444 | 452478 | 277948 | 281571 | 282266 | 2961838 | 2990309 | 2986182 | 932298 | 944557 | 911935
| STD 3 | 421280 | 437077 | 438620 | 386059 | 378396 | 363384 | 2057392 | 2067050 | 2025135 | 1830047 | 1881044 | 1794250 | 449728 | 460475 | 478084 | 297769 | 288355 | 292676 | 3089524 | 3141672 | 3047218 | 940816 | 939945 | 951191
| STD 4 | 82310 | 84527 | 84643 | 70508 | 72530 | 72600 | 386532 | 384852 | 384505 | 350803 | 352841 | 353106 | 430606 | 437201 | 436857 | 272477 | 295891 | 292843 | 3030582 | 3016261 | 3028995 | 895868 | 900226 | 898424
| STD 5 | 42141 | 43235 | 43858 | 34826 | 34943 | 34099 | 194929 | 194463 | 196497 | 165437 | 168467 | 171588 | 430220 | 437284 | 457476 | 285606 | 291298 | 283856 | 2874850 | 2985551 | 3018436 | 905794 | 903381 | 869772
| STD 6 | 33582 | 33489 | 35864 | 25402 | 23969 | 25542 | 151993 | 152516 | 148828 | 125421 | 139161 | 136970 | 433037 | 433786 | 434641 | 296196 | 284729 | 285875 | 3036256 | 3007670 | 2983583 | 907790 | 909051 | 898317
Table 13 LC-MRM Calibration Line 2 Data (End). Please report the values of each of the three replicates separately for each compound and standard. Each STD prepared in triplicate, injected once from each vial
| Calibration Line 2 | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| STD 1 | 10624359 | 10818773 | 10864964 | 6329774 | 6408436 | 6577740 | 49792641 | 49974047 | 48891081 | 47218366 | 47698186 | 46909362 | 454300 | 447754 | 486745 | 223041 | 230330 | 235177 | 3625952 | 3839084 | 3699266 | 1053430 | 1087302 | 1072668
| STD 2 | 4608008 | 4864820 | 4845659 | 2983848 | 3137638 | 3134421 | 21925664 | 22013030 | 21195536 | 22475439 | 22072569 | 21991826 | 425358 | 452182 | 438084 | 219132 | 227181 | 235310 | 3430784 | 3429294 | 3415142 | 1034734 | 1047763 | 1033581
| STD 3 | 462299 | 455539 | 460709 | 282241 | 301783 | 295122 | 2134754 | 2093282 | 2123219 | 2101373 | 2091993 | 2101774 | 430463 | 434057 | 422453 | 235928 | 222576 | 217620 | 3408862 | 3336109 | 3341792 | 1026601 | 1027287 | 1014988
| STD 4 | 91562 | 111251 | 96310 | 55772 | 56740 | 56173 | 422000 | 425855 | 427319 | 420944 | 420550 | 417381 | 418339 | 439673 | 441951 | 209657 | 212622 | 220644 | 3522429 | 3514213 | 3529178 | 1041478 | 1043812 | 1037763
| STD 5 | 50002 | 52860 | 48880 | 26441 | 28751 | 25952 | 197877 | 219756 | 194288 | 198504 | 193806 | 197135 | 418893 | 409387 | 410488 | 213817 | 224608 | 218153 | 3341690 | 3413024 | 3388994 | 1020229 | 1028503 | 1034691
| STD 6 | 41151 | 45824 | 41512 | 22989 | 19147 | 19946 | 159117 | 163211 | 160716 | 161471 | 147713 | 157076 | 436442 | 433836 | 452759 | 232797 | 218628 | 225878 | 3521247 | 3491458 | 3434450 | 1063186 | 1044052 | 1019058
Table 14 LC-MRM Sample Data. Please report the values of each of the three replicates separately for each compound and standard. Each vial injected 3 times
| NIST SRM | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| SRM 1 | 951641 | 977557 | 987232 | 179598 | 175144 | 181806 | 2300280 | 2500677 | 2448195 | 922852 | 976909 | 994174 | 568656 | 627346 | 595148 | 311936 | 310004 | 304016 | 3064175 | 3207716 | 3170891 | 1066317 | 1098216 | 1109289
| SRM 2 | 976940 | 964182 | 974843 | 162664 | 160173 | 161897 | 2364111 | 2367443 | 2423486 | 967445 | 970891 | 974160 | 620338 | 625051 | 639263 | 310056 | 305370 | 304083 | 3397819 | 3359981 | 3376178 | 1179615 | 1211435 | 1219044
| SRM 3 | 979928 | 984461 | 989080 | 157955 | 157759 | 157208 | 2420143 | 2504024 | 2564194 | 954822 | 977253 | 992764 | 592563 | 604340 | 597262 | 287395 | 294279 | 296406 | 3169847 | 3356087 | 3484770 | 1129547 | 1137731 | 1162426
| SRM 4 | 964612 | 976992 | 967659 | 153530 | 162157 | 162515 | 2497800 | 2555566 | 2501467 | 1027320 | 1020671 | 1020414 | 623018 | 611135 | 600690 | 273841 | 271528 | 281786 | 3602121 | 3513885 | 3528395 | 1210761 | 1201778 | 1204734
| SRM 5 | 939582 | 960426 | 967850 | 151619 | 150863 | 158752 | 2425706 | 2521803 | 2525180 | 973624 | 1023755 | 1020609 | 586006 | 593149 | 595825 | 268051 | 271664 | 278684 | 3304246 | 3455718 | 3480847 | 1154994 | 1246202 | 1202656
| SRM 6 | 933048 | 968988 | 967288 | 137588 | 154804 | 148694 | 2420161 | 2478035 | 2467100 | 974070 | 996662 | 1054677 | 600415 | 611652 | 621488 | 279452 | 267239 | 278582 | 3621994 | 3662528 | 3674703 | 1261305 | 1254595 | 1314269
| NIST hTAG | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| hTAG 1 | 968694 | 961429 | 969537 | 139254 | 142952 | 141504 | 3555814 | 3984863 | 4060767 | 1600624 | 1589900 | 1611986 | 509656 | 522352 | 522486 | 250442 | 248520 | 235332 | 3812122 | 3927794 | 3794746 | 1393678 | 1362270 | 1348206
| hTAG 2 | 988351 | 986539 | 991285 | 141468 | 134945 | 143623 | 4024302 | 4156623 | 3957656 | 1571019 | 1662116 | 1626854 | 524808 | 520377 | 526961 | 244876 | 269096 | 248036 | 3950541 | 4016698 | 4007571 | 1398417 | 1406059 | 1402301
| hTAG 3 | 970067 | 986877 | 965657 | 134164 | 143550 | 150128 | 3847939 | 4089207 | 4144901 | 1577985 | 1687080 | 1647219 | 490747 | 522708 | 541648 | 235678 | 247463 | 236624 | 3906273 | 4163435 | 4176222 | 1339771 | 1446587 | 1444117
| hTAG 4 | 953214 | 965938 | 978083 | 137145 | 150306 | 142736 | 4211268 | 4041510 | 4175284 | 1645836 | 1655280 | 1707753 | 493769 | 513949 | 509364 | 235816 | 247692 | 244490 | 4019088 | 4139427 | 4062508 | 1424113 | 1429142 | 1474299
| hTAG 5 | 909586 | 938196 | 916804 | 135902 | 134610 | 135163 | 4075825 | 4099799 | 4182525 | 1633123 | 1649271 | 1751800 | 497109 | 525475 | 519762 | 239307 | 248891 | 258378 | 3998428 | 4091116 | 4156432 | 1379159 | 1422234 | 1479325
| hTAG 6 | 1296582 | 1337263 | 1322125 | 200956 | 193200 | 196485 | 4596368 | 4621634 | 4345351 | 2141301 | 2083402 | 1946959 | 512282 | 518140 | 500128 | 275955 | 282677 | 279124 | 4376562 | 4452866 | 4461813 | 1484931 | 1560330 | 1555818
| NIST T1D | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| T1D 1 | 811538 | 822791 | 829670 | 157430 | 158965 | 162467 | 2412195 | 2555995 | 2596076 | 1019984 | 1067637 | 1069527 | 531113 | 567519 | 586588 | 266940 | 271955 | 267416 | 3525286 | 3761610 | 3732394 | 1238877 | 1284036 | 1305996
| T1D 2 | 802621 | 841071 | 830044 | 158602 | 159430 | 161871 | 2520997 | 2607866 | 2666362 | 1083970 | 1150494 | 1150173 | 603146 | 638592 | 635953 | 276350 | 294815 | 297264 | 4010016 | 4183383 | 4274475 | 1368915 | 1466865 | 1491509
| T1D 3 | 824896 | 822016 | 825396 | 154751 | 152005 | 151343 | 2559155 | 2537057 | 2575989 | 1105850 | 1089992 | 1053611 | 535668 | 535582 | 536628 | 250983 | 252815 | 240341 | 3566130 | 3611837 | 3552840 | 1197197 | 1203088 | 1208744
| T1D 4 | 822728 | 830960 | 826807 | 156949 | 163372 | 158919 | 2628251 | 2501536 | 2563854 | 1096402 | 1104212 | 1071083 | 543681 | 553590 | 539211 | 255796 | 261285 | 260615 | 3652005 | 3550942 | 3617289 | 1249542 | 1226831 | 1210733
| T1D 5 | 736555 | 745620 | 729470 | 135029 | 140369 | 137327 | 2325353 | 2286299 | 2312998 | 967666 | 926145 | 966249 | 513034 | 505424 | 518678 | 242495 | 235200 | 235840 | 3363803 | 3400017 | 3392383 | 1164532 | 1143853 | 1151097
| T1D 6 | 789340 | 796363 | 801085 | 152523 | 156819 | 154091 | 2438957 | 2564970 | 2556063 | 1039333 | 1105063 | 1100059 | 526395 | 536198 | 522124 | 246063 | 243518 | 248733 | 3508304 | 3760162 | 3652837 | 1191207 | 1250162 | 1212646
| NIST Young AA | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| yA-A 1 | 855107 | 902560 | 902556 | 100071 | 95413 | 97846 | 2090963 | 2089472 | 2078797 | 789724 | 800842 | 815582 | 610244 | 651639 | 640800 | 248725 | 270802 | 267380 | 3720391 | 3607065 | 3882496 | 1226396 | 1251035 | 1292318
| yA-A 2 | 799734 | 803833 | 807757 | 83863 | 97288 | 95257 | 1937822 | 1961633 | 1980565 | 767850 | 766459 | 757706 | 571890 | 594460 | 597038 | 251486 | 249632 | 240846 | 3658728 | 3632679 | 3665335 | 1225262 | 1236238 | 1221663
| yA-A 3 | 915026 | 901878 | 888269 | 102510 | 89943 | 95720 | 2071562 | 1801673 | 2065177 | 772159 | 710877 | 829475 | 636554 | 620268 | 615948 | 264264 | 264279 | 251786 | 3847997 | 3299409 | 3782436 | 1252879 | 1106162 | 1253586
| yA-A 4 | 837404 | 848122 | 847912 | 94452 | 93959 | 82295 | 1971758 | 2010629 | 1949104 | 781995 | 804911 | 787514 | 596408 | 615226 | 601398 | 252756 | 254269 | 258955 | 3896162 | 3751343 | 3738520 | 1263646 | 1269732 | 1243432
| yA-A 5 | 784673 | 824638 | 803913 | 84728 | 94939 | 96821 | 1890219 | 2028206 | 2071240 | 741836 | 780799 | 823117 | 551590 | 612733 | 603461 | 239109 | 272970 | 246280 | 3510106 | 3821604 | 3813480 | 1205798 | 1253725 | 1267990
| yA-A 6 | 847143 | 864570 | 866601 | 84981 | 92967 | 84555 | 2015087 | 2089009 | 2067193 | 782974 | 827098 | 831080 | 575700 | 605571 | 590613 | 225408 | 265095 | 264273 | 3676306 | 3755969 | 3824382 | 1211128 | 1250436 | 1272263
Table 15 LC-MRM Intra Assay QC Data. Please report the values of each of the three replicates separately for each compound and standard. QCs have to be measured in one occasion only. The independent 6 replicates injected 3 times each
| Lowest Level QC | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| LLQC 1 | 281835 | 267018 | 268045 | 44843 | 47033 | 45863 | 846140 | 846744 | 871092 | 314674 | 310306 | 304698 | 576140 | 539542 | 548668 | 254867 | 234090 | 237086 | 3553774 | 3665426 | 3405395 | 1085694 | 1064312 | 1030313
| LLQC 2 | 293058 | 291564 | 290148 | 42844 | 38560 | 40720 | 859806 | 849370 | 841251 | 309582 | 310144 | 313727 | 578308 | 556352 | 553949 | 242735 | 244538 | 253459 | 3662403 | 3688666 | 3655802 | 1095455 | 1085729 | 1098683
| LLQC 3 | 296640 | 300965 | 305576 | 42080 | 44960 | 45065 | 862186 | 901341 | 864852 | 334330 | 323152 | 311229 | 593111 | 608684 | 612954 | 248193 | 248179 | 260493 | 3637392 | 3442176 | 3700439 | 1157270 | 1110766 | 1156618
| LLQC 4 | 289964 | 286300 | 286371 | 40490 | 42703 | 45349 | 864944 | 848223 | 874998 | 340944 | 331238 | 338555 | 599243 | 580914 | 594271 | 253640 | 252131 | 247569 | 3699822 | 3842653 | 3855833 | 1169090 | 1163487 | 1188096
| LLQC 5 | 288569 | 274584 | 281663 | 40921 | 38520 | 36810 | 890744 | 830144 | 812896 | 333053 | 316434 | 310165 | 599979 | 550109 | 578408 | 246364 | 255804 | 239191 | 3684664 | 3484328 | 3391538 | 1153491 | 1109396 | 1096915
| LLQC 6 | 283707 | 277989 | 277337 | 38062 | 34793 | 38254 | 853263 | 823070 | 831756 | 327778 | 310292 | 308805 | 600235 | 570628 | 585080 | 265662 | 265891 | 244410 | 3610201 | 3533630 | 3470418 | 1131134 | 1095577 | 1076781
| Low QC | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| LQC 1 | 421391 | 391252 | 391253 | 61271 | 62977 | 61412 | 1198707 | 1175572 | 1181608 | 451710 | 436531 | 428641 | 602349 | 552804 | 562154 | 260016 | 244042 | 257038 | 3551403 | 3589721 | 3335060 | 1149707 | 1104523 | 1069239
| LQC 2 | 432328 | 421521 | 419474 | 59756 | 56705 | 53557 | 1230681 | 1191428 | 1180039 | 463358 | 454915 | 460171 | 609853 | 574965 | 572482 | 269896 | 266463 | 276183 | 3725205 | 3676880 | 3644121 | 1186992 | 1152925 | 1166680
| LQC 3 | 445994 | 443446 | 450240 | 58740 | 62609 | 61649 | 1220810 | 1318475 | 1200093 | 469168 | 499421 | 468014 | 607633 | 611116 | 615402 | 256460 | 251317 | 263787 | 3767875 | 4006468 | 3756520 | 1192307 | 1223441 | 1167802
| LQC 4 | 441138 | 426852 | 426958 | 56792 | 55949 | 60335 | 1231181 | 1183232 | 1220581 | 474609 | 451875 | 461858 | 614117 | 583427 | 596842 | 268409 | 261476 | 256745 | 3750394 | 3817273 | 3830366 | 1174542 | 1145535 | 1169764
| LQC 5 | 422787 | 394252 | 404416 | 60758 | 58139 | 55106 | 1232876 | 1166019 | 1182624 | 472343 | 459797 | 447187 | 611661 | 589613 | 587904 | 264103 | 246716 | 251285 | 3768169 | 3591806 | 3599032 | 1178163 | 1110463 | 1097970
| LQC 6 | 443115 | 425500 | 424503 | 56819 | 54900 | 56964 | 1260810 | 1191871 | 1204450 | 459398 | 426193 | 424151 | 590912 | 550529 | 564472 | 241396 | 239671 | 227643 | 3795609 | 3640804 | 3575675 | 1143378 | 1085287 | 1066668
| Middle QC | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| MQC 1 | 912271 | 880903 | 880907 | 141106 | 150837 | 147086 | 3116318 | 3178418 | 3194740 | 1223431 | 1229611 | 1207388 | 567548 | 541700 | 550863 | 277094 | 259391 | 262711 | 4105597 | 4315890 | 4009714 | 1387213 | 1386002 | 1341727
| MQC 2 | 935353 | 948451 | 943844 | 141160 | 139484 | 129662 | 3070680 | 3091649 | 3062096 | 1171850 | 1196518 | 1210340 | 558225 | 547342 | 544979 | 266541 | 273677 | 283660 | 4175395 | 4286079 | 4247892 | 1334547 | 1348094 | 1364178
| MQC 3 | 929914 | 961586 | 976318 | 139999 | 153999 | 152809 | 2974367 | 3340810 | 3040848 | 1192941 | 1300659 | 1231832 | 574524 | 600930 | 605145 | 268364 | 273502 | 287072 | 4103867 | 4378107 | 4255160 | 1360070 | 1270042 | 1305402
| MQC 4 | 928232 | 934099 | 934330 | 139183 | 142600 | 145885 | 2999535 | 2998024 | 3092659 | 1199360 | 1187587 | 1213822 | 565525 | 558754 | 571601 | 263923 | 267390 | 262552 | 4129032 | 4370770 | 4385762 | 1355326 | 1374728 | 1403805
| MQC 5 | 935625 | 907375 | 930767 | 144454 | 131393 | 130839 | 2967710 | 2818909 | 2760340 | 1197464 | 1159553 | 1099939 | 588868 | 540285 | 548586 | 276566 | 266911 | 273669 | 4112623 | 3849931 | 3858132 | 1383778 | 1356433 | 1341173
| MQC 6 | 919001 | 917766 | 915616 | 143241 | 133451 | 146727 | 3068215 | 3016468 | 3048302 | 1238802 | 1195231 | 1189504 | 551798 | 534652 | 548193 | 267783 | 253430 | 251091 | 4086879 | 4077002 | 4004070 | 1341220 | 1324001 | 1301286
| High QC | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| HQC 1 | 1376015 | 1421968 | 1424369 | 390590 | 403184 | 398321 | 8252247 | 8306709 | 8571188 | 4471880 | 4528439 | 4669557 | 951549 | 978731 | 987463 | 500863 | 503592 | 497096 | 11658868 | 11912271 | 12440754 | 3470563 | 3540268 | 3610038
| HQC 2 | 1412705 | 1439511 | 1419428 | 401052 | 393921 | 386800 | 8193293 | 8204434 | 8441215 | 4456213 | 4537158 | 4586780 | 967751 | 961901 | 956733 | 480940 | 518307 | 501978 | 11689317 | 11741241 | 12054776 | 3430071 | 3471238 | 3639665
| HQC 3 | 1396327 | 1411919 | 1381486 | 406848 | 400702 | 388914 | 8325846 | 8266320 | 8367226 | 4567848 | 4632263 | 4620950 | 965501 | 973011 | 942177 | 490451 | 502379 | 504918 | 11668371 | 11960601 | 12032033 | 3461689 | 3569762 | 3604516
| HQC 4 | 1403817 | 1358527 | 1385149 | 392537 | 386139 | 395264 | 7869179 | 8439152 | 8373951 | 4377704 | 4654470 | 4636177 | 973887 | 951628 | 941468 | 494722 | 502307 | 487503 | 11638814 | 12089174 | 12077711 | 3442073 | 3572682 | 3587813
| HQC 5 | 1363246 | 1405263 | 1390445 | 376032 | 393955 | 383570 | 7868579 | 8427736 | 8401932 | 4284150 | 4645776 | 4611908 | 968420 | 951433 | 955325 | 480689 | 493751 | 493301 | 11259129 | 11903513 | 12095973 | 3330218 | 3577402 | 3575258
| HQC 6 | 1449677 | 1403691 | 1369326 | 395907 | 396508 | 392074 | 8268151 | 8340548 | 8458720 | 4514605 | 4572989 | 4655906 | 1019521 | 947638 | 948996 | 519030 | 498131 | 508533 | 11873916 | 11990393 | 12121713 | 3477132 | 3605811 | 3542492
| Highest Level QC | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| HLQC 1 | 1815832 | 1824864 | 1885144 | 690752 | 663912 | 660708 | 11817801 | 12012086 | 11954775 | 7752052 | 7790227 | 7869784 | 653889 | 641567 | 681830 | 361294 | 350953 | 355380 | 8715012 | 8944182 | 9013166 | 2562379 | 2634680 | 2637447
| HLQC 2 | 1866936 | 1904698 | 1828751 | 680688 | 672126 | 665044 | 11978763 | 12212190 | 11925062 | 7752909 | 8062304 | 7817501 | 683320 | 690875 | 668351 | 380800 | 362148 | 361374 | 8859933 | 9244652 | 8905848 | 2560662 | 2695469 | 2616918
| HLQC 3 | 1896527 | 1843027 | 1819044 | 701332 | 673702 | 647737 | 11820632 | 12021663 | 11828602 | 7743230 | 7833736 | 7684521 | 693549 | 686038 | 674419 | 372170 | 357563 | 344581 | 8716887 | 9034267 | 8883833 | 2522001 | 2648076 | 2594350
| HLQC 4 | 1877931 | 1916264 | 1837397 | 701666 | 691691 | 634081 | 12047857 | 11823092 | 11672078 | 7873304 | 7813841 | 7740076 | 671236 | 672253 | 677913 | 379975 | 378208 | 360153 | 8929936 | 8799071 | 8793936 | 2670840 | 2590457 | 2565700
| HLQC 5 | 1879066 | 1846571 | 1861464 | 681092 | 665796 | 655763 | 11870455 | 12226392 | 11744685 | 7750496 | 7997979 | 7757449 | 659376 | 665511 | 667853 | 372244 | 367198 | 357526 | 8748756 | 9235672 | 8922929 | 2530355 | 2680480 | 2588407
| HLQC 6 | 1855243 | 1884093 | 1792912 | 671107 | 679743 | 663370 | 11968870 | 12158488 | 11748294 | 7775481 | 7994523 | 7782531 | 674789 | 702563 | 657479 | 359295 | 371509 | 358891 | 8711832 | 8992475 | 8906940 | 2561371 | 2601438 | 2584938
| Blank QC | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| Matrix Blank | 867857 | 931593 | 923091 | 210743 | 140262 | 134735 | 2521393 | 2813804 | 3106443 | 1013751 | 1108378 | 1251167 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| Total Blank | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0

## Raw Peak Areas
Sample_Name | Cer d18:1/16:0 [M>264] | Cer d18:1/16:0-d7 [M>264] | Cer d18:1/18:0 [M>264] | Cer d18:1/18:0-d7 [M>264] | Cer d18:1/24:0 [M>264] | Cer d18:1/24:0-d7 [M>264] | Cer d18:1/24:1 [M>264] | Cer d18:1/24:1-d7 [M>264]
Total Blank 1
Matrix blank
Calibration line 1 from STD6 to 1 (n=3)
[…]
Pooled QC
Pooled QC
SRM 1950-1
SRM 1950-2
SRM 1950-3
SRM 1950-4
SRM 1950-5
SRM 1950-6
Pooled QC
Pooled QC
T1D-1
T1D-2
T1D-3
T1D-4
T1D-5
T1D-6
Total Blank 2
Pooled QC
Pooled QC
Young AA-1
Young AA-2
Young AA-3
Young AA-4
Young AA-5
Young AA-6
Pooled QC
Pooled QC
hTAG-1
hTAG-2
hTAG-3
hTAG-4
hTAG-5
hTAG-6
Pooled QC
Pooled QC
Calibration line 2 from STD1 to 6 (n=3)
[…]
Matrix blank
Total Blank 3