## Participant Information
Participant Information
Participant Information |  | Group Name
|  | Lab code | #32
| Institution | Institution name
Street
City
Postal code
Country
| Contact person | Name
Address
Email
Phone

## Summary
Summary
Summary | This protocol describes the extraction and LC-MRM procedure for absolute quantification of Cer d18:1/16:0, Cer d18:1/18:0, Cer d18:1/24:0 and Cer d18:1/24:1 in human plasma according to Kauhanen et al, Anal Bioanal Chem (2016) 408:3475–3483.
A calibration curve is built by injecting several dilutions of the non-labelled standards mixed with a fixed concentration of labelled standards. Each calibration point is prepared and analysed in triplicates (n=3).
To establish the endogenous concentration of the four selected ceramides in the study samples (NIST SRM1950, NIST high TAG, NIST T1D, and NIST young A-A) each sample is extracted and analyzed in six replicates (n =6).
Prepare the pooled QC by mixing 100 mL each of NIST SRM1950, NIST high TAG, NIST T1D, and NIST young AA (total 400 mL).
We suggest proceeding first to measure the ceramides concentration in the four study samples and then continue with the intra-assay validation to avoid wasting material.

## Extraction - Overview
Extraction Overview
Extraction Overview | Table 1 - Extraction - Materials
Chemicals, solvents and materials
Ethyl acetate (EtAc) (Fisher Chemical)
2-propanol (IPA) - Optima LC-MS-grade (Fisher Chemical)
5% Bovine Serum Albumin (BSA) (Sigma Aldrich)
300 µL LC vials (Verex)
1.5 mL tubes (Extragene)
tips 1000 µL (Extragene); tips 2000 µL (Tarsons)
Table 2 - Extraction - Sample Types
| Sample type | Components
| Pooled Quality Control samples (QC) | 10 µL pooled QC |  |  | 20 µL Labelled IS |  | 570 µL EtAc:IPA 2:8 (v/v)
| Matrix Blank (MB) | 10 µL Matrix (pooled QC) only |  |  | - |  | 590 µL EtAc:IPA 2:8
| Study Samples (S) | 10 µL Human plasma (NIST SRM1950, high TAG, T1D, young AA) |  |  | 20 µL Labelled IS |  | 570 µL EtAc:IPA 2:8
| Total Blank (TB) | 10 µL 5% BSA |  |  | - |  | 590 µL EtAc:IPA 2:8
| Calibration line standards (STD) | 10 µL 5% BSA + 20 µL non labelled STD mixture (Table 5) |  |  | 20 µL Labelled IS |  | 550 µL EtAc:IPA 2:8
Table 3 - Extraction - Labelled Internal Standard Mixture Solution (provided)
| Labelled IS mixture in EtAc:IPA 2:8 (v/v) | Concentration |  |  |  |  | Unit | Vendor
| D7-Cer d18:1/16:0 | 0.125 |  |  |  |  | pmol/µL [µM] | Avanti Polar Lipids
| D7-Cer d18:1/18:0 | 0.050
| D7-Cer d18:1/24:0 | 1500
| D7-Cer d18:1/24:1 | 0.500
Table 4 - Extraction - Non-Labelled Standards - for STD1 Mixture Solution (provided)
| Non-labelled standards in EtAc:IPA 2:8 (v/v) | Concentration |  |  |  |  | Unit | Vendor
| Cer d18:1/16:0 | 2.0 |  |  |  |  | pmol/µL [µM] | Avanti Polar Lipids
| Cer d18:1/18:0 | 2.0
| Cer d18:1/24:0 | 20.0
| Cer d18:1/24:1 | 20.0
Table 5 - Serial dilutions of Non-Labelled Standards from STD1 Mixture Solution for calibration line
| Add 20 µL of | STD1 Conc. | STD2 Conc. | STD3 Conc. | STD4 Conc. | STD5 Conc. | STD6 Conc. | Unit
| Cer d18:1/16:0 | 2.0 | 1.0 | 0.10 | 0.02 | 0.01 | 0.008 | pmol/µL [µM]
| Cer d18:1/18:0 | 2.0 | 1.0 | 0.10 | 0.02 | 0.01 | 0.008
| Cer d18:1/24:0 | 20.0 | 10.0 | 1.0 | 0.2 | 0.1 | 0.08
| Cer d18:1/24:1 | 20.0 | 10.0 | 1.0 | 0.2 | 0.1 | 0.08
We decided to include two additional points in the calibration lines without modifying the SOP procedure for calibration lines
Table 6 - Serial dilutions of Non-Labelled Standards from STD3 (additional calibration line points)
|  | STD3 Conc. | STD3_dilution1/2 Conc. | STD3_dlution1/3 Conc. | Unit
| Cer d18:1/16:0 | 0.10 | 0.05 | 0.033 | pmol/µL [µM]
| Cer d18:1/18:0 | 0.10 | 0.05 | 0.033
| Cer d18:1/24:0 | 1.0 | 0.50 | 0.33
| Cer d18:1/24:1 | 1.0 | 0.50 | 0.33

## Extraction - Procedure
Extraction Procedure
Extraction Procedure | Extraction Procedure
| Step | Description | Reference Table
| 1 | Prepare different dilutions of STD1 by diluting the provided mixture (Extraction - Overview, Table 4) with EtAc:IPA 2:8 as suggested in table | Extraction - Overview'!B34
| 2 | Thaw plasma on ice (+4°C) and bring to room temperature prior to extraction. Mix by pipetting before use.
| 3 | Prepare samples in plates/tubes manually or with a robot, according to specifications in column C of table | Extraction - Overview'!C14
| 4 | Add 20 µL of labelled internal standard mixture to samples according to column F of table | Extraction - Overview'!F14
| 5 | For the calibration line, add 20 µL of different dilutions of the non-labelled mixture (STD1-6).
| 6 | Add 590 µL of ethyl acetate:isopropanol (2:8, vol/vol) to Total Blank (TB) and Matrix Blank (MB); 550 µL to calibration line (STD); 570 µL to study samples and QCs according to column H of table | Extraction - Overview'!H14
| 7 | Mix samples 10 min by pipetting or vortexing
| 8 | Centrifuge samples for 10 min at 3000 x gat room temperature
| 9 | Transfer 50 µL of the clear supernatant to MS-vials or MS plates, close/seal, store samples at -20 °C prior to LC-MRM analysis.

## LC-MRM - Overview
LC-MRM Overview
LC-MRM Overview | Table 6 Chemicals and solvents
Ammonium acetate LC-MS-grade (Fisher Chemical)
Formic acid LC-MS-grade (Fisher Chemical)
Ultra-pure water 18.2 MΩ·cm resistivity (Thermo Scientific Barnstead Micropure UV ultrapure water system)
2-propanol (IPA) - Optima LC-MS-grade (Fisher Chemical)
Acetonitrile - Optima LC-MS-grade (Fisher Chemical)
Table 7 LC-MRM LC-MS Details
| Property | Specifications
| Mass spectrometry | Waters Xevo G2S QTOF mass spectrometer
| LC | Waters ACQUITY UPLC I-Class
| Injection volume | 5 µL
| Wash | 2-propanol (IPA) - Optima LC-MS-grade (Fisher Chemical)
| Pre-column | Waters Acquity BEH C18, 1.7 µm VanGuard Pre-Column
| Column | Waters Acquity BEH C18, 2.1 × 50 mm id. 1.7 µm
| Column temperature | 60 °C
| Mobile phase A | 10 mM ammonium acetate in water with 0.1% formic acid
| Mobile phase B | 10 mM ammonium acetate in acetonitrile:2-propanol (4:3, vol/vol) with 0.1% formic acid
Table 8 LC-MRM LC gradient
| Time (minutes) | A% | B% |  | Flow (µL/min)
| 0.00 | 15 | 85 |  | 500
| 0.50 | 15 | 85 |  | 500
| 1.50 | 0 | 100 |  | 500
| 4.00 | 0 | 100 |  | 500
| 4.10 | 15 | 85 |  | 500
| 5.00 | 15 | 85 |  | 500
Table 9 LC-MRM MS conditions - Xevo G2S QTOF mass spectrometer (Waters Corporation, Manchester, U.K.)
| Polarity mode | Positive
| Source capillary (ESI) | 2500 V
| Source Sampling cone | 30 V
| Source Offset | 80 V
| Source temperature | 150 °C
| Desolvation gas temperature | 500 °C
| Gas flow - Cone gas | 10 L/h
| Gas flow - Desolvation gas | 600 L/h
| MS continuum mode - Scan Time | 1 s
| MS continuum mode - m/z acquisition range | 50-1200
| LockSpray solution | Leucine Enkephalin (m/z 556.2771) Scan time: 0.5 s / Interval: 54 s / Flow rate: 5 μL min−1
Table 10 LC-MRM Mass transitions used for absolute quantification. On different systems, RT and CE might be different!
| Analyte | m/z ([M-H2O+H]+) | IS | m/z ([M-H2O+H]+) | RT (min) | CE*
| Cer d18:1/16:0 | 520.5094 | D7- Cer d18:1/16:0 | 527.5528 | 1.46 | -
| Cer d18:1/18:0 | 548.5407 | D7- Cer d18:1/18:0 | 555.5841 | 1.63 | -
| Cer d18:1/24:0 | 632.6346 | D7- Cer d18:1/24:0 | 639.678 | 2.01 | -
| Cer d18:1/24:1 | 630.6189 | D7- Cer d18:1/24:1 | 637.6623 | 1.89 | -
Table 11 Data processing
| Processing software | MassLynx V4.1
*The collision cell was operated with a transmission voltage of 6V since data was acquired in MS mode only.

## LC-MRM - Sample Sequence
Sample Sequence (Example)
Sample Sequence (Example) | Sample Sequence 1
| Number | Sample Name | Remarks
| 1 | Total Blank 1
| 2 | Matrix blank
| 3 | Calibration line 1 from STD6 to 1 (n=3)
| 4 | […]
| 5 | Pooled QC
| 6 | Pooled QC
| 7 | SRM 1950-1
| 8 | SRM 1950-2
| 9 | SRM 1950-3
| 10 | SRM 1950-4
| 11 | SRM 1950-5
| 12 | SRM 1950-6
| 13 | Pooled QC
| 14 | Pooled QC
| 15 | T1D-1
| 16 | T1D-2
| 17 | T1D-3
| 18 | T1D-4
| 19 | T1D-5
| 20 | T1D-6
| 21 | Total Blank 2
| 22 | Pooled QC
| 23 | Pooled QC
| 24 | Young AA-1
| 25 | Young AA-2
| 26 | Young AA-3
| 27 | Young AA-4
| 28 | Young AA-5
| 29 | Young AA-6
| 30 | Pooled QC
| 31 | Pooled QC
| 32 | hTAG-1
| 33 | hTAG-2
| 34 | hTAG-3
| 35 | hTAG-4
| 36 | hTAG-5
| 37 | hTAG-6
| 38 | Pooled QC
| 39 | Pooled QC
| 40 | Calibration line 2 from STD1 to 6 (n=3)
| 41 | Intra-assay QCs: LLQC, LQC, MQC, HQC
| 42 | Matrix blank
| 43 | Total Blank 3

## LC-MRM - Intra Assay QC Proc
Intra Assay Procedure
Intra Assay Procedure | Quality Controls for Intra-Assay Variation
| Step | Description | Reference Table
| 1 | Prepare the Pooled NIST QC sample containing equal volumes of the four NIST reference plasma you received.
| 2 | The pooled QC representing middle quality control (MQC), low QC (LQC), and lower limit of quantification QC (LLOQ) is prepared by diluting MQC (e.g. two- and three-fold) with water.
| 3 | High QC (HQC) and upper limit of quantification (ULOQ) QC are prepared by spiking in 10 µL and 20 µL respectively of STD3 endogenous standards from table | Extraction - Overview'!B34
| 4 | Internal standards are added to all QC samples.
Intra-Assay Variation
| 6 | The precision and accuracy of the assay is determined for each of the four ceramides measured in LLOQ, LQC, MQC, HQC, and HLOQ.
| 7 | Intra-assay variation, precision, and accuracy is calculated for each of the ceramide in replicates of six (n = 6) at each QC concentration.
| 8 | The intra-assay precision (percentage coefficient variance, %CV) and accuracy (percentage accuracy, %Accuracy) are calculated from the nominal concentrations according to the formulas reported in corresponding sheets and tables of this document.
| 9 | Recovery estimation is not necessary for the published method but only for different ones. Stability estimation is not necessary for the published method but only for different ones. Calibration line is prepared in 5% BSA.

## LC-MRM - Intra Assay QC Res
INTRA ASSAY - QC RESULTS
INTRA ASSAY - QC RESULTS | Table 12 LC-MRM Calibration Line 1 Data (Beginning). Please report the values of each of the three replicates separately for each compound and standard. Each STD prepared in triplicate, injected once from each vial
| Calibration Line 1 | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| STD 1 | 1118 | 1002 | 1084 | 1770 | 1670 | 1595 | 17888 | 16621 | 12253 | 17177 | 18598 | 16799 | 65 | 47 | 52 | 52 | 39 | 48 | 1110 | 925 | 796 | 406 | 361 | 357
| STD 2 | 534 | 621 | 447 | 837 | 930 | 736 | 10285 | 10174 | 10617 | 7357 | 8963 | 8676 | 52 | 62 | 50 | 51 | 50 | 42 | 1281 | 1097 | 1218 | 325 | 376 | 369
| STD 3 | 67 | 44 | 51 | 80 | 87 | 57 | 966 | 999 | 959 | 824 | 790 | 755 | 64 | 42 | 54 | 53 | 48 | 54 | 1426 | 1368 | 1284 | 373 | 405 | 386
| STD 4 | ND | ND | ND | ND | ND | ND | 238 | 222 | 236 | 176 | 184 | 176 | 58 | 65 | 47 | 54 | 57 | 48 | 1613 | 1550 | 1454 | 479 | 458 | 412
| STD 5 | ND | ND | ND | ND | ND | ND | 131 | 135 | 129 | 98 | 100 | 93 | 84 | 66 | 72 | 80 | 64 | 57 | 1906 | 1676 | 1585 | 449 | 463 | 424
| STD 6 | ND | ND | ND | ND | ND | ND | 182 | 181 | 121 | 146 | 145 | 108 | 141 | 86 | 74 | 142 | 103 | 88 | 3258 | 2712 | 1919 | 938 | 791 | 633
Table 13 LC-MRM Calibration Line 2 Data (End). Please report the values of each of the three replicates separately for each compound and standard. Each STD prepared in triplicate, injected once from each vial
| Calibration Line 2 | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| STD 1 | 805 | 843 | 696 | 911 | 1147 | 1332 | 13194 | 15569 | 16751 | 12703 | 11974 | 13667 | 41 | 46 | 36 | 33 | 37 | 40 | 825 | 922 | 922 | 266 | 271 | 267
| STD 2 | 308 | 355 | 384 | 565 | 597 | 641 | 6450 | 7683 | 6346 | 6898 | 6206 | 6215 | 32 | 42 | 35 | 30 | 36 | 38 | 752 | 1010 | 793 | 273 | 280 | 299
| STD 3 | 41 | 45 | 44 | 74 | 58 | 62 | 674 | 640 | 613 | 535 | 587 | 586 | 41 | 42 | 37 | 43 | 38 | 33 | 954 | 891 | 865 | 231 | 276 | 271
| STD 4 | ND | ND | ND | ND | ND | ND | 127 | 143 | 127 | 108 | 102 | 95 | 32 | 48 | 43 | 31 | 29 | 40 | 843 | 864 | 963 | 284 | 286 | 230
| STD 5 | ND | ND | ND | ND | ND | ND | 70 | 65 | 62 | 59 | 51 | 44 | 33 | 41 | 24 | 33 | 37 | 22 | 903 | 768 | 803 | 269 | 287 | 258
| STD 6 | ND | ND | ND | ND | ND | ND | 40 | 56 | 54 | 52 | 45 | 51 | 35 | 33 | 39 | 36 | 28 | 34 | 655 | 794 | 965 | 260 | 279 | 259
Table 14 LC-MRM Sample Data. Please report the values of each of the three replicates separately for each compound and standard. Each vial injected 3 times
| NIST SRM | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| SRM 1 | 71 | 85 | 87 | 65 | 66 | 61 | 975 | 1056 | 1091 | 547 | 404 | 438 | 52 | 55 | 60 | 49 | 51 | 40 | 1020 | 1037 | 1067 | 393 | 272 | 301
| SRM 2 | 51 | 46 | 57 | 28 | 35 | 31 | 553 | 485 | 468 | 275 | 299 | 311 | 32 | 31 | 42 | 27 | 38 | 37 | 791 | 764 | 784 | 235 | 221 | 229
| SRM 3 | 87 | 94 | 48 | 53 | 59 | 50 | 942 | 1021 | 798 | 426 | 457 | 417 | 50 | 55 | 32 | 42 | 42 | 38 | 933 | 978 | 812 | 277 | 321 | 287
| SRM 4 | 76 | 93 | 72 | 57 | 46 | 50 | 855 | 918 | 939 | 428 | 364 | 376 | 41 | 46 | 40 | 45 | 41 | 44 | 805 | 975 | 1001 | 292 | 264 | 262
| SRM 5 | 77 | 75 | 60 | 48 | 39 | 47 | 830 | 829 | 928 | 426 | 440 | 470 | 41 | 43 | 37 | 42 | 33 | 43 | 784 | 867 | 935 | 279 | 302 | 297
| SRM 6 | 84 | 57 | 79 | 38 | 37 | 51 | 917 | 825 | 923 | 360 | 392 | 394 | 41 | 29 | 42 | 32 | 30 | 42 | 951 | 793 | 935 | 271 | 295 | 270
| NIST hTAG | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| hTAG 1 | 58 | 58 | 68 | 45 | 43 | 42 | 1161 | 1078 | 1206 | 486 | 464 | 358 | 34 | 35 | 36 | 33 | 33 | 38 | 864 | 833 | 825 | 242 | 257 | 190
| hTAG 2 | 63 | 69 | 58 | 35 | 35 | 39 | 1116 | 1067 | 1166 | 438 | 329 | 419 | 38 | 35 | 38 | 31 | 31 | 37 | 814 | 849 | 868 | 238 | 179 | 218
| hTAG 3 | 69 | 53 | 56 | 43 | 26 | 35 | 636 | 892 | 723 | 292 | 367 | 322 | 34 | 29 | 32 | 43 | 26 | 33 | 533 | 796 | 694 | 191 | 223 | 232
| hTAG 4 | 63 | 46 | 52 | 31 | 28 | 28 | 820 | 863 | 924 | 339 | 374 | 324 | 37 | 30 | 29 | 32 | 27 | 29 | 706 | 707 | 742 | 195 | 240 | 207
| hTAG 5 | 51 | 68 | 55 | 27 | 29 | 33 | 637 | 745 | 730 | 306 | 328 | 324 | 28 | 41 | 36 | 23 | 27 | 32 | 542 | 610 | 675 | 198 | 187 | 211
| hTAG 6 | 51 | 50 | 48 | 33 | 32 | 35 | 669 | 771 | 736 | 268 | 264 | 302 | 41 | 42 | 36 | 35 | 33 | 36 | 651 | 767 | 806 | 218 | 210 | 217
| NIST T1D | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| T1D 1 | 80 | 67 | 65 | 43 | 49 | 50 | 829 | 832 | 666 | 303 | 331 | 357 | 37 | 38 | 36 | 38 | 37 | 42 | 900 | 941 | 706 | 250 | 258 | 274
| T1D 2 | 70 | 73 | 75 | 59 | 57 | 56 | 902 | 907 | 872 | 386 | 380 | 412 | 42 | 42 | 45 | 44 | 46 | 38 | 900 | 1003 | 925 | 266 | 245 | 276
| T1D 3 | 72 | 74 | 65 | 60 | 53 | 51 | 927 | 953 | 958 | 361 | 358 | 373 | 48 | 45 | 39 | 43 | 37 | 37 | 995 | 958 | 908 | 264 | 256 | 241
| T1D 4 | 57 | 55 | 50 | 33 | 33 | 40 | 723 | 717 | 780 | 358 | 320 | 323 | 35 | 35 | 33 | 27 | 30 | 36 | 749 | 725 | 777 | 263 | 269 | 235
| T1D 5 | 75 | 80 | 69 | 47 | 55 | 53 | 866 | 921 | 834 | 326 | 391 | 378 | 42 | 43 | 39 | 35 | 42 | 40 | 847 | 934 | 866 | 270 | 289 | 242
| T1D 6 | 67 | 70 | 73 | 54 | 45 | 41 | 862 | 850 | 872 | 342 | 347 | 368 | 38 | 35 | 36 | 39 | 38 | 33 | 826 | 855 | 915 | 275 | 276 | 250
| NIST Young AA | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| yA-A 1 | 41 | 38 | 55 | 25 | 28 | 24 | 419 | 395 | 476 | 202 | 208 | 189 | 39 | 37 | 39 | 32 | 34 | 35 | 657 | 646 | 799 | 248 | 275 | 266
| yA-A 2 | 50 | 48 | 35 | 29 | 22 | 27 | 548 | 596 | 551 | 249 | 241 | 237 | 43 | 45 | 34 | 32 | 26 | 35 | 816 | 904 | 846 | 270 | 289 | 269
| yA-A 3 | 44 | 37 | 25 | 25 | 20 | 22 | 604 | 546 | 485 | 203 | 223 | 261 | 35 | 42 | 31 | 37 | 38 | 32 | 806 | 827 | 703 | 223 | 233 | 266
| yA-A 4 | 29 | 32 | 36 | 22 | 15 | 18 | 342 | 392 | 346 | 190 | 192 | 175 | 33 | 33 | 37 | 29 | 24 | 28 | 625 | 690 | 654 | 218 | 227 | 226
| yA-A 5 | 59 | 53 | 44 | 20 | 23 | 22 | 491 | 506 | 537 | 223 | 235 | 203 | 33 | 41 | 37 | 38 | 33 | 38 | 701 | 735 | 796 | 238 | 245 | 220
| yA-A 6 | 39 | 44 | 41 | ND | 20 | ND | 212 | 422 | 422 | 168 | 181 | 195 | 33 | 44 | 36 | 26 | 32 | 35 | 414 | 770 | 754 | 210 | 228 | 253
Table 15 LC-MRM Intra Assay QC Data. Please report the values of each of the three replicates separately for each compound and standard. QCs have to be measured in one occasion only. The independent 6 replicates injected 3 times each
| Lowest Level QC | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| LLQC 1 | 32 | 47 | 31 | 27 | 22 | 24 | 322 | 306 | 293 | 147 | 131 | 138 | 34 | 57 | 36 | 39 | 33 | 33 | 555 | 841 | 659 | 253 | 252 | 246
| LLQC 2 | 28 | 26 | 30 | 20 | ND | ND | 208 | 224 | 269 | 118 | 128 | 121 | 33 | 32 | 33 | 30 |  |  | 732 | 667 | 754 | 222 | 245 | 227
| LLQC 3 | 33 | 26 | 23 | 24 | 17 | 23 | 284 | 249 | 276 | 125 | 140 | 127 | 41 | 33 | 39 | 35 | 26 | 34 | 719 | 720 | 780 | 247 | 242 | 264
| LLQC 4 | 38 | 29 | 28 | 21 | 17 | 22 | 281 | 250 | 260 | 137 | 139 | 136 | 38 | 25 | 32 | 35 | 22 | 29 | 707 | 777 | 564 | 225 | 206 | 254
| LLQC 5 | 37 | 32 | 38 | 20 | 15 | 21 | 223 | 189 | 282 | 107 | 114 | 129 | 38 | 42 | 41 | 32 | 28 | 33 | 784 | 653 | 744 | 213 | 229 | 236
| LLQC 6 | 25 | ND | 28 | 17 | 24 |  | 250 | 259 | 224 | 143 | 120 | 140 | 35 |  | 45 | 25 | 36 | 29 | 569 | 688 | 652 | 238 | 208 | 250
| Low QC | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| LQC 1 | 29 | 32 | 32 | 28 | 27 | 32 | 306 | 434 | 358 | 168 | 205 | 189 | 28 | 28 | 26 | 29 | 26 | 30 | 555 | 841 | 659 | 227 | 235 | 261
| LQC 2 | 39 | 51 | 47 | 29 | 18 | 26 | 403 | 351 | 425 | 182 | 186 | 184 | 33 | 37 | 41 | 34 | 30 | 34 | 732 | 667 | 754 | 218 | 255 | 229
| LQC 3 | 41 | 30 | 41 | 20 | 30 | 23 | 386 | 395 | 404 | 174 | 165 | 159 | 43 | 36 | 42 | 27 | 34 | 33 | 719 | 720 | 780 | 228 | 208 | 232
| LQC 4 | 33 | 45 | 46 | 22 | 25 | 27 | 425 | 381 | 328 | 149 | 165 | 150 | 34 | 36 | 37 | 25 | 31 | 32 | 707 | 777 | 564 | 197 | 211 | 198
| LQC 5 | 45 | 37 | 39 | 30 | 25 | 25 | 401 | 351 | 394 | 154 | 173 | 190 | 40 | 40 | 41 | 37 | 35 | 27 | 784 | 653 | 744 | 186 | 207 | 235
| LQC 6 | 28 | 35 | 34 | 28 | 33 | 29 | 311 | 339 | 367 | 171 | 175 | 183 | 30 | 31 | 34 | 25 | 32 | 29 | 569 | 688 | 652 | 233 | 209 | 235
| Middle QC | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| MQC 1 | 58 | 43 | 56 | 32 | 40 | 38 | 720 | 669 | 660 | 324 | 344 | 309 | 34 | 28 | 35 | 26 | 27 | 29 | 665 | 626 | 625 | 202 | 218 | 207
| MQC 2 | 62 | 52 | 61 | 24 | 38 | 34 | 775 | 716 | 706 | 273 | 300 | 327 | 46 | 28 | 34 | 32 | 31 | 28 | 696 | 687 | 648 | 205 | 214 | 207
| MQC 3 | 65 | 49 | 50 | 40 | 28 | 41 | 599 | 596 | 695 | 271 | 341 | 300 | 41 | 25 | 28 | 29 | 22 | 27 | 568 | 544 | 575 | 189 | 190 | 185
| MQC 4 | 48 | 54 | 60 | 32 | 36 | 40 | 734 | 717 | 741 | 260 | 307 | 292 | 29 | 37 | 41 | 22 | 27 | 29 | 613 | 648 | 697 | 175 | 203 | 219
| MQC 5 | 55 | 60 | 32 | 36 | 33 | 25 | 595 | 599 | 650 | 293 | 289 | 282 | 28 | 33 | 32 | 35 | 26 | 22 | 586 | 600 | 654 | 196 | 185 | 206
| MQC 6 | 45 | 60 | 43 | 27 | 26 | 24 | 578 | 654 | 584 | 238 | 292 | 253 | 23 | 33 | 25 | 29 | 26 | 24 | 593 | 632 | 646 | 154 | 185 | 189
| High QC | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| HQC 1 | 65 | 55 | 63 | 48 | 40 | 47 | 703 | 699 | 769 | 400 | 383 | 431 | 31 | 28 | 38 | 25 | 28 | 26 | 564 | 569 | 680 | 195 | 217 | 208
| HQC 2 | 55 | 54 | 62 | 48 | 50 | 42 | 814 | 919 | 880 | 499 | 467 | 467 | 26 | 28 | 31 | 25 | 26 | 27 | 605 | 687 | 596 | 211 | 209 | 225
| HQC 3 | 72 | 61 | 65 | 47 | 49 | 48 | 890 | 942 | 721 | 470 | 476 | 456 | 33 | 29 | 27 | 22 | 32 | 30 | 682 | 675 | 535 | 196 | 192 | 212
| HQC 4 | 64 | 67 | 63 | 55 | 59 | 66 | 816 | 843 | 958 | 471 | 445 | 460 | 28 | 34 | 32 | 27 | 25 | 31 | 649 | 662 | 641 | 181 | 183 | 211
| HQC 5 | 60 | 65 | 63 | 55 | 48 | 47 | 736 | 761 | 839 | 387 | 420 | 426 | 30 | 29 | 37 | 30 | 32 | 29 | 608 | 621 | 730 | 199 | 193 | 192
| HQC 6 | 59 | 77 | 71 | 41 | 34 | 41 | 716 | 768 | 806 | 429 | 451 | 475 | 27 | 33 | 29 | 22 | 30 | 27 | 590 | 656 | 636 | 198 | 173 | 179
| Highest Level QC | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| HLQC 1 | 53 | 53 | 65 | 54 | 71 | 70 | 905 | 915 | 937 | 598 | 589 | 605 | 23 | 22 | 31 | 21 | 26 | 28 | 680 | 607 | 644 | 201 | 232 | 209
| HLQC 2 | 87 | 83 | 93 | 66 | 68 | 61 | 872 | 1008 | 1080 | 632 | 629 | 616 | 31 | 26 | 32 | 24 | 32 | 27 | 528 | 597 | 614 | 192 | 188 | 163
| HLQC 3 | 77 | 83 | 88 | 43 | 69 | 61 | 906 | 968 | 929 | 321 | 651 | 648 | 34 | 32 | 38 | 18 | 28 | 26 | 637 | 608 | 612 | 187 | 194 | 216
| HLQC 4 | 69 | 83 | 76 | 62 | 72 | 65 | 1006 | 1064 | 1118 | 602 | 611 | 573 | 32 | 36 | 31 | 26 | 26 | 29 | 608 | 612 | 624 | 175 | 163 | 171
| HLQC 5 | 86 | 96 | 83 | 76 | 91 | 72 | 1007 | 1019 | 1068 | 616 | 634 | 617 | 32 | 30 | 29 | 27 | 27 | 25 | 654 | 691 | 686 | 197 | 201 | 196
| HLQC 6 | 82 | 77 | 74 | 64 | 54 | 63 | 859 | 731 | 999 | 604 | 630 | 623 | 27 | 22 | 24 | 25 | 25 | 30 | 519 | 453 | 632 | 195 | 173 | 207
| Blank QC | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| Matrix Blank | 263 | 81 | - | 189 | 44 | - | 3908 | 898 | ND | 1615 | 340 | ND | ND | ND | - | ND | ND | - | ND | ND | - | ND | ND | -
| Total Blank | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND
BATCH 2 - QC analysis
INTRA ASSAY - QC RESULTS | Calibration Line 1 | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| STD 1 | 814 | 808 | 809 | 1202 | 1288 | 1268 | 14608 | 12250 | 11471 | 10499 | 14028 | 13160 | 37 | 33 | 32 | 41 | 34 | 31 | 918 | 699 | 690 | 228 | 301 | 293
| STD 2 | 465 | 429 | 458 | 660 | 645 | 623 | 7559 | 7033 | 5608 | 6226 | 6337 | 6460 | 42 | 34 | 40 | 41 | 40 | 37 | 1005 | 820 | 689 | 276 | 247 | 273
| STD 3 | 49 | 38 | 44 | 56 | 76 | 55 | 565 | 666 | 686 | 576 | 624 | 589 | 38 | 35 | 42 | 42 | 39 | 39 | 819 | 911 | 932 | 284 | 256 | 275
| STD 4 | 15 | 24 | 24 | 35 | 33 | 40 | 273 | 324 | 283 | 278 | 299 | 305 | 32 | 33 | 39 | 27 | 41 | 35 | 797 | 783 | 743 | 290 | 274 | 270
| STD 5 | 13 | 13 | 16 | 22 | 20 | 12 | 182 | 227 | 208 | 177 | 190 | 185 | 33 | 35 | 30 | 32 | 32 | 29 | 916 | 853 | 856 | 249 | 207 | 262
| STD 6 | 3 | 2 | 4 | 3 | 8 | 6 | 54 | 56 | 54 | 43 | 53 | 60 | 41 | 36 | 40 | 36 | 31 | 42 | 796 | 820 | 856 | 324 | 238 | 291
Table 13 LC-MRM Calibration Line 2 Data (End). Please report the values of each of the three replicates separately for each compound and standard. Each STD prepared in triplicate, injected once from each vial
| Calibration Line 2 | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| STD 1 | 620 | 651 | 686 | 1001 | 945 | 922 | 11110 | 12317 | 11718 | 10042 | 10382 | 9254 | 27 | 32 | 32 | 23 | 28 | 19 | 672 | 753 | 628 | 195 | 217 | 192
| STD 2 | 348 | 333 | 330 | 487 | 502 | 430 | 6843 | 5529 | 5566 | 4579 | 5063 | 4849 | 29 | 29 | 30 | 26 | 29 | 26 | 726 | 737 | 679 | 177 | 236 | 210
| STD 3 | 31 | 27 | 32 | 45 | 48 | 48 | 659 | 485 | 552 | 454 | 488 | 477 | 31 | 27 | 28 | 31 | 34 | 23 | 756 | 741 | 691 | 210 | 231 | 215
| STD 4 | 15 | 23 | 18 | 27 | 19 | 27 | 238 | 211 | 238 | 252 | 243 | 232 | 28 | 33 | 31 | 28 | 27 | 30 | 606 | 683 | 729 | 217 | 251 | 241
| STD 5 | 9 | 9 | 12 | 17 | 11 | 14 | 159 | 172 | 178 | 168 | 133 | 165 | 36 | 30 | 36 | 32 | 25 | 33 | 686 | 625 | 753 | 218 | 215 | 216
| STD 6 | ND | 3 | 3 | 3 | ND | ND | 34 | 31 | 39 | 39 | 39 | 39 | 35 | 29 | 32 | 23 | 24 | 11 | 624 | 656 | 634 | 223 | 213 | 194
| Blank QC | Area endogenous Cer | [11 empty] | Labelled IS Area
|  | Cer d18:1/16:0 |  |  | Cer d18:1/18:0 |  |  | Cer d18:1/24:0 |  |  | Cer d18:1/24:1 |  |  | d7-Cer d18:1/16:0 |  |  | d7-Cer d18:1/18:0 |  |  | d7-Cer d18:1/24:0 |  |  | d7-Cer d18:1/24:1
|  | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3 | Repl 1 | Repl 2 | Repl 3
| Matrix Blank | 81 | 47 | - | 46 | 30 | - | 983 | 527 | - | 466 | 286 | - | ND | ND | - | ND | ND | - | ND | ND | - | ND | ND | -
| Total Blank | ND | - | ND | ND | - | ND | ND | - | ND | ND | - | ND | ND | - | ND | ND | - | ND | ND | - | ND | ND | - | ND

## Raw Peak Areas
Sample_Name | Cer d18:1/16:0 [M-H2O+H]+ | Cer d18:1/16:0-d7 [M-H2O+H]+ | Cer d18:1/18:0 [M-H2O+H]+ | Cer d18:1/18:0-d7 [M-H2O+H]+ | Cer d18:1/24:0 [M-H2O+H]+ | Cer d18:1/24:0-d7 [M-H2O+H]+ | Cer d18:1/24:1 [M-H2O+H]+ | Cer d18:1/24:1-d7 [M-H2O+H]+
Total_Blank1 | ND | ND | ND | ND | ND | ND | ND | ND
Matrix_Blank_1 | 263 | ND | 189 | ND | 3908 | ND | 1615 | ND
Cal_A_STD6 | 23 | 141 | 10 | 142 | 182 | 3258 | 148 | 960
Cal_B_STD6 | 13 | 86 | 21 | 103 | 181 | 2712 | 147 | 812
Cal_C_STD6 | 10 | 74 | 11 | 88 | 121 | 1919 | 109 | 644
Cal_A_STD5 | 13 | 84 | 8 | 80 | 131 | 1906 | 104 | 465
Cal_B_STD5 | 9 | 66 | 19 | 64 | 135 | 1676 | 104 | 477
Cal_C_STD5 | 8 | 72 | 10 | 57 | 129 | 1585 | 95 | 441
Cal_A_STD4 | 13 | 58 | 20 | 54 | 238 | 1613 | 189 | 490
Cal_B_STD4 | 16 | 65 | 25 | 57 | 222 | 1550 | 184 | 483
Cal_C_STD4 | 10 | 47 | 17 | 48 | 236 | 1454 | 145 | 416
Cal_A_STD3_dilution1/3 | 25 | 63 | 33 | 59 | 312 | 1849 | 307 | 455
Cal_B_STD3_dilution1/3 | 32 | 57 | 32 | 48 | 311 | 1266 | 325 | 423
Cal_C_STD3_dilution1/3 | 19 | 47 | 34 | 54 | 328 | 1412 | 260 | 362
Cal_A_STD3_dilution1/2 | 34 | 54 | 50 | 53 | 499 | 1383 | 389 | 391
Cal_B_STD3_dilution1/2 | 34 | 60 | 46 | 51 | 494 | 1349 | 439 | 377
Cal_C_STD3_dilution1/2 | 29 | 47 | 39 | 49 | 487 | 1415 | 474 | 425
Cal_A_STD3 | 67 | 57 | 80 | 42 | 966 | 1426 | 685 | 323
Cal_B_STD3 | 44 | 46 | 87 | 48 | 999 | 1368 | 888 | 409
Cal_C_STD3 | 48 | 53 | 51 | 51 | 959 | 1284 | 760 | 391
Cal_A_STD2 | 534 | 48 | 837 | 46 | 10285 | 1281 | 7426 | 335
Cal_B_STD2 | 621 | 57 | 930 | 47 | 10174 | 1097 | 9028 | 379
Cal_C_STD2 | 443 | 48 | 736 | 40 | 10617 | 1218 | 8750 | 373
Cal_A_STD1 | 1118 | 53 | 1770 | 49 | 17888 | 1110 | 17284 | 412
Cal_B_STD1 | 1011 | 42 | 1670 | 38 | 16621 | 925 | 18707 | 363
Cal_C_STD1 | 1084 | 49 | 1595 | 41 | 12253 | 796 | 16898 | 358
pool_QC1A | 61 | 52 | 38 | 44 | 747 | 1032 | 388 | 341
pool_QC1B | 55 | 48 | 35 | 37 | 784 | 1070 | 361 | 324
SRM1950_1A | 86 | 52 | 69 | 45 | 975 | 1020 | 480 | 326
SRM1950_1B | 85 | 48 | 59 | 49 | 1056 | 1037 | 427 | 275
SRM1950_1C | 86 | 48 | 55 | 34 | 1091 | 1067 | 463 | 305
SRM1950_2A | 51 | 30 | 24 | 24 | 553 | 791 | 253 | 245
SRM1950_2B | 46 | 31 | 32 | 38 | 485 | 764 | 222 | 226
SRM1950_2C | 57 | 43 | 26 | 35 | 468 | 784 | 233 | 230
SRM1950_3A | 87 | 44 | 47 | 38 | 942 | 933 | 459 | 271
SRM1950_3B | 94 | 49 | 59 | 37 | 1021 | 978 | 446 | 289
SRM1950_3C | 48 | 20 | 50 | 37 | 798 | 812 | 449 | 290
SRM1950_4A | 76 | 36 | 51 | 42 | 855 | 805 | 428 | 292
SRM1950_4B | 93 | 41 | 46 | 37 | 918 | 975 | 383 | 268
SRM1950_4C | 72 | 37 | 40 | 41 | 939 | 1001 | 409 | 266
SRM1950_5A | 77 | 40 | 48 | 37 | 830 | 784 | 454 | 284
SRM1950_5B | 75 | 41 | 39 | 31 | 829 | 867 | 442 | 303
SRM1950_5C | 60 | 37 | 38 | 41 | 928 | 935 | 470 | 297
SRM1950_6A | 82 | 33 | 44 | 25 | 917 | 951 | 377 | 271
SRM1950_6B | 65 | 28 | 37 | 30 | 825 | 793 | 428 | 302
SRM1950_6C | 77 | 44 | 51 | 38 | 923 | 935 | 388 | 272
pool_QC2A | 42 | 32 | 28 | 24 | 640 | 943 | 283 | 236
pool_QC2B | 43 | 38 | 31 | 34 | 606 | 901 | 303 | 273
T1D_1A | 77 | 29 | 41 | 35 | 829 | 900 | 327 | 250
T1D_1B | 67 | 34 | 42 | 35 | 832 | 941 | 360 | 260
T1D_1C | 68 | 34 | 62 | 38 | 666 | 706 | 357 | 275
T1D_2A | 70 | 37 | 59 | 40 | 902 | 900 | 388 | 268
T1D_2B | 73 | 37 | 53 | 39 | 907 | 1003 | 380 | 246
T1D_2C | 64 | 43 | 57 | 32 | 872 | 925 | 412 | 284
T1D_3A | 65 | 41 | 67 | 37 | 927 | 995 | 386 | 268
T1D_3B | 74 | 40 | 53 | 33 | 953 | 958 | 387 | 262
T1D_3C | 53 | 33 | 51 | 33 | 958 | 908 | 402 | 247
T1D_4A | 56 | 32 | 33 | 24 | 723 | 749 | 345 | 268
T1D_4B | 63 | 32 | 39 | 25 | 717 | 725 | 320 | 275
T1D_4C | 50 | 33 | 40 | 31 | 780 | 777 | 357 | 235
T1D_5A | 65 | 35 | 53 | 31 | 866 | 847 | 354 | 273
T1D_5B | 80 | 41 | 56 | 37 | 921 | 934 | 398 | 293
T1D_5C | 69 | 33 | 53 | 34 | 834 | 866 | 240 | 172
T1D_6A | 65 | 36 | 54 | 38 | 862 | 826 | 375 | 280
T1D_6B | 70 | 31 | 64 | 38 | 850 | 855 | 375 | 277
T1D_6C | 71 | 29 | 59 | 33 | 872 | 915 | 401 | 258
Total_Blank_2 | ND | ND | ND | ND | ND | ND | ND | ND
pool_QC3A | 50 | 34 | 36 | 35 | 634 | 808 | 313 | 261
pool_QC3B | 51 | 37 | 37 | 37 | 649 | 873 | 295 | 273
yAA_1A | 41 | 34 | 25 | 28 | 419 | 657 | 222 | 251
yAA_1B | 38 | 37 | 23 | 34 | 395 | 646 | 208 | 281
yAA_1C | 55 | 36 | 21 | 33 | 476 | 799 | 191 | 175
yAA_2A | 50 | 41 | 29 | 26 | 548 | 816 | 274 | 281
yAA_2B | 48 | 40 | 14 | 25 | 596 | 904 | 263 | 291
yAA_2C | 35 | 32 | 24 | 35 | 551 | 846 | 260 | 275
yAA_3A | 44 | 28 | 19 | 36 | 604 | 806 | 220 | 229
yAA_3B | 37 | 39 | 23 | 33 | 546 | 827 | 242 | 242
yAA_3C | 25 | 29 | 22 | 30 | 485 | 703 | 261 | 266
yAA_4A | 29 | 32 | 22 | 25 | 342 | 625 | 208 | 219
yAA_4B | 32 | 32 | 15 | 20 | 392 | 690 | 193 | 238
yAA_4C | 36 | 36 | 18 | 23 | 346 | 654 | 204 | 226
yAA_5A | 59 | 30 | 21 | 37 | 491 | 701 | 242 | 239
yAA_5B | 53 | 36 | 16 | 30 | 506 | 735 | 252 | 249
yAA_5C | 44 | 34 | 22 | 35 | 537 | 796 | 218 | 228
yAA_6A | 39 | 27 | 11 | 26 | 212 | 414 | 182 | 212
yAA_6B | 44 | 35 | 15 | 31 | 422 | 770 | 197 | 232
yAA_6C | 41 | 34 | 15 | 29 | 422 | 754 | 205 | 258
pool_QC4A | 36 | 28 | 25 | 33 | 589 | 806 | 309 | 251
pool_QC4B | 50 | 36 | 31 | 35 | 580 | 790 | 292 | 254
hTAG_1A | 58 | 26 | 48 | 31 | 1161 | 864 | 513 | 243
hTAG_1B | 68 | 35 | 45 | 30 | 1078 | 833 | 494 | 259
hTAG_1C | 72 | 26 | 64 | 34 | 1206 | 825 | 379 | 198
hTAG_2A | 63 | 36 | 43 | 25 | 1116 | 814 | 430 | 241
hTAG_2B | 72 | 30 | 48 | 30 | 1067 | 849 | 345 | 184
hTAG_2C | 67 | 44 | 46 | 36 | 1166 | 868 | 443 | 226
hTAG_3A | 73 | 29 | 43 | 35 | 636 | 533 | 308 | 194
hTAG_3B | 60 | 26 | 33 | 25 | 892 | 796 | 386 | 228
hTAG_3C | 56 | 28 | 35 | 29 | 723 | 694 | 333 | 236
hTAG_4A | 54 | 34 | 38 | 29 | 820 | 706 | 360 | 199
hTAG_4B | 46 | 24 | 38 | 26 | 863 | 707 | 378 | 242
hTAG_4C | 59 | 25 | 35 | 23 | 924 | 742 | 342 | 214
hTAG_5A | 51 | 28 | 27 | 20 | 637 | 542 | 321 | 207
hTAG_5B | 53 | 33 | 29 | 23 | 745 | 610 | 351 | 187
hTAG_5C | 52 | 34 | 26 | 26 | 730 | 675 | 349 | 217
hTAG_6A | 46 | 42 | 36 | 30 | 669 | 651 | 268 | 223
hTAG_6B | 50 | 37 | 36 | 30 | 771 | 767 | 264 | 214
hTAG_6C | 48 | 34 | 29 | 31 | 736 | 806 | 304 | 223
pool_QC5A | 45 | 42 | 25 | 29 | 591 | 794 | 225 | 203
pool_QC5B | 43 | 32 | 26 | 29 | 610 | 853 | 265 | 246
Cal_D_STD1 | 805 | 39 | 884 | 28 | 13194 | 825 | 12810 | 267
Cal_E_STD1 | 809 | 46 | 1147 | 31 | 15569 | 922 | 12081 | 273
Cal_F_STD1 | 687 | 36 | 1332 | 39 | 16751 | 922 | 13788 | 269
Cal_D_STD2 | 308 | 31 | 565 | 30 | 6450 | 752 | 6954 | 276
Cal_E_STD2 | 355 | 39 | 597 | 33 | 7683 | 1010 | 6266 | 297
Cal_F_STD2 | 385 | 33 | 641 | 36 | 6346 | 793 | 6250 | 309
Cal_D_STD3 | 41 | 37 | 74 | 43 | 674 | 954 | 539 | 234
Cal_E_STD3 | 45 | 37 | 58 | 34 | 640 | 891 | 589 | 281
Cal_F_STD3 | 44 | 36 | 62 | 31 | 613 | 865 | 589 | 272
Cal_D_STD3_dilution1/2 | 17 | 48 | 36 | 41 | 328 | 903 | 280 | 249
Cal_E_STD3_dilution1/2 | 19 | 43 | 23 | 29 | 243 | 749 | 253 | 286
Cal_F_STD3_dilution1/2 | 24 | 46 | 31 | 37 | 320 | 905 | 287 | 297
Cal_D_STD3_dilution1/3 | 11 | 34 | 17 | 30 | 213 | 831 | 197 | 281
Cal_E_STD3_dilution1/3 | 17 | 39 | 21 | 30 | 151 | 766 | 195 | 280
Cal_F_STD3_dilution1/3 | 14 | 39 | 21 | 34 | 207 | 878 | 189 | 273
Cal_D_STD4 | 7 | 37 | 13 | 29 | 127 | 843 | 108 | 285
Cal_E_STD4 | 12 | 48 | 11 | 29 | 143 | 864 | 103 | 286
Cal_F_STD4 | 9 | 37 | 17 | 39 | 127 | 963 | 97 | 231
Cal_D_STD5 | 4 | 33 | 10 | 33 | 70 | 903 | 60 | 269
Cal_E_STD5 | 5 | 41 | 6 | 37 | 65 | 768 | 51 | 292
Cal_F_STD5 | 5 | 24 | 7 | 22 | 62 | 803 | 44 | 258
Cal_D_STD6 | 5 | 35 | 3 | 36 | 40 | 655 | 52 | 263
Cal_E_STD6 | 4 | 33 | 2 | 28 | 56 | 794 | 45 | 281
Cal_F_STD6 | 5 | 39 | 2 | 34 | 54 | 965 | 42 | 236
Matrix_Blank_2 | 81 | ND | 44 | ND | 898 | ND | 340 | ND
Total_Blank_3 | ND | ND | ND | ND | ND | ND | ND | ND
BATCH 2
Total_Blank_1 | ND | ND | ND | ND | ND | ND | ND | ND
Matrix_Blank_1 | 88 | ND | 46 | ND | 983 | ND | 466 | ND
Cal_A_STD6 | 3 | 41 | 3 | 36 | 54 | 796 | 43 | 324
Cal_B_STD6 | 2 | 36 | 8 | 31 | 56 | 820 | 53 | 238
Cal_C_STD6 | 4 | 40 | 6 | 42 | 54 | 856 | 60 | 291
Cal_A_STD5 | 5 | 36 | 6 | 34 | 59 | 847 | 53 | 242
Cal_B_STD5 | 4 | 47 | 4 | 37 | 62 | 742 | 52 | 279
Cal_C_STD5 | 5 | 39 | 7 | 38 | 54 | 769 | 55 | 297
Cal_A_STD4 | 10 | 45 | 12 | 35 | 107 | 758 | 115 | 283
Cal_B_STD4 | 11 | 42 | 12 | 28 | 124 | 871 | 112 | 238
Cal_C_STD4 | 10 | 37 | 8 | 35 | 117 | 873 | 118 | 212
Cal_A_STD3_dilution1/3 | 13 | 33 | 22 | 32 | 182 | 916 | 177 | 249
Cal_B_STD3_dilution1/3 | 13 | 35 | 20 | 32 | 227 | 853 | 190 | 207
Cal_C_STD3_dilution1/3 | 16 | 30 | 12 | 29 | 208 | 856 | 185 | 262
Cal_A_STD3_dilution1/2 | 15 | 32 | 35 | 27 | 273 | 797 | 278 | 290
Cal_B_STD3_dilution1/2 | 24 | 33 | 33 | 41 | 324 | 783 | 299 | 274
Cal_C_STD3_dilution1/2 | 24 | 39 | 40 | 35 | 283 | 743 | 305 | 270
Cal_A_STD3 | 49 | 38 | 56 | 42 | 565 | 819 | 576 | 284
Cal_B_STD3 | 38 | 35 | 76 | 39 | 666 | 911 | 624 | 256
Cal_C_STD3 | 44 | 42 | 55 | 39 | 686 | 932 | 589 | 275
Cal_A_STD2 | 465 | 42 | 660 | 41 | 7559 | 1005 | 6226 | 276
Cal_B_STD2 | 429 | 34 | 645 | 40 | 7033 | 820 | 6337 | 247
Cal_C_STD2 | 458 | 40 | 623 | 37 | 5608 | 689 | 6460 | 273
Cal_A_STD1 | 814 | 37 | 1202 | 41 | 14608 | 918 | 10499 | 228
Cal_B_STD1 | 808 | 33 | 1288 | 34 | 12250 | 699 | 14028 | 301
Cal_C_STD1 | 809 | 32 | 1268 | 31 | 11471 | 690 | 13160 | 293
LLOQ_1A | 32 | 32 | 21 | 32 | 322 | 812 | 148 | 254
LLOQ_1B | 42 | 50 | 18 | 32 | 306 | 787 | 133 | 254
LLOQ_1C | 22 | 26 | 11 | 30 | 293 | 795 | 138 | 251
LLOQ_2A | 27 | 31 | 19 | 30 | 208 | 709 | 123 | 225
LLOQ_2B | 23 | 27 | 15 | 21 | 224 | 629 | 129 | 245
LLOQ_2C | 28 | 30 | 22 | 38 | 269 | 803 | 122 | 228
LLOQ_3A | 34 | 41 | 18 | 33 | 284 | 744 | 125 | 251
LLOQ_3B | 25 | 33 | 10 | 26 | 249 | 713 | 141 | 245
LLOQ_3C | 22 | 43 | 23 | 31 | 276 | 772 | 130 | 264
LLOQ_4A | 31 | 40 | 21 | 36 | 281 | 763 | 138 | 226
LLOQ_4B | 17 | 21 | 17 | 28 | 250 | 755 | 141 | 209
LLOQ_4C | 16 | 34 | 19 | 24 | 260 | 662 | 136 | 253
LLOQ_5A | 30 | 34 | 18 | 27 | 223 | 757 | 110 | 214
LLOQ_5B | 23 | 35 | 15 | 26 | 189 | 541 | 114 | 229
LLOQ_5C | 27 | 39 | 21 | 33 | 282 | 758 | 129 | 238
LLOQ_6A | 20 | 32 | 17 | 23 | 250 | 653 | 146 | 243
LLOQ_6B | 33 | 42 | 24 | 35 | 259 | 752 | 123 | 208
LLOQ_6C | 26 | 45 | 21 | 25 | 224 | 580 | 141 | 251
LQC_1A | 29 | 25 | 18 | 25 | 306 | 555 | 169 | 232
LQC_1B | 32 | 21 | 22 | 22 | 434 | 841 | 221 | 237
LQC_1C | 32 | 24 | 20 | 27 | 358 | 659 | 191 | 262
LQC_2A | 39 | 33 | 29 | 29 | 403 | 732 | 187 | 219
LQC_2B | 38 | 33 | 21 | 24 | 351 | 667 | 187 | 255
LQC_2C | 38 | 41 | 26 | 34 | 425 | 754 | 186 | 231
LQC_3A | 36 | 32 | 19 | 24 | 386 | 719 | 191 | 236
LQC_3B | 36 | 34 | 26 | 30 | 395 | 720 | 178 | 212
LQC_3C | 41 | 41 | 20 | 27 | 404 | 780 | 174 | 238
LQC_4A | 38 | 33 | 21 | 22 | 425 | 707 | 166 | 204
LQC_4B | 37 | 36 | 21 | 28 | 381 | 777 | 186 | 214
LQC_4C | 40 | 34 | 27 | 29 | 328 | 564 | 164 | 201
LQC_5A | 41 | 38 | 30 | 36 | 401 | 784 | 166 | 192
LQC_5B | 33 | 39 | 25 | 29 | 351 | 653 | 186 | 213
LQC_5C | 39 | 34 | 25 | 21 | 394 | 744 | 190 | 241
LQC_6A | 28 | 30 | 29 | 23 | 311 | 569 | 171 | 240
LQC_6B | 35 | 27 | 35 | 29 | 339 | 688 | 175 | 214
LQC_6C | 39 | 36 | 22 | 26 | 367 | 652 | 183 | 243
MQC_1A | 58 | 29 | 32 | 24 | 720 | 665 | 325 | 205
MQC_1B | 43 | 24 | 40 | 23 | 669 | 626 | 347 | 219
MQC_1C | 56 | 26 | 38 | 22 | 660 | 625 | 316 | 211
MQC_2A | 54 | 29 | 24 | 30 | 775 | 696 | 279 | 205
MQC_2B | 54 | 25 | 38 | 24 | 716 | 687 | 296 | 232
MQC_2C | 61 | 31 | 34 | 24 | 706 | 648 | 330 | 207
MQC_3A | 58 | 32 | 31 | 27 | 599 | 568 | 271 | 188
MQC_3B | 52 | 22 | 28 | 20 | 596 | 544 | 348 | 207
MQC_3C | 50 | 21 | 41 | 23 | 695 | 575 | 302 | 191
MQC_4A | 27 | 21 | 32 | 17 | 734 | 613 | 260 | 178
MQC_4B | 50 | 33 | 36 | 24 | 717 | 648 | 307 | 204
MQC_4C | 51 | 38 | 40 | 25 | 741 | 697 | 292 | 219
MQC_5A | 55 | 25 | 36 | 31 | 595 | 586 | 295 | 196
MQC_5B | 54 | 28 | 33 | 24 | 599 | 600 | 289 | 186
MQC_5C | 34 | 19 | 25 | 18 | 650 | 654 | 283 | 206
MQC_6A | 39 | 18 | 27 | 29 | 578 | 593 | 214 | 158
MQC_6B | 60 | 32 | 26 | 25 | 654 | 632 | 297 | 179
MQC_6C | 43 | 19 | 24 | 15 | 584 | 646 | 254 | 193
HQC_1A | 62 | 29 | 48 | 22 | 703 | 564 | 403 | 196
HQC_1B | 48 | 24 | 40 | 26 | 699 | 569 | 385 | 217
HQC_1C | 50 | 21 | 47 | 24 | 769 | 680 | 438 | 211
HQC_2A | 55 | 25 | 48 | 23 | 814 | 605 | 506 | 211
HQC_2B | 50 | 25 | 50 | 20 | 919 | 687 | 472 | 216
HQC_2C | 59 | 30 | 30 | 23 | 880 | 596 | 470 | 255
HQC_3A | 49 | 19 | 47 | 20 | 890 | 682 | 480 | 196
HQC_3B | 61 | 25 | 41 | 31 | 942 | 675 | 477 | 193
HQC_3C | 65 | 25 | 48 | 23 | 721 | 535 | 457 | 212
HQC_4A | 64 | 26 | 55 | 23 | 816 | 649 | 420 | 187
HQC_4B | 67 | 33 | 59 | 22 | 843 | 662 | 445 | 183
HQC_4C | 55 | 24 | 50 | 28 | 958 | 641 | 460 | 218
HQC_5A | 47 | 22 | 40 | 26 | 736 | 608 | 392 | 205
HQC_5B | 65 | 24 | 37 | 26 | 761 | 621 | 427 | 194
HQC_5C | 57 | 31 | 47 | 22 | 839 | 730 | 428 | 195
HQC_6A | 56 | 24 | 43 | 16 | 716 | 590 | 433 | 201
HQC_6B | 77 | 33 | 34 | 26 | 768 | 656 | 464 | 174
HQC_6C | 61 | 23 | 40 | 27 | 806 | 636 | 491 | 176
HLQC_1A | 53 | 19 | 54 | 18 | 905 | 680 | 610 | 203
HLOC_1B | 53 | 19 | 71 | 24 | 915 | 607 | 593 | 232
HLOC_1C | 65 | 28 | 70 | 25 | 937 | 644 | 610 | 209
HLOC_2A | 85 | 28 | 66 | 18 | 872 | 528 | 632 | 194
HLOC_2B | 83 | 31 | 68 | 29 | 1008 | 597 | 630 | 188
HLOC_2C | 91 | 31 | 61 | 26 | 1080 | 614 | 626 | 167
HLOC_3A | 71 | 28 | 43 | 15 | 906 | 637 | 380 | 187
HLOC_3B | 83 | 30 | 69 | 25 | 968 | 608 | 659 | 194
HLOC_3C | 79 | 27 | 61 | 22 | 929 | 612 | 662 | 216
HLOC_4A | 68 | 26 | 62 | 24 | 1006 | 608 | 603 | 175
HLOC_4B | 75 | 31 | 72 | 22 | 1064 | 612 | 619 | 164
HLOC_4C | 67 | 27 | 57 | 24 | 1118 | 624 | 574 | 172
HLOC_5A | 83 | 26 | 76 | 21 | 1007 | 654 | 620 | 198
HLOC_5B | 101 | 28 | 91 | 23 | 1019 | 691 | 640 | 201
HLOC_5C | 84 | 28 | 59 | 25 | 1068 | 686 | 632 | 196
HLOC_6A | 82 | 24 | 64 | 23 | 859 | 519 | 605 | 195
HLOC_6B | 77 | 19 | 45 | 22 | 731 | 453 | 633 | 173
HLOC_6C | 65 | 20 | 63 | 30 | 999 | 632 | 629 | 208
Cal_D_STD1 | 620 | 27 | 1001 | 23 | 11110 | 672 | 10042 | 195
Cal_E_STD1 | 651 | 32 | 945 | 28 | 12317 | 753 | 10382 | 217
Cal_F_STD1 | 686 | 32 | 922 | 19 | 11718 | 628 | 9254 | 192
Cal_D_STD2 | 348 | 29 | 487 | 26 | 6843 | 726 | 4579 | 177
Cal_E_STD2 | 333 | 29 | 502 | 29 | 5529 | 737 | 5063 | 236
Cal_F_STD2 | 330 | 30 | 430 | 26 | 5566 | 679 | 4849 | 210
Cal_D_STD3 | 31 | 31 | 45 | 31 | 659 | 756 | 454 | 210
Cal_E_STD3 | 27 | 27 | 48 | 34 | 485 | 741 | 488 | 231
Cal_F_STD3 | 32 | 28 | 48 | 23 | 552 | 691 | 477 | 215
Cal_D_STD3_dilution1/2 | 15 | 28 | 27 | 28 | 238 | 606 | 252 | 217
Cal_E_STD3_dilution1/2 | 23 | 33 | 19 | 27 | 211 | 683 | 243 | 251
Cal_F_STD3_dilution1/2 | 18 | 31 | 27 | 30 | 238 | 729 | 232 | 241
Cal_D_STD3_dilution1/3 | 9 | 36 | 17 | 32 | 159 | 686 | 168 | 218
Cal_E_STD3_dilution1/3 | 9 | 30 | 11 | 25 | 172 | 625 | 133 | 215
Cal_F_STD3_dilution1/3 | 12 | 36 | 14 | 33 | 178 | 753 | 165 | 216
Cal_D_STD4 | 4 | 28 | 12 | 28 | 96 | 635 | 88 | 244
Cal_E_STD4 | 4 | 21 | 9 | 34 | 93 | 712 | 98 | 223
Cal_F_STD4 | 9 | 38 | 10 | 27 | 92 | 605 | 85 | 205
Cal_D_STD5 | 5 | 32 | 6 | 30 | 51 | 664 | 51 | 218
Cal_E_STD5 | 5 | 30 | 9 | 30 | 49 | 755 | 45 | 243
Cal_F_STD5 | 4 | 31 | 3 | 36 | 53 | 664 | 43 | 199
Cal_D_STD6 | ND | 35 | 3 | 23 | 34 | 624 | 39 | 223
Cal_E_STD6 | 3 | 29 | ND | 24 | 31 | 656 | 39 | 213
Cal_F_STD6 | 3 | 32 | ND | 11 | 39 | 634 | 39 | 194
Matrix_Blank_2 | 47 | ND | 30 | ND | 527 | ND | 286 | ND
Total_Blank_3 | ND | ND | ND | ND | ND | ND | ND | ND